## UMG 801 Physical
Group | Address | Format | RD/WR | Unit | Note | Doc. no. 2.053.265.1.a
basic device | 4100 | String | RD | - | Name
basic device | 4132 | String | RD | - | firmware version
basic device | 4164 | UInt64 | RD | - | hardware version
basic device | 4168 | UInt16 | RD | - | state (0=unavailable, 1=initializing_device, 2=initializing_modules, 3=operating, 4=error)
basic device | 4169 | UInt16 | RD | - | device type (0=5A, 1=undefined)
basic device | 4170 | UInt64 | RD | - | product number
basic device | 4174 | UInt64 | RD | - | serial number
Information Module 01 | 4178 | String | RD | - | Name
Information Module 01 | 4210 | String | RD | - | firmware version
Information Module 01 | 4242 | UInt64 | RD | - | hardware version
Information Module 01 | 4246 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 01 | 4247 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 01 | 4248 | UInt64 | RD | - | product number
Information Module 01 | 4252 | UInt64 | RD | - | serial number
Information Module 01 | 4256 | UInt32 | RD | - | ID
Information Module 02 | 4258 | String | RD | - | Name
Information Module 02 | 4290 | String | RD | - | firmware version
Information Module 02 | 4322 | UInt64 | RD | - | hardware version
Information Module 02 | 4326 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 02 | 4327 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 02 | 4328 | UInt64 | RD | - | product number
Information Module 02 | 4332 | UInt64 | RD | - | serial number
Information Module 02 | 4336 | UInt32 | RD | - | ID
Information Module 03 | 4338 | String | RD | - | Name
Information Module 03 | 4370 | String | RD | - | firmware version
Information Module 03 | 4402 | UInt64 | RD | - | hardware version
Information Module 03 | 4406 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 03 | 4407 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 03 | 4408 | UInt64 | RD | - | product number
Information Module 03 | 4412 | UInt64 | RD | - | serial number
Information Module 03 | 4416 | UInt32 | RD | - | ID
Information Module 04 | 4418 | String | RD | - | Name
Information Module 04 | 4450 | String | RD | - | firmware version
Information Module 04 | 4482 | UInt64 | RD | - | hardware version
Information Module 04 | 4486 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 04 | 4487 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 04 | 4488 | UInt64 | RD | - | product number
Information Module 04 | 4492 | UInt64 | RD | - | serial number
Information Module 04 | 4496 | UInt32 | RD | - | ID
Information Module 05 | 4498 | String | RD | - | Name
Information Module 05 | 4530 | String | RD | - | firmware version
Information Module 05 | 4562 | UInt64 | RD | - | hardware version
Information Module 05 | 4566 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 05 | 4567 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 05 | 4568 | UInt64 | RD | - | product number
Information Module 05 | 4572 | UInt64 | RD | - | serial number
Information Module 05 | 4576 | UInt32 | RD | - | ID
Information Module 06 | 4578 | String | RD | - | Name
Information Module 06 | 4610 | String | RD | - | firmware version
Information Module 06 | 4642 | UInt64 | RD | - | hardware version
Information Module 06 | 4646 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 06 | 4647 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 06 | 4648 | UInt64 | RD | - | product number
Information Module 06 | 4652 | UInt64 | RD | - | serial number
Information Module 06 | 4656 | UInt32 | RD | - | ID
Information Module 07 | 4658 | String | RD | - | Name
Information Module 07 | 4690 | String | RD | - | firmware version
Information Module 07 | 4722 | UInt64 | RD | - | hardware version
Information Module 07 | 4726 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 07 | 4727 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 07 | 4728 | UInt64 | RD | - | product number
Information Module 07 | 4732 | UInt64 | RD | - | serial number
Information Module 07 | 4736 | UInt32 | RD | - | ID
Information Module 08 | 4738 | String | RD | - | Name
Information Module 08 | 4770 | String | RD | - | firmware version
Information Module 08 | 4802 | UInt64 | RD | - | hardware version
Information Module 08 | 4806 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 08 | 4807 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 08 | 4808 | UInt64 | RD | - | product number
Information Module 08 | 4812 | UInt64 | RD | - | serial number
Information Module 08 | 4816 | UInt32 | RD | - | ID
Information Module 09 | 4818 | String | RD | - | Name
Information Module 09 | 4850 | String | RD | - | firmware version
Information Module 09 | 4882 | UInt64 | RD | - | hardware version
Information Module 09 | 4886 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 09 | 4887 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 09 | 4888 | UInt64 | RD | - | product number
Information Module 09 | 4892 | UInt64 | RD | - | serial number
Information Module 09 | 4896 | UInt32 | RD | - | ID
Information Module 10 | 4898 | String | RD | - | Name
Information Module 10 | 4930 | String | RD | - | firmware version
Information Module 10 | 4962 | UInt64 | RD | - | hardware version
Information Module 10 | 4966 | UInt16 | RD | - | state (0=uninitialized, 1=initializing, 2=operating, 3=error, 4=missing)
Information Module 10 | 4967 | UInt16 | RD | - | device type (0=CT8_5A, 1=DI_14, 2=CT24_LP, 3=CT8_LP)
Information Module 10 | 4968 | UInt64 | RD | - | product number
Information Module 10 | 4972 | UInt64 | RD | - | serial number
Information Module 10 | 4976 | UInt32 | RD | - | ID
basic device | 5000 | Float | RD | V | Average Value L1 Voltage L-N (RMS)
basic device | 5002 | Float | RD | V | Average Value L2 Voltage L-N (RMS)
basic device | 5004 | Float | RD | V | Average Value L3 Voltage L-N (RMS)
basic device | 5006 | Float | RD | V | Average Value L1-L2 Voltage L-L (RMS)
basic device | 5008 | Float | RD | V | Average Value L2-L3 Voltage L-L (RMS)
basic device | 5010 | Float | RD | V | Average Value L3-L1 Voltage L-L (RMS)
Basic device measuring group 01 | 5012 | Float | RD | A | Average Value L1 Current (RMS)
Basic device measuring group 01 | 5014 | Float | RD | A | Average Value L2 Current (RMS)
Basic device measuring group 01 | 5016 | Float | RD | A | Average Value L3 Current (RMS)
Basic device measuring group 01 | 5018 | Float | RD | A | Average Value L4 Current (RMS)
Basic device measuring group 01 | 5020 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Basic device measuring group 01 | 5022 | Float | RD | W | Average Value L1 Active Power
Basic device measuring group 01 | 5024 | Float | RD | W | Average Value L2 Active Power
Basic device measuring group 01 | 5026 | Float | RD | W | Average Value L3 Active Power
Basic device measuring group 01 | 5028 | Float | RD | W | Average Value L4 Active Power
Basic device measuring group 01 | 5030 | Float | RD | W | Average Value Sum L1..L4 Active Power
Basic device measuring group 01 | 5032 | Float | RD | var | Average Value L1 Reactive Power
Basic device measuring group 01 | 5034 | Float | RD | var | Average Value L2 Reactive Power
Basic device measuring group 01 | 5036 | Float | RD | var | Average Value L3 Reactive Power
Basic device measuring group 01 | 5038 | Float | RD | var | Average Value L4 Reactive Power
Basic device measuring group 01 | 5040 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Basic device measuring group 01 | 5042 | Float | RD | VA | Average Value L1 Apparent Power
Basic device measuring group 01 | 5044 | Float | RD | VA | Average Value L2 Apparent Power
Basic device measuring group 01 | 5046 | Float | RD | VA | Average Value L3 Apparent Power
Basic device measuring group 01 | 5048 | Float | RD | VA | Average Value L4 Apparent Power
Basic device measuring group 01 | 5050 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Basic device measuring group 01 | 5052 | Float | RD | - | Average Value L1 Cos(phi)
Basic device measuring group 01 | 5054 | Float | RD | - | Average Value L2 Cos(phi)
Basic device measuring group 01 | 5056 | Float | RD | - | Average Value L3 Cos(phi)
Basic device measuring group 01 | 5058 | Float | RD | - | Average Value L4 Cos(phi)
Basic device measuring group 01 | 5060 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Basic device measuring group 01 | 5062 | Float | RD | - | Average Value L1 Power Factor
Basic device measuring group 01 | 5064 | Float | RD | - | Average Value L2 Power Factor
Basic device measuring group 01 | 5066 | Float | RD | - | Average Value L3 Power Factor
Basic device measuring group 01 | 5068 | Float | RD | - | Average Value L4 Power Factor
Basic device measuring group 01 | 5070 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Basic device measuring group 01 | 5072 | Float | RD | W | Average Value L1 Active Power (fundamental)
Basic device measuring group 01 | 5074 | Float | RD | W | Average Value L2 Active Power (fundamental)
Basic device measuring group 01 | 5076 | Float | RD | W | Average Value L3 Active Power (fundamental)
Basic device measuring group 01 | 5078 | Float | RD | W | Average Value L4 Active Power (fundamental)
Basic device measuring group 01 | 5080 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 01 | 5082 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Basic device measuring group 01 | 5084 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Basic device measuring group 01 | 5086 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Basic device measuring group 01 | 5088 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Basic device measuring group 01 | 5090 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 01 | 5092 | Float | RD | var | Average Value L1 Distortion Reactive Power
Basic device measuring group 01 | 5094 | Float | RD | var | Average Value L2 Distortion Reactive Power
Basic device measuring group 01 | 5096 | Float | RD | var | Average Value L3 Distortion Reactive Power
Basic device measuring group 01 | 5098 | Float | RD | var | Average Value L4 Distortion Reactive Power
Basic device measuring group 01 | 5100 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 01 | 5102 | Float | RD | % | Average Value L1 THD Current
Basic device measuring group 01 | 5104 | Float | RD | % | Average Value L2 THD Current
Basic device measuring group 01 | 5106 | Float | RD | % | Average Value L3 THD Current
Basic device measuring group 01 | 5108 | Float | RD | % | Average Value L4 THD Current
Basic device measuring group 01 | 5110 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Basic device measuring group 01 | 5112 | Float | RD | % | Average Value L1 TDD Current
Basic device measuring group 01 | 5114 | Float | RD | % | Average Value L2 TDD Current
Basic device measuring group 01 | 5116 | Float | RD | % | Average Value L3 TDD Current
Basic device measuring group 01 | 5118 | Float | RD | % | Average Value L4 TDD Current
Basic device measuring group 01 | 5120 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Basic device measuring group 01 | 5122 | Float | RD | % | Average Value L1 Load Current (%)
Basic device measuring group 01 | 5124 | Float | RD | % | Average Value L2 Load Current (%)
Basic device measuring group 01 | 5126 | Float | RD | % | Average Value L3 Load Current (%)
Basic device measuring group 01 | 5128 | Float | RD | % | Average Value L4 Load Current (%)
Basic device measuring group 01 | 5130 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Basic device measuring group 02 | 5132 | Float | RD | A | Average Value L1 Current (RMS)
Basic device measuring group 02 | 5134 | Float | RD | A | Average Value L2 Current (RMS)
Basic device measuring group 02 | 5136 | Float | RD | A | Average Value L3 Current (RMS)
Basic device measuring group 02 | 5138 | Float | RD | A | Average Value L4 Current (RMS)
Basic device measuring group 02 | 5140 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Basic device measuring group 02 | 5142 | Float | RD | W | Average Value L1 Active Power
Basic device measuring group 02 | 5144 | Float | RD | W | Average Value L2 Active Power
Basic device measuring group 02 | 5146 | Float | RD | W | Average Value L3 Active Power
Basic device measuring group 02 | 5148 | Float | RD | W | Average Value L4 Active Power
Basic device measuring group 02 | 5150 | Float | RD | W | Average Value Sum L1..L4 Active Power
Basic device measuring group 02 | 5152 | Float | RD | var | Average Value L1 Reactive Power
Basic device measuring group 02 | 5154 | Float | RD | var | Average Value L2 Reactive Power
Basic device measuring group 02 | 5156 | Float | RD | var | Average Value L3 Reactive Power
Basic device measuring group 02 | 5158 | Float | RD | var | Average Value L4 Reactive Power
Basic device measuring group 02 | 5160 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Basic device measuring group 02 | 5162 | Float | RD | VA | Average Value L1 Apparent Power
Basic device measuring group 02 | 5164 | Float | RD | VA | Average Value L2 Apparent Power
Basic device measuring group 02 | 5166 | Float | RD | VA | Average Value L3 Apparent Power
Basic device measuring group 02 | 5168 | Float | RD | VA | Average Value L4 Apparent Power
Basic device measuring group 02 | 5170 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Basic device measuring group 02 | 5172 | Float | RD | - | Average Value L1 Cos(phi)
Basic device measuring group 02 | 5174 | Float | RD | - | Average Value L2 Cos(phi)
Basic device measuring group 02 | 5176 | Float | RD | - | Average Value L3 Cos(phi)
Basic device measuring group 02 | 5178 | Float | RD | - | Average Value L4 Cos(phi)
Basic device measuring group 02 | 5180 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Basic device measuring group 02 | 5182 | Float | RD | - | Average Value L1 Power Factor
Basic device measuring group 02 | 5184 | Float | RD | - | Average Value L2 Power Factor
Basic device measuring group 02 | 5186 | Float | RD | - | Average Value L3 Power Factor
Basic device measuring group 02 | 5188 | Float | RD | - | Average Value L4 Power Factor
Basic device measuring group 02 | 5190 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Basic device measuring group 02 | 5192 | Float | RD | W | Average Value L1 Active Power (fundamental)
Basic device measuring group 02 | 5194 | Float | RD | W | Average Value L2 Active Power (fundamental)
Basic device measuring group 02 | 5196 | Float | RD | W | Average Value L3 Active Power (fundamental)
Basic device measuring group 02 | 5198 | Float | RD | W | Average Value L4 Active Power (fundamental)
Basic device measuring group 02 | 5200 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 02 | 5202 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Basic device measuring group 02 | 5204 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Basic device measuring group 02 | 5206 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Basic device measuring group 02 | 5208 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Basic device measuring group 02 | 5210 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 02 | 5212 | Float | RD | var | Average Value L1 Distortion Reactive Power
Basic device measuring group 02 | 5214 | Float | RD | var | Average Value L2 Distortion Reactive Power
Basic device measuring group 02 | 5216 | Float | RD | var | Average Value L3 Distortion Reactive Power
Basic device measuring group 02 | 5218 | Float | RD | var | Average Value L4 Distortion Reactive Power
Basic device measuring group 02 | 5220 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 02 | 5222 | Float | RD | % | Average Value L1 THD Current
Basic device measuring group 02 | 5224 | Float | RD | % | Average Value L2 THD Current
Basic device measuring group 02 | 5226 | Float | RD | % | Average Value L3 THD Current
Basic device measuring group 02 | 5228 | Float | RD | % | Average Value L4 THD Current
Basic device measuring group 02 | 5230 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Basic device measuring group 02 | 5232 | Float | RD | % | Average Value L1 TDD Current
Basic device measuring group 02 | 5234 | Float | RD | % | Average Value L2 TDD Current
Basic device measuring group 02 | 5236 | Float | RD | % | Average Value L3 TDD Current
Basic device measuring group 02 | 5238 | Float | RD | % | Average Value L4 TDD Current
Basic device measuring group 02 | 5240 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Basic device measuring group 02 | 5242 | Float | RD | % | Average Value L1 Load Current (%)
Basic device measuring group 02 | 5244 | Float | RD | % | Average Value L2 Load Current (%)
Basic device measuring group 02 | 5246 | Float | RD | % | Average Value L3 Load Current (%)
Basic device measuring group 02 | 5248 | Float | RD | % | Average Value L4 Load Current (%)
Basic device measuring group 02 | 5250 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 01 | 5252 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 01 | 5254 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 01 | 5256 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 01 | 5258 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 01 | 5260 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 01 | 5262 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 01 | 5264 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 01 | 5266 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 01 | 5268 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 01 | 5270 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 01 | 5272 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 01 | 5274 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 01 | 5276 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 01 | 5278 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 01 | 5280 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 01 | 5282 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 01 | 5284 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 01 | 5286 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 01 | 5288 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 01 | 5290 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 01 | 5292 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 01 | 5294 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 01 | 5296 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 01 | 5298 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 01 | 5300 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 01 | 5302 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 01 | 5304 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 01 | 5306 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 01 | 5308 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 01 | 5310 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 01 | 5312 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 01 | 5314 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 01 | 5316 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 01 | 5318 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 01 | 5320 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 01 | 5322 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 01 | 5324 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 01 | 5326 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 01 | 5328 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 01 | 5330 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 01 | 5332 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 01 | 5334 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 01 | 5336 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 01 | 5338 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 01 | 5340 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 01 | 5342 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 01 | 5344 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 01 | 5346 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 01 | 5348 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 01 | 5350 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 01 | 5352 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 01 | 5354 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 01 | 5356 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 01 | 5358 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 01 | 5360 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 01 | 5362 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 01 | 5364 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 01 | 5366 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 01 | 5368 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 01 | 5370 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 02 | 5372 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 02 | 5374 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 02 | 5376 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 02 | 5378 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 02 | 5380 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 02 | 5382 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 02 | 5384 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 02 | 5386 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 02 | 5388 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 02 | 5390 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 02 | 5392 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 02 | 5394 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 02 | 5396 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 02 | 5398 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 02 | 5400 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 02 | 5402 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 02 | 5404 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 02 | 5406 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 02 | 5408 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 02 | 5410 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 02 | 5412 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 02 | 5414 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 02 | 5416 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 02 | 5418 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 02 | 5420 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 02 | 5422 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 02 | 5424 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 02 | 5426 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 02 | 5428 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 02 | 5430 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 02 | 5432 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 02 | 5434 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 02 | 5436 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 02 | 5438 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 02 | 5440 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 02 | 5442 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 02 | 5444 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 02 | 5446 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 02 | 5448 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 02 | 5450 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 02 | 5452 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 02 | 5454 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 02 | 5456 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 02 | 5458 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 02 | 5460 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 02 | 5462 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 02 | 5464 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 02 | 5466 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 02 | 5468 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 02 | 5470 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 02 | 5472 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 02 | 5474 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 02 | 5476 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 02 | 5478 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 02 | 5480 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 02 | 5482 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 02 | 5484 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 02 | 5486 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 02 | 5488 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 02 | 5490 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 03 | 5492 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 03 | 5494 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 03 | 5496 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 03 | 5498 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 03 | 5500 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 03 | 5502 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 03 | 5504 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 03 | 5506 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 03 | 5508 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 03 | 5510 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 03 | 5512 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 03 | 5514 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 03 | 5516 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 03 | 5518 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 03 | 5520 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 03 | 5522 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 03 | 5524 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 03 | 5526 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 03 | 5528 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 03 | 5530 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 03 | 5532 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 03 | 5534 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 03 | 5536 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 03 | 5538 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 03 | 5540 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 03 | 5542 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 03 | 5544 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 03 | 5546 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 03 | 5548 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 03 | 5550 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 03 | 5552 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 03 | 5554 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 03 | 5556 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 03 | 5558 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 03 | 5560 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 03 | 5562 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 03 | 5564 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 03 | 5566 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 03 | 5568 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 03 | 5570 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 03 | 5572 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 03 | 5574 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 03 | 5576 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 03 | 5578 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 03 | 5580 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 03 | 5582 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 03 | 5584 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 03 | 5586 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 03 | 5588 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 03 | 5590 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 03 | 5592 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 03 | 5594 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 03 | 5596 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 03 | 5598 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 03 | 5600 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 03 | 5602 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 03 | 5604 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 03 | 5606 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 03 | 5608 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 03 | 5610 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 04 | 5612 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 04 | 5614 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 04 | 5616 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 04 | 5618 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 04 | 5620 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 04 | 5622 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 04 | 5624 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 04 | 5626 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 04 | 5628 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 04 | 5630 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 04 | 5632 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 04 | 5634 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 04 | 5636 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 04 | 5638 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 04 | 5640 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 04 | 5642 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 04 | 5644 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 04 | 5646 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 04 | 5648 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 04 | 5650 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 04 | 5652 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 04 | 5654 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 04 | 5656 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 04 | 5658 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 04 | 5660 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 04 | 5662 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 04 | 5664 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 04 | 5666 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 04 | 5668 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 04 | 5670 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 04 | 5672 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 04 | 5674 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 04 | 5676 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 04 | 5678 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 04 | 5680 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 04 | 5682 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 04 | 5684 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 04 | 5686 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 04 | 5688 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 04 | 5690 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 04 | 5692 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 04 | 5694 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 04 | 5696 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 04 | 5698 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 04 | 5700 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 04 | 5702 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 04 | 5704 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 04 | 5706 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 04 | 5708 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 04 | 5710 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 04 | 5712 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 04 | 5714 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 04 | 5716 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 04 | 5718 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 04 | 5720 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 04 | 5722 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 04 | 5724 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 04 | 5726 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 04 | 5728 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 04 | 5730 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 05 | 5732 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 05 | 5734 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 05 | 5736 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 05 | 5738 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 05 | 5740 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 05 | 5742 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 05 | 5744 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 05 | 5746 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 05 | 5748 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 05 | 5750 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 05 | 5752 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 05 | 5754 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 05 | 5756 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 05 | 5758 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 05 | 5760 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 05 | 5762 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 05 | 5764 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 05 | 5766 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 05 | 5768 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 05 | 5770 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 05 | 5772 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 05 | 5774 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 05 | 5776 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 05 | 5778 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 05 | 5780 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 05 | 5782 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 05 | 5784 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 05 | 5786 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 05 | 5788 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 05 | 5790 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 05 | 5792 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 05 | 5794 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 05 | 5796 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 05 | 5798 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 05 | 5800 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 05 | 5802 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 05 | 5804 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 05 | 5806 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 05 | 5808 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 05 | 5810 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 05 | 5812 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 05 | 5814 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 05 | 5816 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 05 | 5818 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 05 | 5820 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 05 | 5822 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 05 | 5824 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 05 | 5826 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 05 | 5828 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 05 | 5830 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 05 | 5832 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 05 | 5834 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 05 | 5836 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 05 | 5838 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 05 | 5840 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 05 | 5842 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 05 | 5844 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 05 | 5846 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 05 | 5848 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 05 | 5850 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 06 | 5852 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 06 | 5854 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 06 | 5856 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 06 | 5858 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 06 | 5860 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 06 | 5862 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 06 | 5864 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 06 | 5866 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 06 | 5868 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 06 | 5870 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 06 | 5872 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 06 | 5874 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 06 | 5876 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 06 | 5878 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 06 | 5880 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 06 | 5882 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 06 | 5884 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 06 | 5886 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 06 | 5888 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 06 | 5890 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 06 | 5892 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 06 | 5894 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 06 | 5896 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 06 | 5898 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 06 | 5900 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 06 | 5902 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 06 | 5904 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 06 | 5906 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 06 | 5908 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 06 | 5910 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 06 | 5912 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 06 | 5914 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 06 | 5916 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 06 | 5918 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 06 | 5920 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 06 | 5922 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 06 | 5924 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 06 | 5926 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 06 | 5928 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 06 | 5930 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 06 | 5932 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 06 | 5934 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 06 | 5936 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 06 | 5938 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 06 | 5940 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 06 | 5942 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 06 | 5944 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 06 | 5946 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 06 | 5948 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 06 | 5950 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 06 | 5952 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 06 | 5954 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 06 | 5956 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 06 | 5958 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 06 | 5960 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 06 | 5962 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 06 | 5964 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 06 | 5966 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 06 | 5968 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 06 | 5970 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 07 | 5972 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 07 | 5974 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 07 | 5976 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 07 | 5978 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 07 | 5980 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 07 | 5982 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 07 | 5984 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 07 | 5986 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 07 | 5988 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 07 | 5990 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 07 | 5992 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 07 | 5994 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 07 | 5996 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 07 | 5998 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 07 | 6000 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 07 | 6002 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 07 | 6004 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 07 | 6006 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 07 | 6008 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 07 | 6010 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 07 | 6012 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 07 | 6014 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 07 | 6016 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 07 | 6018 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 07 | 6020 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 07 | 6022 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 07 | 6024 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 07 | 6026 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 07 | 6028 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 07 | 6030 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 07 | 6032 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 07 | 6034 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 07 | 6036 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 07 | 6038 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 07 | 6040 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 07 | 6042 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 07 | 6044 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 07 | 6046 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 07 | 6048 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 07 | 6050 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 07 | 6052 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 07 | 6054 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 07 | 6056 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 07 | 6058 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 07 | 6060 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 07 | 6062 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 07 | 6064 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 07 | 6066 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 07 | 6068 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 07 | 6070 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 07 | 6072 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 07 | 6074 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 07 | 6076 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 07 | 6078 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 07 | 6080 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 07 | 6082 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 07 | 6084 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 07 | 6086 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 07 | 6088 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 07 | 6090 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 08 | 6092 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 08 | 6094 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 08 | 6096 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 08 | 6098 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 08 | 6100 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 08 | 6102 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 08 | 6104 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 08 | 6106 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 08 | 6108 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 08 | 6110 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 08 | 6112 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 08 | 6114 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 08 | 6116 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 08 | 6118 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 08 | 6120 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 08 | 6122 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 08 | 6124 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 08 | 6126 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 08 | 6128 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 08 | 6130 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 08 | 6132 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 08 | 6134 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 08 | 6136 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 08 | 6138 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 08 | 6140 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 08 | 6142 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 08 | 6144 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 08 | 6146 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 08 | 6148 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 08 | 6150 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 08 | 6152 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 08 | 6154 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 08 | 6156 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 08 | 6158 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 08 | 6160 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 08 | 6162 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 08 | 6164 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 08 | 6166 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 08 | 6168 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 08 | 6170 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 08 | 6172 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 08 | 6174 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 08 | 6176 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 08 | 6178 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 08 | 6180 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 08 | 6182 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 08 | 6184 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 08 | 6186 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 08 | 6188 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 08 | 6190 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 08 | 6192 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 08 | 6194 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 08 | 6196 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 08 | 6198 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 08 | 6200 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 08 | 6202 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 08 | 6204 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 08 | 6206 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 08 | 6208 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 08 | 6210 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 09 | 6212 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 09 | 6214 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 09 | 6216 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 09 | 6218 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 09 | 6220 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 09 | 6222 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 09 | 6224 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 09 | 6226 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 09 | 6228 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 09 | 6230 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 09 | 6232 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 09 | 6234 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 09 | 6236 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 09 | 6238 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 09 | 6240 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 09 | 6242 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 09 | 6244 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 09 | 6246 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 09 | 6248 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 09 | 6250 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 09 | 6252 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 09 | 6254 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 09 | 6256 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 09 | 6258 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 09 | 6260 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 09 | 6262 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 09 | 6264 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 09 | 6266 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 09 | 6268 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 09 | 6270 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 09 | 6272 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 09 | 6274 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 09 | 6276 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 09 | 6278 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 09 | 6280 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 09 | 6282 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 09 | 6284 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 09 | 6286 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 09 | 6288 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 09 | 6290 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 09 | 6292 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 09 | 6294 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 09 | 6296 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 09 | 6298 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 09 | 6300 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 09 | 6302 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 09 | 6304 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 09 | 6306 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 09 | 6308 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 09 | 6310 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 09 | 6312 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 09 | 6314 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 09 | 6316 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 09 | 6318 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 09 | 6320 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 09 | 6322 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 09 | 6324 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 09 | 6326 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 09 | 6328 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 09 | 6330 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 10 | 6332 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 10 | 6334 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 10 | 6336 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 10 | 6338 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 10 | 6340 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 10 | 6342 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 10 | 6344 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 10 | 6346 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 10 | 6348 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 10 | 6350 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 10 | 6352 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 10 | 6354 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 10 | 6356 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 10 | 6358 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 10 | 6360 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 10 | 6362 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 10 | 6364 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 10 | 6366 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 10 | 6368 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 10 | 6370 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 10 | 6372 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 10 | 6374 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 10 | 6376 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 10 | 6378 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 10 | 6380 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 10 | 6382 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 10 | 6384 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 10 | 6386 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 10 | 6388 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 10 | 6390 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 10 | 6392 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 10 | 6394 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 10 | 6396 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 10 | 6398 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 10 | 6400 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 10 | 6402 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 10 | 6404 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 10 | 6406 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 10 | 6408 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 10 | 6410 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 10 | 6412 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 10 | 6414 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 10 | 6416 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 10 | 6418 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 10 | 6420 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 10 | 6422 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 10 | 6424 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 10 | 6426 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 10 | 6428 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 10 | 6430 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 10 | 6432 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 10 | 6434 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 10 | 6436 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 10 | 6438 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 10 | 6440 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 10 | 6442 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 10 | 6444 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 10 | 6446 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 10 | 6448 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 10 | 6450 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 11 | 6452 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 11 | 6454 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 11 | 6456 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 11 | 6458 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 11 | 6460 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 11 | 6462 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 11 | 6464 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 11 | 6466 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 11 | 6468 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 11 | 6470 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 11 | 6472 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 11 | 6474 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 11 | 6476 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 11 | 6478 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 11 | 6480 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 11 | 6482 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 11 | 6484 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 11 | 6486 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 11 | 6488 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 11 | 6490 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 11 | 6492 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 11 | 6494 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 11 | 6496 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 11 | 6498 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 11 | 6500 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 11 | 6502 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 11 | 6504 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 11 | 6506 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 11 | 6508 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 11 | 6510 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 11 | 6512 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 11 | 6514 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 11 | 6516 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 11 | 6518 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 11 | 6520 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 11 | 6522 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 11 | 6524 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 11 | 6526 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 11 | 6528 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 11 | 6530 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 11 | 6532 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 11 | 6534 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 11 | 6536 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 11 | 6538 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 11 | 6540 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 11 | 6542 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 11 | 6544 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 11 | 6546 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 11 | 6548 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 11 | 6550 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 11 | 6552 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 11 | 6554 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 11 | 6556 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 11 | 6558 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 11 | 6560 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 11 | 6562 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 11 | 6564 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 11 | 6566 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 11 | 6568 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 11 | 6570 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 12 | 6572 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 12 | 6574 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 12 | 6576 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 12 | 6578 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 12 | 6580 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 12 | 6582 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 12 | 6584 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 12 | 6586 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 12 | 6588 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 12 | 6590 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 12 | 6592 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 12 | 6594 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 12 | 6596 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 12 | 6598 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 12 | 6600 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 12 | 6602 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 12 | 6604 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 12 | 6606 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 12 | 6608 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 12 | 6610 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 12 | 6612 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 12 | 6614 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 12 | 6616 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 12 | 6618 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 12 | 6620 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 12 | 6622 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 12 | 6624 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 12 | 6626 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 12 | 6628 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 12 | 6630 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 12 | 6632 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 12 | 6634 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 12 | 6636 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 12 | 6638 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 12 | 6640 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 12 | 6642 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 12 | 6644 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 12 | 6646 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 12 | 6648 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 12 | 6650 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 12 | 6652 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 12 | 6654 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 12 | 6656 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 12 | 6658 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 12 | 6660 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 12 | 6662 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 12 | 6664 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 12 | 6666 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 12 | 6668 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 12 | 6670 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 12 | 6672 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 12 | 6674 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 12 | 6676 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 12 | 6678 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 12 | 6680 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 12 | 6682 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 12 | 6684 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 12 | 6686 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 12 | 6688 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 12 | 6690 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 13 | 6692 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 13 | 6694 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 13 | 6696 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 13 | 6698 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 13 | 6700 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 13 | 6702 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 13 | 6704 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 13 | 6706 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 13 | 6708 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 13 | 6710 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 13 | 6712 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 13 | 6714 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 13 | 6716 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 13 | 6718 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 13 | 6720 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 13 | 6722 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 13 | 6724 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 13 | 6726 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 13 | 6728 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 13 | 6730 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 13 | 6732 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 13 | 6734 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 13 | 6736 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 13 | 6738 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 13 | 6740 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 13 | 6742 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 13 | 6744 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 13 | 6746 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 13 | 6748 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 13 | 6750 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 13 | 6752 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 13 | 6754 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 13 | 6756 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 13 | 6758 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 13 | 6760 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 13 | 6762 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 13 | 6764 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 13 | 6766 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 13 | 6768 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 13 | 6770 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 13 | 6772 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 13 | 6774 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 13 | 6776 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 13 | 6778 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 13 | 6780 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 13 | 6782 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 13 | 6784 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 13 | 6786 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 13 | 6788 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 13 | 6790 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 13 | 6792 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 13 | 6794 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 13 | 6796 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 13 | 6798 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 13 | 6800 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 13 | 6802 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 13 | 6804 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 13 | 6806 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 13 | 6808 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 13 | 6810 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 14 | 6812 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 14 | 6814 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 14 | 6816 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 14 | 6818 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 14 | 6820 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 14 | 6822 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 14 | 6824 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 14 | 6826 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 14 | 6828 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 14 | 6830 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 14 | 6832 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 14 | 6834 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 14 | 6836 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 14 | 6838 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 14 | 6840 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 14 | 6842 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 14 | 6844 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 14 | 6846 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 14 | 6848 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 14 | 6850 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 14 | 6852 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 14 | 6854 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 14 | 6856 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 14 | 6858 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 14 | 6860 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 14 | 6862 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 14 | 6864 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 14 | 6866 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 14 | 6868 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 14 | 6870 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 14 | 6872 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 14 | 6874 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 14 | 6876 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 14 | 6878 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 14 | 6880 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 14 | 6882 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 14 | 6884 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 14 | 6886 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 14 | 6888 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 14 | 6890 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 14 | 6892 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 14 | 6894 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 14 | 6896 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 14 | 6898 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 14 | 6900 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 14 | 6902 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 14 | 6904 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 14 | 6906 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 14 | 6908 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 14 | 6910 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 14 | 6912 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 14 | 6914 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 14 | 6916 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 14 | 6918 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 14 | 6920 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 14 | 6922 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 14 | 6924 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 14 | 6926 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 14 | 6928 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 14 | 6930 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 15 | 6932 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 15 | 6934 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 15 | 6936 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 15 | 6938 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 15 | 6940 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 15 | 6942 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 15 | 6944 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 15 | 6946 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 15 | 6948 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 15 | 6950 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 15 | 6952 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 15 | 6954 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 15 | 6956 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 15 | 6958 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 15 | 6960 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 15 | 6962 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 15 | 6964 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 15 | 6966 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 15 | 6968 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 15 | 6970 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 15 | 6972 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 15 | 6974 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 15 | 6976 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 15 | 6978 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 15 | 6980 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 15 | 6982 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 15 | 6984 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 15 | 6986 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 15 | 6988 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 15 | 6990 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 15 | 6992 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 15 | 6994 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 15 | 6996 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 15 | 6998 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 15 | 7000 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 15 | 7002 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 15 | 7004 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 15 | 7006 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 15 | 7008 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 15 | 7010 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 15 | 7012 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 15 | 7014 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 15 | 7016 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 15 | 7018 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 15 | 7020 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 15 | 7022 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 15 | 7024 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 15 | 7026 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 15 | 7028 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 15 | 7030 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 15 | 7032 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 15 | 7034 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 15 | 7036 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 15 | 7038 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 15 | 7040 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 15 | 7042 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 15 | 7044 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 15 | 7046 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 15 | 7048 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 15 | 7050 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 16 | 7052 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 16 | 7054 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 16 | 7056 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 16 | 7058 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 16 | 7060 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 16 | 7062 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 16 | 7064 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 16 | 7066 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 16 | 7068 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 16 | 7070 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 16 | 7072 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 16 | 7074 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 16 | 7076 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 16 | 7078 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 16 | 7080 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 16 | 7082 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 16 | 7084 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 16 | 7086 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 16 | 7088 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 16 | 7090 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 16 | 7092 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 16 | 7094 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 16 | 7096 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 16 | 7098 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 16 | 7100 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 16 | 7102 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 16 | 7104 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 16 | 7106 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 16 | 7108 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 16 | 7110 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 16 | 7112 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 16 | 7114 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 16 | 7116 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 16 | 7118 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 16 | 7120 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 16 | 7122 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 16 | 7124 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 16 | 7126 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 16 | 7128 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 16 | 7130 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 16 | 7132 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 16 | 7134 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 16 | 7136 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 16 | 7138 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 16 | 7140 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 16 | 7142 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 16 | 7144 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 16 | 7146 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 16 | 7148 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 16 | 7150 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 16 | 7152 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 16 | 7154 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 16 | 7156 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 16 | 7158 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 16 | 7160 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 16 | 7162 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 16 | 7164 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 16 | 7166 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 16 | 7168 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 16 | 7170 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 17 | 7172 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 17 | 7174 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 17 | 7176 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 17 | 7178 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 17 | 7180 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 17 | 7182 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 17 | 7184 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 17 | 7186 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 17 | 7188 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 17 | 7190 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 17 | 7192 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 17 | 7194 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 17 | 7196 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 17 | 7198 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 17 | 7200 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 17 | 7202 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 17 | 7204 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 17 | 7206 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 17 | 7208 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 17 | 7210 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 17 | 7212 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 17 | 7214 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 17 | 7216 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 17 | 7218 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 17 | 7220 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 17 | 7222 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 17 | 7224 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 17 | 7226 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 17 | 7228 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 17 | 7230 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 17 | 7232 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 17 | 7234 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 17 | 7236 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 17 | 7238 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 17 | 7240 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 17 | 7242 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 17 | 7244 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 17 | 7246 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 17 | 7248 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 17 | 7250 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 17 | 7252 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 17 | 7254 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 17 | 7256 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 17 | 7258 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 17 | 7260 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 17 | 7262 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 17 | 7264 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 17 | 7266 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 17 | 7268 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 17 | 7270 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 17 | 7272 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 17 | 7274 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 17 | 7276 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 17 | 7278 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 17 | 7280 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 17 | 7282 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 17 | 7284 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 17 | 7286 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 17 | 7288 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 17 | 7290 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 18 | 7292 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 18 | 7294 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 18 | 7296 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 18 | 7298 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 18 | 7300 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 18 | 7302 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 18 | 7304 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 18 | 7306 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 18 | 7308 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 18 | 7310 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 18 | 7312 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 18 | 7314 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 18 | 7316 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 18 | 7318 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 18 | 7320 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 18 | 7322 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 18 | 7324 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 18 | 7326 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 18 | 7328 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 18 | 7330 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 18 | 7332 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 18 | 7334 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 18 | 7336 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 18 | 7338 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 18 | 7340 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 18 | 7342 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 18 | 7344 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 18 | 7346 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 18 | 7348 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 18 | 7350 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 18 | 7352 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 18 | 7354 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 18 | 7356 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 18 | 7358 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 18 | 7360 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 18 | 7362 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 18 | 7364 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 18 | 7366 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 18 | 7368 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 18 | 7370 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 18 | 7372 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 18 | 7374 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 18 | 7376 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 18 | 7378 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 18 | 7380 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 18 | 7382 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 18 | 7384 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 18 | 7386 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 18 | 7388 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 18 | 7390 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 18 | 7392 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 18 | 7394 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 18 | 7396 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 18 | 7398 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 18 | 7400 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 18 | 7402 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 18 | 7404 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 18 | 7406 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 18 | 7408 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 18 | 7410 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 19 | 7412 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 19 | 7414 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 19 | 7416 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 19 | 7418 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 19 | 7420 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 19 | 7422 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 19 | 7424 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 19 | 7426 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 19 | 7428 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 19 | 7430 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 19 | 7432 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 19 | 7434 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 19 | 7436 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 19 | 7438 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 19 | 7440 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 19 | 7442 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 19 | 7444 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 19 | 7446 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 19 | 7448 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 19 | 7450 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 19 | 7452 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 19 | 7454 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 19 | 7456 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 19 | 7458 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 19 | 7460 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 19 | 7462 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 19 | 7464 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 19 | 7466 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 19 | 7468 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 19 | 7470 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 19 | 7472 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 19 | 7474 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 19 | 7476 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 19 | 7478 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 19 | 7480 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 19 | 7482 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 19 | 7484 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 19 | 7486 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 19 | 7488 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 19 | 7490 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 19 | 7492 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 19 | 7494 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 19 | 7496 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 19 | 7498 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 19 | 7500 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 19 | 7502 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 19 | 7504 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 19 | 7506 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 19 | 7508 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 19 | 7510 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 19 | 7512 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 19 | 7514 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 19 | 7516 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 19 | 7518 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 19 | 7520 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 19 | 7522 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 19 | 7524 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 19 | 7526 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 19 | 7528 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 19 | 7530 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 20 | 7532 | Float | RD | A | Average Value L1 Current (RMS)
Current measurement module group 20 | 7534 | Float | RD | A | Average Value L2 Current (RMS)
Current measurement module group 20 | 7536 | Float | RD | A | Average Value L3 Current (RMS)
Current measurement module group 20 | 7538 | Float | RD | A | Average Value L4 Current (RMS)
Current measurement module group 20 | 7540 | Float | RD | A | Average Value Sum L1..L4 Current (RMS)
Current measurement module group 20 | 7542 | Float | RD | W | Average Value L1 Active Power
Current measurement module group 20 | 7544 | Float | RD | W | Average Value L2 Active Power
Current measurement module group 20 | 7546 | Float | RD | W | Average Value L3 Active Power
Current measurement module group 20 | 7548 | Float | RD | W | Average Value L4 Active Power
Current measurement module group 20 | 7550 | Float | RD | W | Average Value Sum L1..L4 Active Power
Current measurement module group 20 | 7552 | Float | RD | var | Average Value L1 Reactive Power
Current measurement module group 20 | 7554 | Float | RD | var | Average Value L2 Reactive Power
Current measurement module group 20 | 7556 | Float | RD | var | Average Value L3 Reactive Power
Current measurement module group 20 | 7558 | Float | RD | var | Average Value L4 Reactive Power
Current measurement module group 20 | 7560 | Float | RD | var | Average Value Sum L1..L4 Reactive Power
Current measurement module group 20 | 7562 | Float | RD | VA | Average Value L1 Apparent Power
Current measurement module group 20 | 7564 | Float | RD | VA | Average Value L2 Apparent Power
Current measurement module group 20 | 7566 | Float | RD | VA | Average Value L3 Apparent Power
Current measurement module group 20 | 7568 | Float | RD | VA | Average Value L4 Apparent Power
Current measurement module group 20 | 7570 | Float | RD | VA | Average Value Sum L1..L4 Apparent Power
Current measurement module group 20 | 7572 | Float | RD | - | Average Value L1 Cos(phi)
Current measurement module group 20 | 7574 | Float | RD | - | Average Value L2 Cos(phi)
Current measurement module group 20 | 7576 | Float | RD | - | Average Value L3 Cos(phi)
Current measurement module group 20 | 7578 | Float | RD | - | Average Value L4 Cos(phi)
Current measurement module group 20 | 7580 | Float | RD | - | Average Value Sum L1..L4 Cos(phi)
Current measurement module group 20 | 7582 | Float | RD | - | Average Value L1 Power Factor
Current measurement module group 20 | 7584 | Float | RD | - | Average Value L2 Power Factor
Current measurement module group 20 | 7586 | Float | RD | - | Average Value L3 Power Factor
Current measurement module group 20 | 7588 | Float | RD | - | Average Value L4 Power Factor
Current measurement module group 20 | 7590 | Float | RD | - | Average Value Sum L1..L4 Power Factor
Current measurement module group 20 | 7592 | Float | RD | W | Average Value L1 Active Power (fundamental)
Current measurement module group 20 | 7594 | Float | RD | W | Average Value L2 Active Power (fundamental)
Current measurement module group 20 | 7596 | Float | RD | W | Average Value L3 Active Power (fundamental)
Current measurement module group 20 | 7598 | Float | RD | W | Average Value L4 Active Power (fundamental)
Current measurement module group 20 | 7600 | Float | RD | W | Average Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 20 | 7602 | Float | RD | var | Average Value L1 Reactive Power (fundamental)
Current measurement module group 20 | 7604 | Float | RD | var | Average Value L2 Reactive Power (fundamental)
Current measurement module group 20 | 7606 | Float | RD | var | Average Value L3 Reactive Power (fundamental)
Current measurement module group 20 | 7608 | Float | RD | var | Average Value L4 Reactive Power (fundamental)
Current measurement module group 20 | 7610 | Float | RD | var | Average Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 20 | 7612 | Float | RD | var | Average Value L1 Distortion Reactive Power
Current measurement module group 20 | 7614 | Float | RD | var | Average Value L2 Distortion Reactive Power
Current measurement module group 20 | 7616 | Float | RD | var | Average Value L3 Distortion Reactive Power
Current measurement module group 20 | 7618 | Float | RD | var | Average Value L4 Distortion Reactive Power
Current measurement module group 20 | 7620 | Float | RD | var | Average Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 20 | 7622 | Float | RD | % | Average Value L1 THD Current
Current measurement module group 20 | 7624 | Float | RD | % | Average Value L2 THD Current
Current measurement module group 20 | 7626 | Float | RD | % | Average Value L3 THD Current
Current measurement module group 20 | 7628 | Float | RD | % | Average Value L4 THD Current
Current measurement module group 20 | 7630 | Float | RD | % | Average Value Sum L1..L4 Sum THD Current
Current measurement module group 20 | 7632 | Float | RD | % | Average Value L1 TDD Current
Current measurement module group 20 | 7634 | Float | RD | % | Average Value L2 TDD Current
Current measurement module group 20 | 7636 | Float | RD | % | Average Value L3 TDD Current
Current measurement module group 20 | 7638 | Float | RD | % | Average Value L4 TDD Current
Current measurement module group 20 | 7640 | Float | RD | % | Average Value Sum L1..L4 Sum TDD Current
Current measurement module group 20 | 7642 | Float | RD | % | Average Value L1 Load Current (%)
Current measurement module group 20 | 7644 | Float | RD | % | Average Value L2 Load Current (%)
Current measurement module group 20 | 7646 | Float | RD | % | Average Value L3 Load Current (%)
Current measurement module group 20 | 7648 | Float | RD | % | Average Value L4 Load Current (%)
Current measurement module group 20 | 7650 | Float | RD | % | Average Value Sum L1..L4 Sum Load Current (%)
basic device | 8000 | Float | RD | V | Minimum Value L1 Voltage L-N (RMS)
basic device | 8002 | Float | RD | V | Minimum Value L2 Voltage L-N (RMS)
basic device | 8004 | Float | RD | V | Minimum Value L3 Voltage L-N (RMS)
basic device | 8006 | Float | RD | V | Minimum Value L1-L2 Voltage L-L (RMS)
basic device | 8008 | Float | RD | V | Minimum Value L2-L3 Voltage L-L (RMS)
basic device | 8010 | Float | RD | V | Minimum Value L3-L1 Voltage L-L (RMS)
Basic device measuring group 01 | 8012 | Float | RD | A | Minimum Value L1 Current (RMS)
Basic device measuring group 01 | 8014 | Float | RD | A | Minimum Value L2 Current (RMS)
Basic device measuring group 01 | 8016 | Float | RD | A | Minimum Value L3 Current (RMS)
Basic device measuring group 01 | 8018 | Float | RD | A | Minimum Value L4 Current (RMS)
Basic device measuring group 01 | 8020 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Basic device measuring group 01 | 8022 | Float | RD | W | Minimum Value L1 Active Power
Basic device measuring group 01 | 8024 | Float | RD | W | Minimum Value L2 Active Power
Basic device measuring group 01 | 8026 | Float | RD | W | Minimum Value L3 Active Power
Basic device measuring group 01 | 8028 | Float | RD | W | Minimum Value L4 Active Power
Basic device measuring group 01 | 8030 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Basic device measuring group 01 | 8032 | Float | RD | var | Minimum Value L1 Reactive Power
Basic device measuring group 01 | 8034 | Float | RD | var | Minimum Value L2 Reactive Power
Basic device measuring group 01 | 8036 | Float | RD | var | Minimum Value L3 Reactive Power
Basic device measuring group 01 | 8038 | Float | RD | var | Minimum Value L4 Reactive Power
Basic device measuring group 01 | 8040 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Basic device measuring group 01 | 8042 | Float | RD | VA | Minimum Value L1 Apparent Power
Basic device measuring group 01 | 8044 | Float | RD | VA | Minimum Value L2 Apparent Power
Basic device measuring group 01 | 8046 | Float | RD | VA | Minimum Value L3 Apparent Power
Basic device measuring group 01 | 8048 | Float | RD | VA | Minimum Value L4 Apparent Power
Basic device measuring group 01 | 8050 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Basic device measuring group 01 | 8052 | Float | RD | - | Minimum Value L1 Cos(phi)
Basic device measuring group 01 | 8054 | Float | RD | - | Minimum Value L2 Cos(phi)
Basic device measuring group 01 | 8056 | Float | RD | - | Minimum Value L3 Cos(phi)
Basic device measuring group 01 | 8058 | Float | RD | - | Minimum Value L4 Cos(phi)
Basic device measuring group 01 | 8060 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Basic device measuring group 01 | 8062 | Float | RD | - | Minimum Value L1 Power Factor
Basic device measuring group 01 | 8064 | Float | RD | - | Minimum Value L2 Power Factor
Basic device measuring group 01 | 8066 | Float | RD | - | Minimum Value L3 Power Factor
Basic device measuring group 01 | 8068 | Float | RD | - | Minimum Value L4 Power Factor
Basic device measuring group 01 | 8070 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Basic device measuring group 01 | 8072 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Basic device measuring group 01 | 8074 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Basic device measuring group 01 | 8076 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Basic device measuring group 01 | 8078 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Basic device measuring group 01 | 8080 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 01 | 8082 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Basic device measuring group 01 | 8084 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Basic device measuring group 01 | 8086 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Basic device measuring group 01 | 8088 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Basic device measuring group 01 | 8090 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 01 | 8092 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Basic device measuring group 01 | 8094 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Basic device measuring group 01 | 8096 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Basic device measuring group 01 | 8098 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Basic device measuring group 01 | 8100 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 01 | 8102 | Float | RD | % | Minimum Value L1 THD Current
Basic device measuring group 01 | 8104 | Float | RD | % | Minimum Value L2 THD Current
Basic device measuring group 01 | 8106 | Float | RD | % | Minimum Value L3 THD Current
Basic device measuring group 01 | 8108 | Float | RD | % | Minimum Value L4 THD Current
Basic device measuring group 01 | 8110 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Basic device measuring group 01 | 8112 | Float | RD | % | Minimum Value L1 TDD Current
Basic device measuring group 01 | 8114 | Float | RD | % | Minimum Value L2 TDD Current
Basic device measuring group 01 | 8116 | Float | RD | % | Minimum Value L3 TDD Current
Basic device measuring group 01 | 8118 | Float | RD | % | Minimum Value L4 TDD Current
Basic device measuring group 01 | 8120 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 01 | 8122 | Float | RD | % | Minimum Value L1 Load Current (%)
Basic device measuring group 01 | 8124 | Float | RD | % | Minimum Value L2 Load Current (%)
Basic device measuring group 01 | 8126 | Float | RD | % | Minimum Value L3 Load Current (%)
Basic device measuring group 01 | 8128 | Float | RD | % | Minimum Value L4 Load Current (%)
Basic device measuring group 01 | 8130 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
basic device | 8132 | DateTime | RD | - | DateTime Minimum Value L1 Voltage L-N (RMS)
basic device | 8134 | DateTime | RD | - | DateTime Minimum Value L2 Voltage L-N (RMS)
basic device | 8136 | DateTime | RD | - | DateTime Minimum Value L3 Voltage L-N (RMS)
basic device | 8138 | DateTime | RD | - | DateTime Minimum Value L1-L2 Voltage L-L (RMS)
basic device | 8140 | DateTime | RD | - | DateTime Minimum Value L2-L3 Voltage L-L (RMS)
basic device | 8142 | DateTime | RD | - | DateTime Minimum Value L3-L1 Voltage L-L (RMS)
Basic device measuring group 01 | 8144 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Basic device measuring group 01 | 8146 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Basic device measuring group 01 | 8148 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Basic device measuring group 01 | 8150 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Basic device measuring group 01 | 8152 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Basic device measuring group 01 | 8154 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Basic device measuring group 01 | 8156 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Basic device measuring group 01 | 8158 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Basic device measuring group 01 | 8160 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Basic device measuring group 01 | 8162 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Basic device measuring group 01 | 8164 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Basic device measuring group 01 | 8166 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Basic device measuring group 01 | 8168 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Basic device measuring group 01 | 8170 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Basic device measuring group 01 | 8172 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Basic device measuring group 01 | 8174 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Basic device measuring group 01 | 8176 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Basic device measuring group 01 | 8178 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Basic device measuring group 01 | 8180 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Basic device measuring group 01 | 8182 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Basic device measuring group 01 | 8184 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Basic device measuring group 01 | 8186 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Basic device measuring group 01 | 8188 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Basic device measuring group 01 | 8190 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Basic device measuring group 01 | 8192 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Basic device measuring group 01 | 8194 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Basic device measuring group 01 | 8196 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Basic device measuring group 01 | 8198 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Basic device measuring group 01 | 8200 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Basic device measuring group 01 | 8202 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Basic device measuring group 01 | 8204 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Basic device measuring group 01 | 8206 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Basic device measuring group 01 | 8208 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Basic device measuring group 01 | 8210 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Basic device measuring group 01 | 8212 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 01 | 8214 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Basic device measuring group 01 | 8216 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Basic device measuring group 01 | 8218 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Basic device measuring group 01 | 8220 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Basic device measuring group 01 | 8222 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 01 | 8224 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Basic device measuring group 01 | 8226 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Basic device measuring group 01 | 8228 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Basic device measuring group 01 | 8230 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Basic device measuring group 01 | 8232 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 01 | 8234 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Basic device measuring group 01 | 8236 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Basic device measuring group 01 | 8238 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Basic device measuring group 01 | 8240 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Basic device measuring group 01 | 8242 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Basic device measuring group 01 | 8244 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Basic device measuring group 01 | 8246 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Basic device measuring group 01 | 8248 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Basic device measuring group 01 | 8250 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Basic device measuring group 01 | 8252 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 01 | 8254 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Basic device measuring group 01 | 8256 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Basic device measuring group 01 | 8258 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Basic device measuring group 01 | 8260 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Basic device measuring group 01 | 8262 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Basic device measuring group 02 | 8264 | Float | RD | A | Minimum Value L1 Current (RMS)
Basic device measuring group 02 | 8266 | Float | RD | A | Minimum Value L2 Current (RMS)
Basic device measuring group 02 | 8268 | Float | RD | A | Minimum Value L3 Current (RMS)
Basic device measuring group 02 | 8270 | Float | RD | A | Minimum Value L4 Current (RMS)
Basic device measuring group 02 | 8272 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Basic device measuring group 02 | 8274 | Float | RD | W | Minimum Value L1 Active Power
Basic device measuring group 02 | 8276 | Float | RD | W | Minimum Value L2 Active Power
Basic device measuring group 02 | 8278 | Float | RD | W | Minimum Value L3 Active Power
Basic device measuring group 02 | 8280 | Float | RD | W | Minimum Value L4 Active Power
Basic device measuring group 02 | 8282 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Basic device measuring group 02 | 8284 | Float | RD | var | Minimum Value L1 Reactive Power
Basic device measuring group 02 | 8286 | Float | RD | var | Minimum Value L2 Reactive Power
Basic device measuring group 02 | 8288 | Float | RD | var | Minimum Value L3 Reactive Power
Basic device measuring group 02 | 8290 | Float | RD | var | Minimum Value L4 Reactive Power
Basic device measuring group 02 | 8292 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Basic device measuring group 02 | 8294 | Float | RD | VA | Minimum Value L1 Apparent Power
Basic device measuring group 02 | 8296 | Float | RD | VA | Minimum Value L2 Apparent Power
Basic device measuring group 02 | 8298 | Float | RD | VA | Minimum Value L3 Apparent Power
Basic device measuring group 02 | 8300 | Float | RD | VA | Minimum Value L4 Apparent Power
Basic device measuring group 02 | 8302 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Basic device measuring group 02 | 8304 | Float | RD | - | Minimum Value L1 Cos(phi)
Basic device measuring group 02 | 8306 | Float | RD | - | Minimum Value L2 Cos(phi)
Basic device measuring group 02 | 8308 | Float | RD | - | Minimum Value L3 Cos(phi)
Basic device measuring group 02 | 8310 | Float | RD | - | Minimum Value L4 Cos(phi)
Basic device measuring group 02 | 8312 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Basic device measuring group 02 | 8314 | Float | RD | - | Minimum Value L1 Power Factor
Basic device measuring group 02 | 8316 | Float | RD | - | Minimum Value L2 Power Factor
Basic device measuring group 02 | 8318 | Float | RD | - | Minimum Value L3 Power Factor
Basic device measuring group 02 | 8320 | Float | RD | - | Minimum Value L4 Power Factor
Basic device measuring group 02 | 8322 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Basic device measuring group 02 | 8324 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Basic device measuring group 02 | 8326 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Basic device measuring group 02 | 8328 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Basic device measuring group 02 | 8330 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Basic device measuring group 02 | 8332 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 02 | 8334 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Basic device measuring group 02 | 8336 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Basic device measuring group 02 | 8338 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Basic device measuring group 02 | 8340 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Basic device measuring group 02 | 8342 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 02 | 8344 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Basic device measuring group 02 | 8346 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Basic device measuring group 02 | 8348 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Basic device measuring group 02 | 8350 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Basic device measuring group 02 | 8352 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 02 | 8354 | Float | RD | % | Minimum Value L1 THD Current
Basic device measuring group 02 | 8356 | Float | RD | % | Minimum Value L2 THD Current
Basic device measuring group 02 | 8358 | Float | RD | % | Minimum Value L3 THD Current
Basic device measuring group 02 | 8360 | Float | RD | % | Minimum Value L4 THD Current
Basic device measuring group 02 | 8362 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Basic device measuring group 02 | 8364 | Float | RD | % | Minimum Value L1 TDD Current
Basic device measuring group 02 | 8366 | Float | RD | % | Minimum Value L2 TDD Current
Basic device measuring group 02 | 8368 | Float | RD | % | Minimum Value L3 TDD Current
Basic device measuring group 02 | 8370 | Float | RD | % | Minimum Value L4 TDD Current
Basic device measuring group 02 | 8372 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 02 | 8374 | Float | RD | % | Minimum Value L1 Load Current (%)
Basic device measuring group 02 | 8376 | Float | RD | % | Minimum Value L2 Load Current (%)
Basic device measuring group 02 | 8378 | Float | RD | % | Minimum Value L3 Load Current (%)
Basic device measuring group 02 | 8380 | Float | RD | % | Minimum Value L4 Load Current (%)
Basic device measuring group 02 | 8382 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Basic device measuring group 02 | 8384 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Basic device measuring group 02 | 8386 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Basic device measuring group 02 | 8388 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Basic device measuring group 02 | 8390 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Basic device measuring group 02 | 8392 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Basic device measuring group 02 | 8394 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Basic device measuring group 02 | 8396 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Basic device measuring group 02 | 8398 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Basic device measuring group 02 | 8400 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Basic device measuring group 02 | 8402 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Basic device measuring group 02 | 8404 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Basic device measuring group 02 | 8406 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Basic device measuring group 02 | 8408 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Basic device measuring group 02 | 8410 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Basic device measuring group 02 | 8412 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Basic device measuring group 02 | 8414 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Basic device measuring group 02 | 8416 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Basic device measuring group 02 | 8418 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Basic device measuring group 02 | 8420 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Basic device measuring group 02 | 8422 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Basic device measuring group 02 | 8424 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Basic device measuring group 02 | 8426 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Basic device measuring group 02 | 8428 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Basic device measuring group 02 | 8430 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Basic device measuring group 02 | 8432 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Basic device measuring group 02 | 8434 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Basic device measuring group 02 | 8436 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Basic device measuring group 02 | 8438 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Basic device measuring group 02 | 8440 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Basic device measuring group 02 | 8442 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Basic device measuring group 02 | 8444 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Basic device measuring group 02 | 8446 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Basic device measuring group 02 | 8448 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Basic device measuring group 02 | 8450 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Basic device measuring group 02 | 8452 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 02 | 8454 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Basic device measuring group 02 | 8456 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Basic device measuring group 02 | 8458 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Basic device measuring group 02 | 8460 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Basic device measuring group 02 | 8462 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 02 | 8464 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Basic device measuring group 02 | 8466 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Basic device measuring group 02 | 8468 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Basic device measuring group 02 | 8470 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Basic device measuring group 02 | 8472 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 02 | 8474 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Basic device measuring group 02 | 8476 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Basic device measuring group 02 | 8478 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Basic device measuring group 02 | 8480 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Basic device measuring group 02 | 8482 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Basic device measuring group 02 | 8484 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Basic device measuring group 02 | 8486 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Basic device measuring group 02 | 8488 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Basic device measuring group 02 | 8490 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Basic device measuring group 02 | 8492 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 02 | 8494 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Basic device measuring group 02 | 8496 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Basic device measuring group 02 | 8498 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Basic device measuring group 02 | 8500 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Basic device measuring group 02 | 8502 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 01 | 8504 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 01 | 8506 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 01 | 8508 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 01 | 8510 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 01 | 8512 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 01 | 8514 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 01 | 8516 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 01 | 8518 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 01 | 8520 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 01 | 8522 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 01 | 8524 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 01 | 8526 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 01 | 8528 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 01 | 8530 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 01 | 8532 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 01 | 8534 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 01 | 8536 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 01 | 8538 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 01 | 8540 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 01 | 8542 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 01 | 8544 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 01 | 8546 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 01 | 8548 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 01 | 8550 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 01 | 8552 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 01 | 8554 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 01 | 8556 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 01 | 8558 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 01 | 8560 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 01 | 8562 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 01 | 8564 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 01 | 8566 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 01 | 8568 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 01 | 8570 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 01 | 8572 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 01 | 8574 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 01 | 8576 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 01 | 8578 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 01 | 8580 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 01 | 8582 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 01 | 8584 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 01 | 8586 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 01 | 8588 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 01 | 8590 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 01 | 8592 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 01 | 8594 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 01 | 8596 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 01 | 8598 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 01 | 8600 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 01 | 8602 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 01 | 8604 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 01 | 8606 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 01 | 8608 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 01 | 8610 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 01 | 8612 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 01 | 8614 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 01 | 8616 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 01 | 8618 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 01 | 8620 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 01 | 8622 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 01 | 8624 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 01 | 8626 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 01 | 8628 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 01 | 8630 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 01 | 8632 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 01 | 8634 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 01 | 8636 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 01 | 8638 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 01 | 8640 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 01 | 8642 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 01 | 8644 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 01 | 8646 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 01 | 8648 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 01 | 8650 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 01 | 8652 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 01 | 8654 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 01 | 8656 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 01 | 8658 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 01 | 8660 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 01 | 8662 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 01 | 8664 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 01 | 8666 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 01 | 8668 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 01 | 8670 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 01 | 8672 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 01 | 8674 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 01 | 8676 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 01 | 8678 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 01 | 8680 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 01 | 8682 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 01 | 8684 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 01 | 8686 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 01 | 8688 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 01 | 8690 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 01 | 8692 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 01 | 8694 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 01 | 8696 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 01 | 8698 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 01 | 8700 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 01 | 8702 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 01 | 8704 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 01 | 8706 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 01 | 8708 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 01 | 8710 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 01 | 8712 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 01 | 8714 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 01 | 8716 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 01 | 8718 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 01 | 8720 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 01 | 8722 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 01 | 8724 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 01 | 8726 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 01 | 8728 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 01 | 8730 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 01 | 8732 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 01 | 8734 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 01 | 8736 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 01 | 8738 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 01 | 8740 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 01 | 8742 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 02 | 8744 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 02 | 8746 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 02 | 8748 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 02 | 8750 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 02 | 8752 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 02 | 8754 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 02 | 8756 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 02 | 8758 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 02 | 8760 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 02 | 8762 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 02 | 8764 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 02 | 8766 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 02 | 8768 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 02 | 8770 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 02 | 8772 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 02 | 8774 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 02 | 8776 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 02 | 8778 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 02 | 8780 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 02 | 8782 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 02 | 8784 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 02 | 8786 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 02 | 8788 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 02 | 8790 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 02 | 8792 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 02 | 8794 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 02 | 8796 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 02 | 8798 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 02 | 8800 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 02 | 8802 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 02 | 8804 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 02 | 8806 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 02 | 8808 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 02 | 8810 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 02 | 8812 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 02 | 8814 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 02 | 8816 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 02 | 8818 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 02 | 8820 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 02 | 8822 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 02 | 8824 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 02 | 8826 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 02 | 8828 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 02 | 8830 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 02 | 8832 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 02 | 8834 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 02 | 8836 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 02 | 8838 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 02 | 8840 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 02 | 8842 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 02 | 8844 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 02 | 8846 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 02 | 8848 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 02 | 8850 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 02 | 8852 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 02 | 8854 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 02 | 8856 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 02 | 8858 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 02 | 8860 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 02 | 8862 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 02 | 8864 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 02 | 8866 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 02 | 8868 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 02 | 8870 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 02 | 8872 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 02 | 8874 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 02 | 8876 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 02 | 8878 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 02 | 8880 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 02 | 8882 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 02 | 8884 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 02 | 8886 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 02 | 8888 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 02 | 8890 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 02 | 8892 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 02 | 8894 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 02 | 8896 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 02 | 8898 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 02 | 8900 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 02 | 8902 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 02 | 8904 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 02 | 8906 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 02 | 8908 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 02 | 8910 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 02 | 8912 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 02 | 8914 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 02 | 8916 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 02 | 8918 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 02 | 8920 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 02 | 8922 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 02 | 8924 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 02 | 8926 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 02 | 8928 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 02 | 8930 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 02 | 8932 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 02 | 8934 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 02 | 8936 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 02 | 8938 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 02 | 8940 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 02 | 8942 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 02 | 8944 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 02 | 8946 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 02 | 8948 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 02 | 8950 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 02 | 8952 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 02 | 8954 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 02 | 8956 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 02 | 8958 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 02 | 8960 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 02 | 8962 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 02 | 8964 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 02 | 8966 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 02 | 8968 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 02 | 8970 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 02 | 8972 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 02 | 8974 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 02 | 8976 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 02 | 8978 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 02 | 8980 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 02 | 8982 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 03 | 8984 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 03 | 8986 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 03 | 8988 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 03 | 8990 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 03 | 8992 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 03 | 8994 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 03 | 8996 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 03 | 8998 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 03 | 9000 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 03 | 9002 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 03 | 9004 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 03 | 9006 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 03 | 9008 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 03 | 9010 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 03 | 9012 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 03 | 9014 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 03 | 9016 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 03 | 9018 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 03 | 9020 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 03 | 9022 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 03 | 9024 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 03 | 9026 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 03 | 9028 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 03 | 9030 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 03 | 9032 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 03 | 9034 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 03 | 9036 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 03 | 9038 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 03 | 9040 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 03 | 9042 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 03 | 9044 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 03 | 9046 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 03 | 9048 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 03 | 9050 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 03 | 9052 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 03 | 9054 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 03 | 9056 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 03 | 9058 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 03 | 9060 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 03 | 9062 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 03 | 9064 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 03 | 9066 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 03 | 9068 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 03 | 9070 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 03 | 9072 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 03 | 9074 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 03 | 9076 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 03 | 9078 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 03 | 9080 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 03 | 9082 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 03 | 9084 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 03 | 9086 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 03 | 9088 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 03 | 9090 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 03 | 9092 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 03 | 9094 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 03 | 9096 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 03 | 9098 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 03 | 9100 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 03 | 9102 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 03 | 9104 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 03 | 9106 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 03 | 9108 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 03 | 9110 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 03 | 9112 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 03 | 9114 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 03 | 9116 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 03 | 9118 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 03 | 9120 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 03 | 9122 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 03 | 9124 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 03 | 9126 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 03 | 9128 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 03 | 9130 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 03 | 9132 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 03 | 9134 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 03 | 9136 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 03 | 9138 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 03 | 9140 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 03 | 9142 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 03 | 9144 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 03 | 9146 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 03 | 9148 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 03 | 9150 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 03 | 9152 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 03 | 9154 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 03 | 9156 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 03 | 9158 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 03 | 9160 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 03 | 9162 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 03 | 9164 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 03 | 9166 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 03 | 9168 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 03 | 9170 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 03 | 9172 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 03 | 9174 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 03 | 9176 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 03 | 9178 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 03 | 9180 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 03 | 9182 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 03 | 9184 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 03 | 9186 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 03 | 9188 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 03 | 9190 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 03 | 9192 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 03 | 9194 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 03 | 9196 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 03 | 9198 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 03 | 9200 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 03 | 9202 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 03 | 9204 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 03 | 9206 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 03 | 9208 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 03 | 9210 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 03 | 9212 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 03 | 9214 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 03 | 9216 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 03 | 9218 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 03 | 9220 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 03 | 9222 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 04 | 9224 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 04 | 9226 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 04 | 9228 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 04 | 9230 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 04 | 9232 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 04 | 9234 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 04 | 9236 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 04 | 9238 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 04 | 9240 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 04 | 9242 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 04 | 9244 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 04 | 9246 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 04 | 9248 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 04 | 9250 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 04 | 9252 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 04 | 9254 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 04 | 9256 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 04 | 9258 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 04 | 9260 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 04 | 9262 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 04 | 9264 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 04 | 9266 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 04 | 9268 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 04 | 9270 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 04 | 9272 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 04 | 9274 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 04 | 9276 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 04 | 9278 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 04 | 9280 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 04 | 9282 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 04 | 9284 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 04 | 9286 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 04 | 9288 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 04 | 9290 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 04 | 9292 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 04 | 9294 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 04 | 9296 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 04 | 9298 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 04 | 9300 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 04 | 9302 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 04 | 9304 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 04 | 9306 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 04 | 9308 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 04 | 9310 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 04 | 9312 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 04 | 9314 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 04 | 9316 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 04 | 9318 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 04 | 9320 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 04 | 9322 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 04 | 9324 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 04 | 9326 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 04 | 9328 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 04 | 9330 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 04 | 9332 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 04 | 9334 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 04 | 9336 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 04 | 9338 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 04 | 9340 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 04 | 9342 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 04 | 9344 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 04 | 9346 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 04 | 9348 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 04 | 9350 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 04 | 9352 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 04 | 9354 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 04 | 9356 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 04 | 9358 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 04 | 9360 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 04 | 9362 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 04 | 9364 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 04 | 9366 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 04 | 9368 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 04 | 9370 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 04 | 9372 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 04 | 9374 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 04 | 9376 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 04 | 9378 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 04 | 9380 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 04 | 9382 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 04 | 9384 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 04 | 9386 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 04 | 9388 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 04 | 9390 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 04 | 9392 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 04 | 9394 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 04 | 9396 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 04 | 9398 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 04 | 9400 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 04 | 9402 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 04 | 9404 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 04 | 9406 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 04 | 9408 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 04 | 9410 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 04 | 9412 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 04 | 9414 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 04 | 9416 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 04 | 9418 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 04 | 9420 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 04 | 9422 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 04 | 9424 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 04 | 9426 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 04 | 9428 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 04 | 9430 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 04 | 9432 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 04 | 9434 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 04 | 9436 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 04 | 9438 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 04 | 9440 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 04 | 9442 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 04 | 9444 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 04 | 9446 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 04 | 9448 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 04 | 9450 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 04 | 9452 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 04 | 9454 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 04 | 9456 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 04 | 9458 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 04 | 9460 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 04 | 9462 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 05 | 9464 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 05 | 9466 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 05 | 9468 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 05 | 9470 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 05 | 9472 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 05 | 9474 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 05 | 9476 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 05 | 9478 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 05 | 9480 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 05 | 9482 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 05 | 9484 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 05 | 9486 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 05 | 9488 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 05 | 9490 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 05 | 9492 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 05 | 9494 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 05 | 9496 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 05 | 9498 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 05 | 9500 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 05 | 9502 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 05 | 9504 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 05 | 9506 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 05 | 9508 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 05 | 9510 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 05 | 9512 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 05 | 9514 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 05 | 9516 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 05 | 9518 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 05 | 9520 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 05 | 9522 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 05 | 9524 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 05 | 9526 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 05 | 9528 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 05 | 9530 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 05 | 9532 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 05 | 9534 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 05 | 9536 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 05 | 9538 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 05 | 9540 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 05 | 9542 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 05 | 9544 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 05 | 9546 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 05 | 9548 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 05 | 9550 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 05 | 9552 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 05 | 9554 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 05 | 9556 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 05 | 9558 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 05 | 9560 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 05 | 9562 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 05 | 9564 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 05 | 9566 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 05 | 9568 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 05 | 9570 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 05 | 9572 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 05 | 9574 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 05 | 9576 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 05 | 9578 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 05 | 9580 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 05 | 9582 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 05 | 9584 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 05 | 9586 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 05 | 9588 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 05 | 9590 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 05 | 9592 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 05 | 9594 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 05 | 9596 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 05 | 9598 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 05 | 9600 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 05 | 9602 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 05 | 9604 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 05 | 9606 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 05 | 9608 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 05 | 9610 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 05 | 9612 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 05 | 9614 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 05 | 9616 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 05 | 9618 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 05 | 9620 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 05 | 9622 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 05 | 9624 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 05 | 9626 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 05 | 9628 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 05 | 9630 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 05 | 9632 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 05 | 9634 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 05 | 9636 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 05 | 9638 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 05 | 9640 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 05 | 9642 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 05 | 9644 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 05 | 9646 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 05 | 9648 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 05 | 9650 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 05 | 9652 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 05 | 9654 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 05 | 9656 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 05 | 9658 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 05 | 9660 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 05 | 9662 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 05 | 9664 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 05 | 9666 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 05 | 9668 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 05 | 9670 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 05 | 9672 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 05 | 9674 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 05 | 9676 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 05 | 9678 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 05 | 9680 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 05 | 9682 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 05 | 9684 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 05 | 9686 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 05 | 9688 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 05 | 9690 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 05 | 9692 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 05 | 9694 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 05 | 9696 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 05 | 9698 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 05 | 9700 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 05 | 9702 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 06 | 9704 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 06 | 9706 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 06 | 9708 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 06 | 9710 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 06 | 9712 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 06 | 9714 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 06 | 9716 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 06 | 9718 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 06 | 9720 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 06 | 9722 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 06 | 9724 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 06 | 9726 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 06 | 9728 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 06 | 9730 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 06 | 9732 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 06 | 9734 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 06 | 9736 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 06 | 9738 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 06 | 9740 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 06 | 9742 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 06 | 9744 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 06 | 9746 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 06 | 9748 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 06 | 9750 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 06 | 9752 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 06 | 9754 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 06 | 9756 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 06 | 9758 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 06 | 9760 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 06 | 9762 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 06 | 9764 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 06 | 9766 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 06 | 9768 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 06 | 9770 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 06 | 9772 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 06 | 9774 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 06 | 9776 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 06 | 9778 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 06 | 9780 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 06 | 9782 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 06 | 9784 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 06 | 9786 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 06 | 9788 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 06 | 9790 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 06 | 9792 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 06 | 9794 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 06 | 9796 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 06 | 9798 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 06 | 9800 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 06 | 9802 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 06 | 9804 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 06 | 9806 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 06 | 9808 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 06 | 9810 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 06 | 9812 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 06 | 9814 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 06 | 9816 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 06 | 9818 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 06 | 9820 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 06 | 9822 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 06 | 9824 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 06 | 9826 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 06 | 9828 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 06 | 9830 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 06 | 9832 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 06 | 9834 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 06 | 9836 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 06 | 9838 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 06 | 9840 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 06 | 9842 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 06 | 9844 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 06 | 9846 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 06 | 9848 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 06 | 9850 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 06 | 9852 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 06 | 9854 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 06 | 9856 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 06 | 9858 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 06 | 9860 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 06 | 9862 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 06 | 9864 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 06 | 9866 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 06 | 9868 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 06 | 9870 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 06 | 9872 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 06 | 9874 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 06 | 9876 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 06 | 9878 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 06 | 9880 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 06 | 9882 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 06 | 9884 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 06 | 9886 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 06 | 9888 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 06 | 9890 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 06 | 9892 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 06 | 9894 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 06 | 9896 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 06 | 9898 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 06 | 9900 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 06 | 9902 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 06 | 9904 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 06 | 9906 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 06 | 9908 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 06 | 9910 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 06 | 9912 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 06 | 9914 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 06 | 9916 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 06 | 9918 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 06 | 9920 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 06 | 9922 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 06 | 9924 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 06 | 9926 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 06 | 9928 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 06 | 9930 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 06 | 9932 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 06 | 9934 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 06 | 9936 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 06 | 9938 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 06 | 9940 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 06 | 9942 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 07 | 9944 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 07 | 9946 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 07 | 9948 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 07 | 9950 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 07 | 9952 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 07 | 9954 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 07 | 9956 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 07 | 9958 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 07 | 9960 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 07 | 9962 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 07 | 9964 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 07 | 9966 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 07 | 9968 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 07 | 9970 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 07 | 9972 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 07 | 9974 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 07 | 9976 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 07 | 9978 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 07 | 9980 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 07 | 9982 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 07 | 9984 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 07 | 9986 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 07 | 9988 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 07 | 9990 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 07 | 9992 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 07 | 9994 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 07 | 9996 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 07 | 9998 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 07 | 10000 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 07 | 10002 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 07 | 10004 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 07 | 10006 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 07 | 10008 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 07 | 10010 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 07 | 10012 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 07 | 10014 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 07 | 10016 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 07 | 10018 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 07 | 10020 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 07 | 10022 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 07 | 10024 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 07 | 10026 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 07 | 10028 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 07 | 10030 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 07 | 10032 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 07 | 10034 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 07 | 10036 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 07 | 10038 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 07 | 10040 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 07 | 10042 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 07 | 10044 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 07 | 10046 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 07 | 10048 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 07 | 10050 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 07 | 10052 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 07 | 10054 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 07 | 10056 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 07 | 10058 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 07 | 10060 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 07 | 10062 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 07 | 10064 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 07 | 10066 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 07 | 10068 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 07 | 10070 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 07 | 10072 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 07 | 10074 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 07 | 10076 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 07 | 10078 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 07 | 10080 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 07 | 10082 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 07 | 10084 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 07 | 10086 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 07 | 10088 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 07 | 10090 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 07 | 10092 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 07 | 10094 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 07 | 10096 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 07 | 10098 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 07 | 10100 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 07 | 10102 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 07 | 10104 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 07 | 10106 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 07 | 10108 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 07 | 10110 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 07 | 10112 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 07 | 10114 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 07 | 10116 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 07 | 10118 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 07 | 10120 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 07 | 10122 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 07 | 10124 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 07 | 10126 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 07 | 10128 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 07 | 10130 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 07 | 10132 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 07 | 10134 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 07 | 10136 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 07 | 10138 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 07 | 10140 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 07 | 10142 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 07 | 10144 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 07 | 10146 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 07 | 10148 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 07 | 10150 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 07 | 10152 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 07 | 10154 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 07 | 10156 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 07 | 10158 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 07 | 10160 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 07 | 10162 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 07 | 10164 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 07 | 10166 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 07 | 10168 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 07 | 10170 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 07 | 10172 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 07 | 10174 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 07 | 10176 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 07 | 10178 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 07 | 10180 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 07 | 10182 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 08 | 10184 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 08 | 10186 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 08 | 10188 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 08 | 10190 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 08 | 10192 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 08 | 10194 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 08 | 10196 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 08 | 10198 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 08 | 10200 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 08 | 10202 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 08 | 10204 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 08 | 10206 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 08 | 10208 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 08 | 10210 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 08 | 10212 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 08 | 10214 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 08 | 10216 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 08 | 10218 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 08 | 10220 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 08 | 10222 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 08 | 10224 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 08 | 10226 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 08 | 10228 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 08 | 10230 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 08 | 10232 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 08 | 10234 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 08 | 10236 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 08 | 10238 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 08 | 10240 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 08 | 10242 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 08 | 10244 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 08 | 10246 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 08 | 10248 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 08 | 10250 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 08 | 10252 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 08 | 10254 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 08 | 10256 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 08 | 10258 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 08 | 10260 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 08 | 10262 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 08 | 10264 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 08 | 10266 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 08 | 10268 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 08 | 10270 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 08 | 10272 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 08 | 10274 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 08 | 10276 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 08 | 10278 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 08 | 10280 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 08 | 10282 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 08 | 10284 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 08 | 10286 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 08 | 10288 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 08 | 10290 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 08 | 10292 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 08 | 10294 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 08 | 10296 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 08 | 10298 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 08 | 10300 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 08 | 10302 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 08 | 10304 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 08 | 10306 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 08 | 10308 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 08 | 10310 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 08 | 10312 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 08 | 10314 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 08 | 10316 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 08 | 10318 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 08 | 10320 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 08 | 10322 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 08 | 10324 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 08 | 10326 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 08 | 10328 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 08 | 10330 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 08 | 10332 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 08 | 10334 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 08 | 10336 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 08 | 10338 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 08 | 10340 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 08 | 10342 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 08 | 10344 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 08 | 10346 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 08 | 10348 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 08 | 10350 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 08 | 10352 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 08 | 10354 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 08 | 10356 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 08 | 10358 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 08 | 10360 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 08 | 10362 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 08 | 10364 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 08 | 10366 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 08 | 10368 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 08 | 10370 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 08 | 10372 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 08 | 10374 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 08 | 10376 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 08 | 10378 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 08 | 10380 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 08 | 10382 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 08 | 10384 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 08 | 10386 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 08 | 10388 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 08 | 10390 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 08 | 10392 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 08 | 10394 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 08 | 10396 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 08 | 10398 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 08 | 10400 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 08 | 10402 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 08 | 10404 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 08 | 10406 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 08 | 10408 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 08 | 10410 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 08 | 10412 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 08 | 10414 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 08 | 10416 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 08 | 10418 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 08 | 10420 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 08 | 10422 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 09 | 10424 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 09 | 10426 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 09 | 10428 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 09 | 10430 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 09 | 10432 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 09 | 10434 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 09 | 10436 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 09 | 10438 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 09 | 10440 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 09 | 10442 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 09 | 10444 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 09 | 10446 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 09 | 10448 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 09 | 10450 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 09 | 10452 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 09 | 10454 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 09 | 10456 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 09 | 10458 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 09 | 10460 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 09 | 10462 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 09 | 10464 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 09 | 10466 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 09 | 10468 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 09 | 10470 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 09 | 10472 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 09 | 10474 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 09 | 10476 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 09 | 10478 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 09 | 10480 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 09 | 10482 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 09 | 10484 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 09 | 10486 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 09 | 10488 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 09 | 10490 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 09 | 10492 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 09 | 10494 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 09 | 10496 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 09 | 10498 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 09 | 10500 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 09 | 10502 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 09 | 10504 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 09 | 10506 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 09 | 10508 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 09 | 10510 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 09 | 10512 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 09 | 10514 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 09 | 10516 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 09 | 10518 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 09 | 10520 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 09 | 10522 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 09 | 10524 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 09 | 10526 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 09 | 10528 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 09 | 10530 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 09 | 10532 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 09 | 10534 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 09 | 10536 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 09 | 10538 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 09 | 10540 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 09 | 10542 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 09 | 10544 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 09 | 10546 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 09 | 10548 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 09 | 10550 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 09 | 10552 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 09 | 10554 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 09 | 10556 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 09 | 10558 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 09 | 10560 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 09 | 10562 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 09 | 10564 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 09 | 10566 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 09 | 10568 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 09 | 10570 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 09 | 10572 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 09 | 10574 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 09 | 10576 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 09 | 10578 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 09 | 10580 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 09 | 10582 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 09 | 10584 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 09 | 10586 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 09 | 10588 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 09 | 10590 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 09 | 10592 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 09 | 10594 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 09 | 10596 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 09 | 10598 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 09 | 10600 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 09 | 10602 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 09 | 10604 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 09 | 10606 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 09 | 10608 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 09 | 10610 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 09 | 10612 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 09 | 10614 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 09 | 10616 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 09 | 10618 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 09 | 10620 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 09 | 10622 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 09 | 10624 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 09 | 10626 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 09 | 10628 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 09 | 10630 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 09 | 10632 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 09 | 10634 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 09 | 10636 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 09 | 10638 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 09 | 10640 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 09 | 10642 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 09 | 10644 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 09 | 10646 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 09 | 10648 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 09 | 10650 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 09 | 10652 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 09 | 10654 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 09 | 10656 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 09 | 10658 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 09 | 10660 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 09 | 10662 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 10 | 10664 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 10 | 10666 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 10 | 10668 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 10 | 10670 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 10 | 10672 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 10 | 10674 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 10 | 10676 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 10 | 10678 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 10 | 10680 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 10 | 10682 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 10 | 10684 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 10 | 10686 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 10 | 10688 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 10 | 10690 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 10 | 10692 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 10 | 10694 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 10 | 10696 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 10 | 10698 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 10 | 10700 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 10 | 10702 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 10 | 10704 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 10 | 10706 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 10 | 10708 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 10 | 10710 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 10 | 10712 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 10 | 10714 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 10 | 10716 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 10 | 10718 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 10 | 10720 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 10 | 10722 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 10 | 10724 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 10 | 10726 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 10 | 10728 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 10 | 10730 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 10 | 10732 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 10 | 10734 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 10 | 10736 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 10 | 10738 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 10 | 10740 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 10 | 10742 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 10 | 10744 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 10 | 10746 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 10 | 10748 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 10 | 10750 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 10 | 10752 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 10 | 10754 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 10 | 10756 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 10 | 10758 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 10 | 10760 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 10 | 10762 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 10 | 10764 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 10 | 10766 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 10 | 10768 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 10 | 10770 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 10 | 10772 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 10 | 10774 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 10 | 10776 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 10 | 10778 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 10 | 10780 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 10 | 10782 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 10 | 10784 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 10 | 10786 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 10 | 10788 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 10 | 10790 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 10 | 10792 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 10 | 10794 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 10 | 10796 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 10 | 10798 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 10 | 10800 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 10 | 10802 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 10 | 10804 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 10 | 10806 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 10 | 10808 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 10 | 10810 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 10 | 10812 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 10 | 10814 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 10 | 10816 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 10 | 10818 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 10 | 10820 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 10 | 10822 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 10 | 10824 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 10 | 10826 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 10 | 10828 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 10 | 10830 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 10 | 10832 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 10 | 10834 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 10 | 10836 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 10 | 10838 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 10 | 10840 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 10 | 10842 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 10 | 10844 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 10 | 10846 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 10 | 10848 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 10 | 10850 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 10 | 10852 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 10 | 10854 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 10 | 10856 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 10 | 10858 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 10 | 10860 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 10 | 10862 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 10 | 10864 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 10 | 10866 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 10 | 10868 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 10 | 10870 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 10 | 10872 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 10 | 10874 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 10 | 10876 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 10 | 10878 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 10 | 10880 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 10 | 10882 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 10 | 10884 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 10 | 10886 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 10 | 10888 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 10 | 10890 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 10 | 10892 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 10 | 10894 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 10 | 10896 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 10 | 10898 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 10 | 10900 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 10 | 10902 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 11 | 10904 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 11 | 10906 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 11 | 10908 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 11 | 10910 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 11 | 10912 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 11 | 10914 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 11 | 10916 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 11 | 10918 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 11 | 10920 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 11 | 10922 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 11 | 10924 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 11 | 10926 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 11 | 10928 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 11 | 10930 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 11 | 10932 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 11 | 10934 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 11 | 10936 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 11 | 10938 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 11 | 10940 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 11 | 10942 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 11 | 10944 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 11 | 10946 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 11 | 10948 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 11 | 10950 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 11 | 10952 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 11 | 10954 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 11 | 10956 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 11 | 10958 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 11 | 10960 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 11 | 10962 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 11 | 10964 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 11 | 10966 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 11 | 10968 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 11 | 10970 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 11 | 10972 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 11 | 10974 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 11 | 10976 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 11 | 10978 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 11 | 10980 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 11 | 10982 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 11 | 10984 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 11 | 10986 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 11 | 10988 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 11 | 10990 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 11 | 10992 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 11 | 10994 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 11 | 10996 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 11 | 10998 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 11 | 11000 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 11 | 11002 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 11 | 11004 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 11 | 11006 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 11 | 11008 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 11 | 11010 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 11 | 11012 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 11 | 11014 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 11 | 11016 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 11 | 11018 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 11 | 11020 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 11 | 11022 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 11 | 11024 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 11 | 11026 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 11 | 11028 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 11 | 11030 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 11 | 11032 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 11 | 11034 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 11 | 11036 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 11 | 11038 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 11 | 11040 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 11 | 11042 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 11 | 11044 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 11 | 11046 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 11 | 11048 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 11 | 11050 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 11 | 11052 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 11 | 11054 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 11 | 11056 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 11 | 11058 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 11 | 11060 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 11 | 11062 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 11 | 11064 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 11 | 11066 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 11 | 11068 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 11 | 11070 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 11 | 11072 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 11 | 11074 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 11 | 11076 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 11 | 11078 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 11 | 11080 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 11 | 11082 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 11 | 11084 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 11 | 11086 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 11 | 11088 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 11 | 11090 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 11 | 11092 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 11 | 11094 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 11 | 11096 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 11 | 11098 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 11 | 11100 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 11 | 11102 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 11 | 11104 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 11 | 11106 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 11 | 11108 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 11 | 11110 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 11 | 11112 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 11 | 11114 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 11 | 11116 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 11 | 11118 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 11 | 11120 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 11 | 11122 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 11 | 11124 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 11 | 11126 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 11 | 11128 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 11 | 11130 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 11 | 11132 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 11 | 11134 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 11 | 11136 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 11 | 11138 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 11 | 11140 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 11 | 11142 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 12 | 11144 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 12 | 11146 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 12 | 11148 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 12 | 11150 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 12 | 11152 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 12 | 11154 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 12 | 11156 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 12 | 11158 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 12 | 11160 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 12 | 11162 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 12 | 11164 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 12 | 11166 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 12 | 11168 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 12 | 11170 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 12 | 11172 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 12 | 11174 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 12 | 11176 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 12 | 11178 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 12 | 11180 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 12 | 11182 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 12 | 11184 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 12 | 11186 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 12 | 11188 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 12 | 11190 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 12 | 11192 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 12 | 11194 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 12 | 11196 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 12 | 11198 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 12 | 11200 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 12 | 11202 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 12 | 11204 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 12 | 11206 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 12 | 11208 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 12 | 11210 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 12 | 11212 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 12 | 11214 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 12 | 11216 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 12 | 11218 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 12 | 11220 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 12 | 11222 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 12 | 11224 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 12 | 11226 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 12 | 11228 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 12 | 11230 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 12 | 11232 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 12 | 11234 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 12 | 11236 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 12 | 11238 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 12 | 11240 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 12 | 11242 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 12 | 11244 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 12 | 11246 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 12 | 11248 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 12 | 11250 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 12 | 11252 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 12 | 11254 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 12 | 11256 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 12 | 11258 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 12 | 11260 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 12 | 11262 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 12 | 11264 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 12 | 11266 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 12 | 11268 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 12 | 11270 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 12 | 11272 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 12 | 11274 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 12 | 11276 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 12 | 11278 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 12 | 11280 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 12 | 11282 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 12 | 11284 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 12 | 11286 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 12 | 11288 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 12 | 11290 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 12 | 11292 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 12 | 11294 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 12 | 11296 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 12 | 11298 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 12 | 11300 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 12 | 11302 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 12 | 11304 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 12 | 11306 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 12 | 11308 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 12 | 11310 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 12 | 11312 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 12 | 11314 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 12 | 11316 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 12 | 11318 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 12 | 11320 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 12 | 11322 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 12 | 11324 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 12 | 11326 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 12 | 11328 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 12 | 11330 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 12 | 11332 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 12 | 11334 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 12 | 11336 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 12 | 11338 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 12 | 11340 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 12 | 11342 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 12 | 11344 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 12 | 11346 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 12 | 11348 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 12 | 11350 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 12 | 11352 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 12 | 11354 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 12 | 11356 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 12 | 11358 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 12 | 11360 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 12 | 11362 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 12 | 11364 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 12 | 11366 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 12 | 11368 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 12 | 11370 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 12 | 11372 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 12 | 11374 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 12 | 11376 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 12 | 11378 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 12 | 11380 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 12 | 11382 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 13 | 11384 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 13 | 11386 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 13 | 11388 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 13 | 11390 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 13 | 11392 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 13 | 11394 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 13 | 11396 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 13 | 11398 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 13 | 11400 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 13 | 11402 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 13 | 11404 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 13 | 11406 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 13 | 11408 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 13 | 11410 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 13 | 11412 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 13 | 11414 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 13 | 11416 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 13 | 11418 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 13 | 11420 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 13 | 11422 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 13 | 11424 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 13 | 11426 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 13 | 11428 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 13 | 11430 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 13 | 11432 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 13 | 11434 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 13 | 11436 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 13 | 11438 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 13 | 11440 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 13 | 11442 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 13 | 11444 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 13 | 11446 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 13 | 11448 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 13 | 11450 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 13 | 11452 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 13 | 11454 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 13 | 11456 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 13 | 11458 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 13 | 11460 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 13 | 11462 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 13 | 11464 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 13 | 11466 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 13 | 11468 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 13 | 11470 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 13 | 11472 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 13 | 11474 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 13 | 11476 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 13 | 11478 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 13 | 11480 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 13 | 11482 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 13 | 11484 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 13 | 11486 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 13 | 11488 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 13 | 11490 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 13 | 11492 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 13 | 11494 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 13 | 11496 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 13 | 11498 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 13 | 11500 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 13 | 11502 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 13 | 11504 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 13 | 11506 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 13 | 11508 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 13 | 11510 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 13 | 11512 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 13 | 11514 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 13 | 11516 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 13 | 11518 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 13 | 11520 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 13 | 11522 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 13 | 11524 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 13 | 11526 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 13 | 11528 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 13 | 11530 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 13 | 11532 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 13 | 11534 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 13 | 11536 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 13 | 11538 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 13 | 11540 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 13 | 11542 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 13 | 11544 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 13 | 11546 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 13 | 11548 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 13 | 11550 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 13 | 11552 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 13 | 11554 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 13 | 11556 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 13 | 11558 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 13 | 11560 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 13 | 11562 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 13 | 11564 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 13 | 11566 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 13 | 11568 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 13 | 11570 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 13 | 11572 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 13 | 11574 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 13 | 11576 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 13 | 11578 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 13 | 11580 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 13 | 11582 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 13 | 11584 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 13 | 11586 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 13 | 11588 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 13 | 11590 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 13 | 11592 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 13 | 11594 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 13 | 11596 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 13 | 11598 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 13 | 11600 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 13 | 11602 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 13 | 11604 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 13 | 11606 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 13 | 11608 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 13 | 11610 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 13 | 11612 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 13 | 11614 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 13 | 11616 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 13 | 11618 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 13 | 11620 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 13 | 11622 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 14 | 11624 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 14 | 11626 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 14 | 11628 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 14 | 11630 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 14 | 11632 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 14 | 11634 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 14 | 11636 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 14 | 11638 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 14 | 11640 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 14 | 11642 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 14 | 11644 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 14 | 11646 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 14 | 11648 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 14 | 11650 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 14 | 11652 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 14 | 11654 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 14 | 11656 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 14 | 11658 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 14 | 11660 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 14 | 11662 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 14 | 11664 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 14 | 11666 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 14 | 11668 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 14 | 11670 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 14 | 11672 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 14 | 11674 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 14 | 11676 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 14 | 11678 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 14 | 11680 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 14 | 11682 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 14 | 11684 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 14 | 11686 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 14 | 11688 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 14 | 11690 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 14 | 11692 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 14 | 11694 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 14 | 11696 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 14 | 11698 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 14 | 11700 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 14 | 11702 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 14 | 11704 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 14 | 11706 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 14 | 11708 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 14 | 11710 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 14 | 11712 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 14 | 11714 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 14 | 11716 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 14 | 11718 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 14 | 11720 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 14 | 11722 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 14 | 11724 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 14 | 11726 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 14 | 11728 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 14 | 11730 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 14 | 11732 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 14 | 11734 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 14 | 11736 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 14 | 11738 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 14 | 11740 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 14 | 11742 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 14 | 11744 | Float | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 14 | 11746 | Float | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 14 | 11748 | Float | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 14 | 11750 | Float | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 14 | 11752 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 14 | 11754 | Float | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 14 | 11756 | Float | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 14 | 11758 | Float | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 14 | 11760 | Float | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 14 | 11762 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 14 | 11764 | Float | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 14 | 11766 | Float | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 14 | 11768 | Float | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 14 | 11770 | Float | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 14 | 11772 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 14 | 11774 | Float | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 14 | 11776 | Float | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 14 | 11778 | Float | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 14 | 11780 | Float | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 14 | 11782 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 14 | 11784 | Float | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 14 | 11786 | Float | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 14 | 11788 | Float | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 14 | 11790 | Float | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 14 | 11792 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 14 | 11794 | Float | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 14 | 11796 | Float | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 14 | 11798 | Float | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 14 | 11800 | Float | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 14 | 11802 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 14 | 11804 | Float | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 14 | 11806 | Float | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 14 | 11808 | Float | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 14 | 11810 | Float | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 14 | 11812 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 14 | 11814 | Float | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 14 | 11816 | Float | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 14 | 11818 | Float | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 14 | 11820 | Float | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 14 | 11822 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 14 | 11824 | Float | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 14 | 11826 | Float | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 14 | 11828 | Float | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 14 | 11830 | Float | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 14 | 11832 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 14 | 11834 | Float | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 14 | 11836 | Float | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 14 | 11838 | Float | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 14 | 11840 | Float | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 14 | 11842 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 14 | 11844 | Float | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 14 | 11846 | Float | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 14 | 11848 | Float | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 14 | 11850 | Float | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 14 | 11852 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 14 | 11854 | Float | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 14 | 11856 | Float | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 14 | 11858 | Float | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 14 | 11860 | Float | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 14 | 11862 | Float | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 15 | 11864 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 15 | 11866 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 15 | 11868 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 15 | 11870 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 15 | 11872 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 15 | 11874 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 15 | 11876 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 15 | 11878 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 15 | 11880 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 15 | 11882 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 15 | 11884 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 15 | 11886 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 15 | 11888 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 15 | 11890 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 15 | 11892 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 15 | 11894 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 15 | 11896 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 15 | 11898 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 15 | 11900 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 15 | 11902 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 15 | 11904 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 15 | 11906 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 15 | 11908 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 15 | 11910 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 15 | 11912 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 15 | 11914 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 15 | 11916 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 15 | 11918 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 15 | 11920 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 15 | 11922 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 15 | 11924 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 15 | 11926 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 15 | 11928 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 15 | 11930 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 15 | 11932 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 15 | 11934 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 15 | 11936 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 15 | 11938 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 15 | 11940 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 15 | 11942 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 15 | 11944 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 15 | 11946 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 15 | 11948 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 15 | 11950 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 15 | 11952 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 15 | 11954 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 15 | 11956 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 15 | 11958 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 15 | 11960 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 15 | 11962 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 15 | 11964 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 15 | 11966 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 15 | 11968 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 15 | 11970 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 15 | 11972 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 15 | 11974 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 15 | 11976 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 15 | 11978 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 15 | 11980 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 15 | 11982 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 15 | 11984 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 15 | 11986 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 15 | 11988 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 15 | 11990 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 15 | 11992 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 15 | 11994 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 15 | 11996 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 15 | 11998 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 15 | 12000 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 15 | 12002 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 15 | 12004 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 15 | 12006 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 15 | 12008 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 15 | 12010 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 15 | 12012 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 15 | 12014 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 15 | 12016 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 15 | 12018 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 15 | 12020 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 15 | 12022 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 15 | 12024 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 15 | 12026 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 15 | 12028 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 15 | 12030 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 15 | 12032 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 15 | 12034 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 15 | 12036 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 15 | 12038 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 15 | 12040 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 15 | 12042 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 15 | 12044 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 15 | 12046 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 15 | 12048 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 15 | 12050 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 15 | 12052 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 15 | 12054 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 15 | 12056 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 15 | 12058 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 15 | 12060 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 15 | 12062 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 15 | 12064 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 15 | 12066 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 15 | 12068 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 15 | 12070 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 15 | 12072 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 15 | 12074 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 15 | 12076 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 15 | 12078 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 15 | 12080 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 15 | 12082 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 15 | 12084 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 15 | 12086 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 15 | 12088 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 15 | 12090 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 15 | 12092 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 15 | 12094 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 15 | 12096 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 15 | 12098 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 15 | 12100 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 15 | 12102 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 16 | 12104 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 16 | 12106 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 16 | 12108 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 16 | 12110 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 16 | 12112 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 16 | 12114 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 16 | 12116 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 16 | 12118 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 16 | 12120 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 16 | 12122 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 16 | 12124 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 16 | 12126 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 16 | 12128 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 16 | 12130 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 16 | 12132 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 16 | 12134 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 16 | 12136 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 16 | 12138 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 16 | 12140 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 16 | 12142 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 16 | 12144 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 16 | 12146 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 16 | 12148 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 16 | 12150 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 16 | 12152 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 16 | 12154 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 16 | 12156 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 16 | 12158 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 16 | 12160 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 16 | 12162 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 16 | 12164 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 16 | 12166 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 16 | 12168 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 16 | 12170 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 16 | 12172 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 16 | 12174 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 16 | 12176 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 16 | 12178 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 16 | 12180 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 16 | 12182 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 16 | 12184 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 16 | 12186 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 16 | 12188 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 16 | 12190 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 16 | 12192 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 16 | 12194 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 16 | 12196 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 16 | 12198 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 16 | 12200 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 16 | 12202 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 16 | 12204 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 16 | 12206 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 16 | 12208 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 16 | 12210 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 16 | 12212 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 16 | 12214 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 16 | 12216 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 16 | 12218 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 16 | 12220 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 16 | 12222 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 16 | 12224 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 16 | 12226 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 16 | 12228 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 16 | 12230 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 16 | 12232 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 16 | 12234 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 16 | 12236 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 16 | 12238 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 16 | 12240 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 16 | 12242 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 16 | 12244 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 16 | 12246 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 16 | 12248 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 16 | 12250 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 16 | 12252 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 16 | 12254 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 16 | 12256 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 16 | 12258 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 16 | 12260 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 16 | 12262 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 16 | 12264 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 16 | 12266 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 16 | 12268 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 16 | 12270 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 16 | 12272 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 16 | 12274 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 16 | 12276 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 16 | 12278 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 16 | 12280 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 16 | 12282 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 16 | 12284 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 16 | 12286 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 16 | 12288 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 16 | 12290 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 16 | 12292 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 16 | 12294 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 16 | 12296 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 16 | 12298 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 16 | 12300 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 16 | 12302 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 16 | 12304 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 16 | 12306 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 16 | 12308 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 16 | 12310 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 16 | 12312 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 16 | 12314 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 16 | 12316 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 16 | 12318 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 16 | 12320 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 16 | 12322 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 16 | 12324 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 16 | 12326 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 16 | 12328 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 16 | 12330 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 16 | 12332 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 16 | 12334 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 16 | 12336 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 16 | 12338 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 16 | 12340 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 16 | 12342 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 17 | 12344 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 17 | 12346 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 17 | 12348 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 17 | 12350 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 17 | 12352 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 17 | 12354 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 17 | 12356 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 17 | 12358 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 17 | 12360 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 17 | 12362 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 17 | 12364 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 17 | 12366 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 17 | 12368 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 17 | 12370 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 17 | 12372 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 17 | 12374 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 17 | 12376 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 17 | 12378 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 17 | 12380 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 17 | 12382 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 17 | 12384 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 17 | 12386 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 17 | 12388 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 17 | 12390 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 17 | 12392 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 17 | 12394 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 17 | 12396 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 17 | 12398 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 17 | 12400 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 17 | 12402 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 17 | 12404 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 17 | 12406 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 17 | 12408 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 17 | 12410 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 17 | 12412 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 17 | 12414 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 17 | 12416 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 17 | 12418 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 17 | 12420 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 17 | 12422 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 17 | 12424 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 17 | 12426 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 17 | 12428 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 17 | 12430 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 17 | 12432 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 17 | 12434 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 17 | 12436 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 17 | 12438 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 17 | 12440 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 17 | 12442 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 17 | 12444 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 17 | 12446 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 17 | 12448 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 17 | 12450 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 17 | 12452 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 17 | 12454 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 17 | 12456 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 17 | 12458 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 17 | 12460 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 17 | 12462 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 17 | 12464 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 17 | 12466 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 17 | 12468 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 17 | 12470 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 17 | 12472 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 17 | 12474 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 17 | 12476 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 17 | 12478 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 17 | 12480 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 17 | 12482 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 17 | 12484 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 17 | 12486 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 17 | 12488 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 17 | 12490 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 17 | 12492 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 17 | 12494 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 17 | 12496 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 17 | 12498 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 17 | 12500 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 17 | 12502 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 17 | 12504 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 17 | 12506 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 17 | 12508 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 17 | 12510 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 17 | 12512 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 17 | 12514 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 17 | 12516 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 17 | 12518 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 17 | 12520 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 17 | 12522 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 17 | 12524 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 17 | 12526 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 17 | 12528 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 17 | 12530 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 17 | 12532 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 17 | 12534 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 17 | 12536 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 17 | 12538 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 17 | 12540 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 17 | 12542 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 17 | 12544 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 17 | 12546 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 17 | 12548 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 17 | 12550 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 17 | 12552 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 17 | 12554 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 17 | 12556 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 17 | 12558 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 17 | 12560 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 17 | 12562 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 17 | 12564 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 17 | 12566 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 17 | 12568 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 17 | 12570 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 17 | 12572 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 17 | 12574 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 17 | 12576 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 17 | 12578 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 17 | 12580 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 17 | 12582 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 18 | 12584 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 18 | 12586 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 18 | 12588 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 18 | 12590 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 18 | 12592 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 18 | 12594 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 18 | 12596 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 18 | 12598 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 18 | 12600 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 18 | 12602 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 18 | 12604 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 18 | 12606 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 18 | 12608 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 18 | 12610 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 18 | 12612 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 18 | 12614 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 18 | 12616 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 18 | 12618 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 18 | 12620 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 18 | 12622 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 18 | 12624 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 18 | 12626 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 18 | 12628 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 18 | 12630 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 18 | 12632 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 18 | 12634 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 18 | 12636 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 18 | 12638 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 18 | 12640 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 18 | 12642 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 18 | 12644 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 18 | 12646 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 18 | 12648 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 18 | 12650 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 18 | 12652 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 18 | 12654 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 18 | 12656 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 18 | 12658 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 18 | 12660 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 18 | 12662 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 18 | 12664 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 18 | 12666 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 18 | 12668 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 18 | 12670 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 18 | 12672 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 18 | 12674 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 18 | 12676 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 18 | 12678 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 18 | 12680 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 18 | 12682 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 18 | 12684 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 18 | 12686 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 18 | 12688 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 18 | 12690 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 18 | 12692 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 18 | 12694 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 18 | 12696 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 18 | 12698 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 18 | 12700 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 18 | 12702 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 18 | 12704 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 18 | 12706 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 18 | 12708 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 18 | 12710 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 18 | 12712 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 18 | 12714 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 18 | 12716 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 18 | 12718 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 18 | 12720 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 18 | 12722 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 18 | 12724 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 18 | 12726 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 18 | 12728 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 18 | 12730 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 18 | 12732 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 18 | 12734 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 18 | 12736 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 18 | 12738 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 18 | 12740 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 18 | 12742 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 18 | 12744 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 18 | 12746 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 18 | 12748 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 18 | 12750 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 18 | 12752 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 18 | 12754 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 18 | 12756 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 18 | 12758 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 18 | 12760 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 18 | 12762 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 18 | 12764 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 18 | 12766 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 18 | 12768 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 18 | 12770 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 18 | 12772 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 18 | 12774 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 18 | 12776 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 18 | 12778 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 18 | 12780 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 18 | 12782 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 18 | 12784 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 18 | 12786 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 18 | 12788 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 18 | 12790 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 18 | 12792 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 18 | 12794 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 18 | 12796 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 18 | 12798 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 18 | 12800 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 18 | 12802 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 18 | 12804 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 18 | 12806 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 18 | 12808 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 18 | 12810 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 18 | 12812 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 18 | 12814 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 18 | 12816 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 18 | 12818 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 18 | 12820 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 18 | 12822 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 19 | 12824 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 19 | 12826 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 19 | 12828 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 19 | 12830 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 19 | 12832 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 19 | 12834 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 19 | 12836 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 19 | 12838 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 19 | 12840 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 19 | 12842 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 19 | 12844 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 19 | 12846 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 19 | 12848 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 19 | 12850 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 19 | 12852 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 19 | 12854 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 19 | 12856 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 19 | 12858 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 19 | 12860 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 19 | 12862 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 19 | 12864 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 19 | 12866 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 19 | 12868 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 19 | 12870 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 19 | 12872 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 19 | 12874 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 19 | 12876 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 19 | 12878 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 19 | 12880 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 19 | 12882 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 19 | 12884 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 19 | 12886 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 19 | 12888 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 19 | 12890 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 19 | 12892 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 19 | 12894 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 19 | 12896 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 19 | 12898 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 19 | 12900 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 19 | 12902 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 19 | 12904 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 19 | 12906 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 19 | 12908 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 19 | 12910 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 19 | 12912 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 19 | 12914 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 19 | 12916 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 19 | 12918 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 19 | 12920 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 19 | 12922 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 19 | 12924 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 19 | 12926 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 19 | 12928 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 19 | 12930 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 19 | 12932 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 19 | 12934 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 19 | 12936 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 19 | 12938 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 19 | 12940 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 19 | 12942 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 19 | 12944 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 19 | 12946 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 19 | 12948 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 19 | 12950 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 19 | 12952 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 19 | 12954 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 19 | 12956 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 19 | 12958 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 19 | 12960 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 19 | 12962 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 19 | 12964 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 19 | 12966 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 19 | 12968 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 19 | 12970 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 19 | 12972 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 19 | 12974 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 19 | 12976 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 19 | 12978 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 19 | 12980 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 19 | 12982 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 19 | 12984 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 19 | 12986 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 19 | 12988 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 19 | 12990 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 19 | 12992 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 19 | 12994 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 19 | 12996 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 19 | 12998 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 19 | 13000 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 19 | 13002 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 19 | 13004 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 19 | 13006 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 19 | 13008 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 19 | 13010 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 19 | 13012 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 19 | 13014 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 19 | 13016 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 19 | 13018 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 19 | 13020 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 19 | 13022 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 19 | 13024 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 19 | 13026 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 19 | 13028 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 19 | 13030 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 19 | 13032 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 19 | 13034 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 19 | 13036 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 19 | 13038 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 19 | 13040 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 19 | 13042 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 19 | 13044 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 19 | 13046 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 19 | 13048 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 19 | 13050 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 19 | 13052 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 19 | 13054 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 19 | 13056 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 19 | 13058 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 19 | 13060 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 19 | 13062 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 20 | 13064 | Float | RD | A | Minimum Value L1 Current (RMS)
Current measurement module group 20 | 13066 | Float | RD | A | Minimum Value L2 Current (RMS)
Current measurement module group 20 | 13068 | Float | RD | A | Minimum Value L3 Current (RMS)
Current measurement module group 20 | 13070 | Float | RD | A | Minimum Value L4 Current (RMS)
Current measurement module group 20 | 13072 | Float | RD | A | Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 20 | 13074 | Float | RD | W | Minimum Value L1 Active Power
Current measurement module group 20 | 13076 | Float | RD | W | Minimum Value L2 Active Power
Current measurement module group 20 | 13078 | Float | RD | W | Minimum Value L3 Active Power
Current measurement module group 20 | 13080 | Float | RD | W | Minimum Value L4 Active Power
Current measurement module group 20 | 13082 | Float | RD | W | Minimum Value Sum L1..L4 Active Power
Current measurement module group 20 | 13084 | Float | RD | var | Minimum Value L1 Reactive Power
Current measurement module group 20 | 13086 | Float | RD | var | Minimum Value L2 Reactive Power
Current measurement module group 20 | 13088 | Float | RD | var | Minimum Value L3 Reactive Power
Current measurement module group 20 | 13090 | Float | RD | var | Minimum Value L4 Reactive Power
Current measurement module group 20 | 13092 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 20 | 13094 | Float | RD | VA | Minimum Value L1 Apparent Power
Current measurement module group 20 | 13096 | Float | RD | VA | Minimum Value L2 Apparent Power
Current measurement module group 20 | 13098 | Float | RD | VA | Minimum Value L3 Apparent Power
Current measurement module group 20 | 13100 | Float | RD | VA | Minimum Value L4 Apparent Power
Current measurement module group 20 | 13102 | Float | RD | VA | Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 20 | 13104 | Float | RD | - | Minimum Value L1 Cos(phi)
Current measurement module group 20 | 13106 | Float | RD | - | Minimum Value L2 Cos(phi)
Current measurement module group 20 | 13108 | Float | RD | - | Minimum Value L3 Cos(phi)
Current measurement module group 20 | 13110 | Float | RD | - | Minimum Value L4 Cos(phi)
Current measurement module group 20 | 13112 | Float | RD | - | Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 20 | 13114 | Float | RD | - | Minimum Value L1 Power Factor
Current measurement module group 20 | 13116 | Float | RD | - | Minimum Value L2 Power Factor
Current measurement module group 20 | 13118 | Float | RD | - | Minimum Value L3 Power Factor
Current measurement module group 20 | 13120 | Float | RD | - | Minimum Value L4 Power Factor
Current measurement module group 20 | 13122 | Float | RD | - | Minimum Value Sum L1..L4 Power Factor
Current measurement module group 20 | 13124 | Float | RD | W | Minimum Value L1 Active Power (fundamental)
Current measurement module group 20 | 13126 | Float | RD | W | Minimum Value L2 Active Power (fundamental)
Current measurement module group 20 | 13128 | Float | RD | W | Minimum Value L3 Active Power (fundamental)
Current measurement module group 20 | 13130 | Float | RD | W | Minimum Value L4 Active Power (fundamental)
Current measurement module group 20 | 13132 | Float | RD | W | Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 20 | 13134 | Float | RD | var | Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 20 | 13136 | Float | RD | var | Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 20 | 13138 | Float | RD | var | Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 20 | 13140 | Float | RD | var | Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 20 | 13142 | Float | RD | var | Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 20 | 13144 | Float | RD | var | Minimum Value L1 Distortion Reactive Power
Current measurement module group 20 | 13146 | Float | RD | var | Minimum Value L2 Distortion Reactive Power
Current measurement module group 20 | 13148 | Float | RD | var | Minimum Value L3 Distortion Reactive Power
Current measurement module group 20 | 13150 | Float | RD | var | Minimum Value L4 Distortion Reactive Power
Current measurement module group 20 | 13152 | Float | RD | var | Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 20 | 13154 | Float | RD | % | Minimum Value L1 THD Current
Current measurement module group 20 | 13156 | Float | RD | % | Minimum Value L2 THD Current
Current measurement module group 20 | 13158 | Float | RD | % | Minimum Value L3 THD Current
Current measurement module group 20 | 13160 | Float | RD | % | Minimum Value L4 THD Current
Current measurement module group 20 | 13162 | Float | RD | % | Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 20 | 13164 | Float | RD | % | Minimum Value L1 TDD Current
Current measurement module group 20 | 13166 | Float | RD | % | Minimum Value L2 TDD Current
Current measurement module group 20 | 13168 | Float | RD | % | Minimum Value L3 TDD Current
Current measurement module group 20 | 13170 | Float | RD | % | Minimum Value L4 TDD Current
Current measurement module group 20 | 13172 | Float | RD | % | Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 20 | 13174 | Float | RD | % | Minimum Value L1 Load Current (%)
Current measurement module group 20 | 13176 | Float | RD | % | Minimum Value L2 Load Current (%)
Current measurement module group 20 | 13178 | Float | RD | % | Minimum Value L3 Load Current (%)
Current measurement module group 20 | 13180 | Float | RD | % | Minimum Value L4 Load Current (%)
Current measurement module group 20 | 13182 | Float | RD | % | Minimum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 20 | 13184 | DateTime | RD | - | DateTime Minimum Value L1 Current (RMS)
Current measurement module group 20 | 13186 | DateTime | RD | - | DateTime Minimum Value L2 Current (RMS)
Current measurement module group 20 | 13188 | DateTime | RD | - | DateTime Minimum Value L3 Current (RMS)
Current measurement module group 20 | 13190 | DateTime | RD | - | DateTime Minimum Value L4 Current (RMS)
Current measurement module group 20 | 13192 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Current (RMS)
Current measurement module group 20 | 13194 | DateTime | RD | - | DateTime Minimum Value L1 Active Power
Current measurement module group 20 | 13196 | DateTime | RD | - | DateTime Minimum Value L2 Active Power
Current measurement module group 20 | 13198 | DateTime | RD | - | DateTime Minimum Value L3 Active Power
Current measurement module group 20 | 13200 | DateTime | RD | - | DateTime Minimum Value L4 Active Power
Current measurement module group 20 | 13202 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power
Current measurement module group 20 | 13204 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power
Current measurement module group 20 | 13206 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power
Current measurement module group 20 | 13208 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power
Current measurement module group 20 | 13210 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power
Current measurement module group 20 | 13212 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power
Current measurement module group 20 | 13214 | DateTime | RD | - | DateTime Minimum Value L1 Apparent Power
Current measurement module group 20 | 13216 | DateTime | RD | - | DateTime Minimum Value L2 Apparent Power
Current measurement module group 20 | 13218 | DateTime | RD | - | DateTime Minimum Value L3 Apparent Power
Current measurement module group 20 | 13220 | DateTime | RD | - | DateTime Minimum Value L4 Apparent Power
Current measurement module group 20 | 13222 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Apparent Power
Current measurement module group 20 | 13224 | DateTime | RD | - | DateTime Minimum Value L1 Cos(phi)
Current measurement module group 20 | 13226 | DateTime | RD | - | DateTime Minimum Value L2 Cos(phi)
Current measurement module group 20 | 13228 | DateTime | RD | - | DateTime Minimum Value L3 Cos(phi)
Current measurement module group 20 | 13230 | DateTime | RD | - | DateTime Minimum Value L4 Cos(phi)
Current measurement module group 20 | 13232 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Cos(phi)
Current measurement module group 20 | 13234 | DateTime | RD | - | DateTime Minimum Value L1 Power Factor
Current measurement module group 20 | 13236 | DateTime | RD | - | DateTime Minimum Value L2 Power Factor
Current measurement module group 20 | 13238 | DateTime | RD | - | DateTime Minimum Value L3 Power Factor
Current measurement module group 20 | 13240 | DateTime | RD | - | DateTime Minimum Value L4 Power Factor
Current measurement module group 20 | 13242 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Power Factor
Current measurement module group 20 | 13244 | DateTime | RD | - | DateTime Minimum Value L1 Active Power (fundamental)
Current measurement module group 20 | 13246 | DateTime | RD | - | DateTime Minimum Value L2 Active Power (fundamental)
Current measurement module group 20 | 13248 | DateTime | RD | - | DateTime Minimum Value L3 Active Power (fundamental)
Current measurement module group 20 | 13250 | DateTime | RD | - | DateTime Minimum Value L4 Active Power (fundamental)
Current measurement module group 20 | 13252 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 20 | 13254 | DateTime | RD | - | DateTime Minimum Value L1 Reactive Power (fundamental)
Current measurement module group 20 | 13256 | DateTime | RD | - | DateTime Minimum Value L2 Reactive Power (fundamental)
Current measurement module group 20 | 13258 | DateTime | RD | - | DateTime Minimum Value L3 Reactive Power (fundamental)
Current measurement module group 20 | 13260 | DateTime | RD | - | DateTime Minimum Value L4 Reactive Power (fundamental)
Current measurement module group 20 | 13262 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 20 | 13264 | DateTime | RD | - | DateTime Minimum Value L1 Distortion Reactive Power
Current measurement module group 20 | 13266 | DateTime | RD | - | DateTime Minimum Value L2 Distortion Reactive Power
Current measurement module group 20 | 13268 | DateTime | RD | - | DateTime Minimum Value L3 Distortion Reactive Power
Current measurement module group 20 | 13270 | DateTime | RD | - | DateTime Minimum Value L4 Distortion Reactive Power
Current measurement module group 20 | 13272 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 20 | 13274 | DateTime | RD | - | DateTime Minimum Value L1 THD Current
Current measurement module group 20 | 13276 | DateTime | RD | - | DateTime Minimum Value L2 THD Current
Current measurement module group 20 | 13278 | DateTime | RD | - | DateTime Minimum Value L3 THD Current
Current measurement module group 20 | 13280 | DateTime | RD | - | DateTime Minimum Value L4 THD Current
Current measurement module group 20 | 13282 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum THD Current
Current measurement module group 20 | 13284 | DateTime | RD | - | DateTime Minimum Value L1 TDD Current
Current measurement module group 20 | 13286 | DateTime | RD | - | DateTime Minimum Value L2 TDD Current
Current measurement module group 20 | 13288 | DateTime | RD | - | DateTime Minimum Value L3 TDD Current
Current measurement module group 20 | 13290 | DateTime | RD | - | DateTime Minimum Value L4 TDD Current
Current measurement module group 20 | 13292 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum TDD Current
Current measurement module group 20 | 13294 | DateTime | RD | - | DateTime Minimum Value L1 Load Current (%)
Current measurement module group 20 | 13296 | DateTime | RD | - | DateTime Minimum Value L2 Load Current (%)
Current measurement module group 20 | 13298 | DateTime | RD | - | DateTime Minimum Value L3 Load Current (%)
Current measurement module group 20 | 13300 | DateTime | RD | - | DateTime Minimum Value L4 Load Current (%)
Current measurement module group 20 | 13302 | DateTime | RD | - | DateTime Minimum Value Sum L1..L4 Sum Load Current (%)
basic device | 13600 | Float | RD | V | Maximum Value L1 Voltage L-N (RMS)
basic device | 13602 | Float | RD | V | Maximum Value L2 Voltage L-N (RMS)
basic device | 13604 | Float | RD | V | Maximum Value L3 Voltage L-N (RMS)
basic device | 13606 | Float | RD | V | Maximum Value L1-L2 Voltage L-L (RMS)
basic device | 13608 | Float | RD | V | Maximum Value L2-L3 Voltage L-L (RMS)
basic device | 13610 | Float | RD | V | Maximum Value L3-L1 Voltage L-L (RMS)
Basic device measuring group 01 | 13612 | Float | RD | A | Maximum Value L1 Current (RMS)
Basic device measuring group 01 | 13614 | Float | RD | A | Maximum Value L2 Current (RMS)
Basic device measuring group 01 | 13616 | Float | RD | A | Maximum Value L3 Current (RMS)
Basic device measuring group 01 | 13618 | Float | RD | A | Maximum Value L4 Current (RMS)
Basic device measuring group 01 | 13620 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Basic device measuring group 01 | 13622 | Float | RD | W | Maximum Value L1 Active Power
Basic device measuring group 01 | 13624 | Float | RD | W | Maximum Value L2 Active Power
Basic device measuring group 01 | 13626 | Float | RD | W | Maximum Value L3 Active Power
Basic device measuring group 01 | 13628 | Float | RD | W | Maximum Value L4 Active Power
Basic device measuring group 01 | 13630 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Basic device measuring group 01 | 13632 | Float | RD | var | Maximum Value L1 Reactive Power
Basic device measuring group 01 | 13634 | Float | RD | var | Maximum Value L2 Reactive Power
Basic device measuring group 01 | 13636 | Float | RD | var | Maximum Value L3 Reactive Power
Basic device measuring group 01 | 13638 | Float | RD | var | Maximum Value L4 Reactive Power
Basic device measuring group 01 | 13640 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Basic device measuring group 01 | 13642 | Float | RD | VA | Maximum Value L1 Apparent Power
Basic device measuring group 01 | 13644 | Float | RD | VA | Maximum Value L2 Apparent Power
Basic device measuring group 01 | 13646 | Float | RD | VA | Maximum Value L3 Apparent Power
Basic device measuring group 01 | 13648 | Float | RD | VA | Maximum Value L4 Apparent Power
Basic device measuring group 01 | 13650 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Basic device measuring group 01 | 13652 | Float | RD | - | Maximum Value L1 Cos(phi)
Basic device measuring group 01 | 13654 | Float | RD | - | Maximum Value L2 Cos(phi)
Basic device measuring group 01 | 13656 | Float | RD | - | Maximum Value L3 Cos(phi)
Basic device measuring group 01 | 13658 | Float | RD | - | Maximum Value L4 Cos(phi)
Basic device measuring group 01 | 13660 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Basic device measuring group 01 | 13662 | Float | RD | - | Maximum Value L1 Power Factor
Basic device measuring group 01 | 13664 | Float | RD | - | Maximum Value L2 Power Factor
Basic device measuring group 01 | 13666 | Float | RD | - | Maximum Value L3 Power Factor
Basic device measuring group 01 | 13668 | Float | RD | - | Maximum Value L4 Power Factor
Basic device measuring group 01 | 13670 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Basic device measuring group 01 | 13672 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Basic device measuring group 01 | 13674 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Basic device measuring group 01 | 13676 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Basic device measuring group 01 | 13678 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Basic device measuring group 01 | 13680 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 01 | 13682 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Basic device measuring group 01 | 13684 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Basic device measuring group 01 | 13686 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Basic device measuring group 01 | 13688 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Basic device measuring group 01 | 13690 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 01 | 13692 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Basic device measuring group 01 | 13694 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Basic device measuring group 01 | 13696 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Basic device measuring group 01 | 13698 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Basic device measuring group 01 | 13700 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 01 | 13702 | Float | RD | % | Maximum Value L1 THD Current
Basic device measuring group 01 | 13704 | Float | RD | % | Maximum Value L2 THD Current
Basic device measuring group 01 | 13706 | Float | RD | % | Maximum Value L3 THD Current
Basic device measuring group 01 | 13708 | Float | RD | % | Maximum Value L4 THD Current
Basic device measuring group 01 | 13710 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Basic device measuring group 01 | 13712 | Float | RD | % | Maximum Value L1 TDD Current
Basic device measuring group 01 | 13714 | Float | RD | % | Maximum Value L2 TDD Current
Basic device measuring group 01 | 13716 | Float | RD | % | Maximum Value L3 TDD Current
Basic device measuring group 01 | 13718 | Float | RD | % | Maximum Value L4 TDD Current
Basic device measuring group 01 | 13720 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 01 | 13722 | Float | RD | % | Maximum Value L1 Load Current (%)
Basic device measuring group 01 | 13724 | Float | RD | % | Maximum Value L2 Load Current (%)
Basic device measuring group 01 | 13726 | Float | RD | % | Maximum Value L3 Load Current (%)
Basic device measuring group 01 | 13728 | Float | RD | % | Maximum Value L4 Load Current (%)
Basic device measuring group 01 | 13730 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
basic device | 13732 | DateTime | RD | - | DateTime Maximum Value L1 Voltage L-N (RMS)
basic device | 13734 | DateTime | RD | - | DateTime Maximum Value L2 Voltage L-N (RMS)
basic device | 13736 | DateTime | RD | - | DateTime Maximum Value L3 Voltage L-N (RMS)
basic device | 13738 | DateTime | RD | - | DateTime Maximum Value L1-L2 Voltage L-L (RMS)
basic device | 13740 | DateTime | RD | - | DateTime Maximum Value L2-L3 Voltage L-L (RMS)
basic device | 13742 | DateTime | RD | - | DateTime Maximum Value L3-L1 Voltage L-L (RMS)
Basic device measuring group 01 | 13744 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Basic device measuring group 01 | 13746 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Basic device measuring group 01 | 13748 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Basic device measuring group 01 | 13750 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Basic device measuring group 01 | 13752 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Basic device measuring group 01 | 13754 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Basic device measuring group 01 | 13756 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Basic device measuring group 01 | 13758 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Basic device measuring group 01 | 13760 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Basic device measuring group 01 | 13762 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Basic device measuring group 01 | 13764 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Basic device measuring group 01 | 13766 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Basic device measuring group 01 | 13768 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Basic device measuring group 01 | 13770 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Basic device measuring group 01 | 13772 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Basic device measuring group 01 | 13774 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Basic device measuring group 01 | 13776 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Basic device measuring group 01 | 13778 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Basic device measuring group 01 | 13780 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Basic device measuring group 01 | 13782 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Basic device measuring group 01 | 13784 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Basic device measuring group 01 | 13786 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Basic device measuring group 01 | 13788 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Basic device measuring group 01 | 13790 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Basic device measuring group 01 | 13792 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Basic device measuring group 01 | 13794 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Basic device measuring group 01 | 13796 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Basic device measuring group 01 | 13798 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Basic device measuring group 01 | 13800 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Basic device measuring group 01 | 13802 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Basic device measuring group 01 | 13804 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Basic device measuring group 01 | 13806 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Basic device measuring group 01 | 13808 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Basic device measuring group 01 | 13810 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Basic device measuring group 01 | 13812 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 01 | 13814 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Basic device measuring group 01 | 13816 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Basic device measuring group 01 | 13818 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Basic device measuring group 01 | 13820 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Basic device measuring group 01 | 13822 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 01 | 13824 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Basic device measuring group 01 | 13826 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Basic device measuring group 01 | 13828 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Basic device measuring group 01 | 13830 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Basic device measuring group 01 | 13832 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 01 | 13834 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Basic device measuring group 01 | 13836 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Basic device measuring group 01 | 13838 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Basic device measuring group 01 | 13840 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Basic device measuring group 01 | 13842 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Basic device measuring group 01 | 13844 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Basic device measuring group 01 | 13846 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Basic device measuring group 01 | 13848 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Basic device measuring group 01 | 13850 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Basic device measuring group 01 | 13852 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 01 | 13854 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Basic device measuring group 01 | 13856 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Basic device measuring group 01 | 13858 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Basic device measuring group 01 | 13860 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Basic device measuring group 01 | 13862 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Basic device measuring group 02 | 13864 | Float | RD | A | Maximum Value L1 Current (RMS)
Basic device measuring group 02 | 13866 | Float | RD | A | Maximum Value L2 Current (RMS)
Basic device measuring group 02 | 13868 | Float | RD | A | Maximum Value L3 Current (RMS)
Basic device measuring group 02 | 13870 | Float | RD | A | Maximum Value L4 Current (RMS)
Basic device measuring group 02 | 13872 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Basic device measuring group 02 | 13874 | Float | RD | W | Maximum Value L1 Active Power
Basic device measuring group 02 | 13876 | Float | RD | W | Maximum Value L2 Active Power
Basic device measuring group 02 | 13878 | Float | RD | W | Maximum Value L3 Active Power
Basic device measuring group 02 | 13880 | Float | RD | W | Maximum Value L4 Active Power
Basic device measuring group 02 | 13882 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Basic device measuring group 02 | 13884 | Float | RD | var | Maximum Value L1 Reactive Power
Basic device measuring group 02 | 13886 | Float | RD | var | Maximum Value L2 Reactive Power
Basic device measuring group 02 | 13888 | Float | RD | var | Maximum Value L3 Reactive Power
Basic device measuring group 02 | 13890 | Float | RD | var | Maximum Value L4 Reactive Power
Basic device measuring group 02 | 13892 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Basic device measuring group 02 | 13894 | Float | RD | VA | Maximum Value L1 Apparent Power
Basic device measuring group 02 | 13896 | Float | RD | VA | Maximum Value L2 Apparent Power
Basic device measuring group 02 | 13898 | Float | RD | VA | Maximum Value L3 Apparent Power
Basic device measuring group 02 | 13900 | Float | RD | VA | Maximum Value L4 Apparent Power
Basic device measuring group 02 | 13902 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Basic device measuring group 02 | 13904 | Float | RD | - | Maximum Value L1 Cos(phi)
Basic device measuring group 02 | 13906 | Float | RD | - | Maximum Value L2 Cos(phi)
Basic device measuring group 02 | 13908 | Float | RD | - | Maximum Value L3 Cos(phi)
Basic device measuring group 02 | 13910 | Float | RD | - | Maximum Value L4 Cos(phi)
Basic device measuring group 02 | 13912 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Basic device measuring group 02 | 13914 | Float | RD | - | Maximum Value L1 Power Factor
Basic device measuring group 02 | 13916 | Float | RD | - | Maximum Value L2 Power Factor
Basic device measuring group 02 | 13918 | Float | RD | - | Maximum Value L3 Power Factor
Basic device measuring group 02 | 13920 | Float | RD | - | Maximum Value L4 Power Factor
Basic device measuring group 02 | 13922 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Basic device measuring group 02 | 13924 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Basic device measuring group 02 | 13926 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Basic device measuring group 02 | 13928 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Basic device measuring group 02 | 13930 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Basic device measuring group 02 | 13932 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 02 | 13934 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Basic device measuring group 02 | 13936 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Basic device measuring group 02 | 13938 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Basic device measuring group 02 | 13940 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Basic device measuring group 02 | 13942 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 02 | 13944 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Basic device measuring group 02 | 13946 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Basic device measuring group 02 | 13948 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Basic device measuring group 02 | 13950 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Basic device measuring group 02 | 13952 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 02 | 13954 | Float | RD | % | Maximum Value L1 THD Current
Basic device measuring group 02 | 13956 | Float | RD | % | Maximum Value L2 THD Current
Basic device measuring group 02 | 13958 | Float | RD | % | Maximum Value L3 THD Current
Basic device measuring group 02 | 13960 | Float | RD | % | Maximum Value L4 THD Current
Basic device measuring group 02 | 13962 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Basic device measuring group 02 | 13964 | Float | RD | % | Maximum Value L1 TDD Current
Basic device measuring group 02 | 13966 | Float | RD | % | Maximum Value L2 TDD Current
Basic device measuring group 02 | 13968 | Float | RD | % | Maximum Value L3 TDD Current
Basic device measuring group 02 | 13970 | Float | RD | % | Maximum Value L4 TDD Current
Basic device measuring group 02 | 13972 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 02 | 13974 | Float | RD | % | Maximum Value L1 Load Current (%)
Basic device measuring group 02 | 13976 | Float | RD | % | Maximum Value L2 Load Current (%)
Basic device measuring group 02 | 13978 | Float | RD | % | Maximum Value L3 Load Current (%)
Basic device measuring group 02 | 13980 | Float | RD | % | Maximum Value L4 Load Current (%)
Basic device measuring group 02 | 13982 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Basic device measuring group 02 | 13984 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Basic device measuring group 02 | 13986 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Basic device measuring group 02 | 13988 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Basic device measuring group 02 | 13990 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Basic device measuring group 02 | 13992 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Basic device measuring group 02 | 13994 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Basic device measuring group 02 | 13996 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Basic device measuring group 02 | 13998 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Basic device measuring group 02 | 14000 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Basic device measuring group 02 | 14002 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Basic device measuring group 02 | 14004 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Basic device measuring group 02 | 14006 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Basic device measuring group 02 | 14008 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Basic device measuring group 02 | 14010 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Basic device measuring group 02 | 14012 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Basic device measuring group 02 | 14014 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Basic device measuring group 02 | 14016 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Basic device measuring group 02 | 14018 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Basic device measuring group 02 | 14020 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Basic device measuring group 02 | 14022 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Basic device measuring group 02 | 14024 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Basic device measuring group 02 | 14026 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Basic device measuring group 02 | 14028 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Basic device measuring group 02 | 14030 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Basic device measuring group 02 | 14032 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Basic device measuring group 02 | 14034 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Basic device measuring group 02 | 14036 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Basic device measuring group 02 | 14038 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Basic device measuring group 02 | 14040 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Basic device measuring group 02 | 14042 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Basic device measuring group 02 | 14044 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Basic device measuring group 02 | 14046 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Basic device measuring group 02 | 14048 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Basic device measuring group 02 | 14050 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Basic device measuring group 02 | 14052 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Basic device measuring group 02 | 14054 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Basic device measuring group 02 | 14056 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Basic device measuring group 02 | 14058 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Basic device measuring group 02 | 14060 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Basic device measuring group 02 | 14062 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Basic device measuring group 02 | 14064 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Basic device measuring group 02 | 14066 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Basic device measuring group 02 | 14068 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Basic device measuring group 02 | 14070 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Basic device measuring group 02 | 14072 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Basic device measuring group 02 | 14074 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Basic device measuring group 02 | 14076 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Basic device measuring group 02 | 14078 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Basic device measuring group 02 | 14080 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Basic device measuring group 02 | 14082 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Basic device measuring group 02 | 14084 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Basic device measuring group 02 | 14086 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Basic device measuring group 02 | 14088 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Basic device measuring group 02 | 14090 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Basic device measuring group 02 | 14092 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Basic device measuring group 02 | 14094 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Basic device measuring group 02 | 14096 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Basic device measuring group 02 | 14098 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Basic device measuring group 02 | 14100 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Basic device measuring group 02 | 14102 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 01 | 14104 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 01 | 14106 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 01 | 14108 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 01 | 14110 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 01 | 14112 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 01 | 14114 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 01 | 14116 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 01 | 14118 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 01 | 14120 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 01 | 14122 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 01 | 14124 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 01 | 14126 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 01 | 14128 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 01 | 14130 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 01 | 14132 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 01 | 14134 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 01 | 14136 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 01 | 14138 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 01 | 14140 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 01 | 14142 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 01 | 14144 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 01 | 14146 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 01 | 14148 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 01 | 14150 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 01 | 14152 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 01 | 14154 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 01 | 14156 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 01 | 14158 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 01 | 14160 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 01 | 14162 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 01 | 14164 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 01 | 14166 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 01 | 14168 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 01 | 14170 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 01 | 14172 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 01 | 14174 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 01 | 14176 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 01 | 14178 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 01 | 14180 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 01 | 14182 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 01 | 14184 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 01 | 14186 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 01 | 14188 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 01 | 14190 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 01 | 14192 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 01 | 14194 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 01 | 14196 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 01 | 14198 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 01 | 14200 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 01 | 14202 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 01 | 14204 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 01 | 14206 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 01 | 14208 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 01 | 14210 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 01 | 14212 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 01 | 14214 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 01 | 14216 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 01 | 14218 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 01 | 14220 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 01 | 14222 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 01 | 14224 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 01 | 14226 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 01 | 14228 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 01 | 14230 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 01 | 14232 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 01 | 14234 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 01 | 14236 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 01 | 14238 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 01 | 14240 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 01 | 14242 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 01 | 14244 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 01 | 14246 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 01 | 14248 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 01 | 14250 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 01 | 14252 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 01 | 14254 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 01 | 14256 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 01 | 14258 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 01 | 14260 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 01 | 14262 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 01 | 14264 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 01 | 14266 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 01 | 14268 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 01 | 14270 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 01 | 14272 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 01 | 14274 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 01 | 14276 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 01 | 14278 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 01 | 14280 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 01 | 14282 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 01 | 14284 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 01 | 14286 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 01 | 14288 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 01 | 14290 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 01 | 14292 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 01 | 14294 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 01 | 14296 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 01 | 14298 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 01 | 14300 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 01 | 14302 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 01 | 14304 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 01 | 14306 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 01 | 14308 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 01 | 14310 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 01 | 14312 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 01 | 14314 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 01 | 14316 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 01 | 14318 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 01 | 14320 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 01 | 14322 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 01 | 14324 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 01 | 14326 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 01 | 14328 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 01 | 14330 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 01 | 14332 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 01 | 14334 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 01 | 14336 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 01 | 14338 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 01 | 14340 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 01 | 14342 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 02 | 14344 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 02 | 14346 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 02 | 14348 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 02 | 14350 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 02 | 14352 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 02 | 14354 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 02 | 14356 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 02 | 14358 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 02 | 14360 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 02 | 14362 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 02 | 14364 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 02 | 14366 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 02 | 14368 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 02 | 14370 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 02 | 14372 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 02 | 14374 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 02 | 14376 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 02 | 14378 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 02 | 14380 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 02 | 14382 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 02 | 14384 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 02 | 14386 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 02 | 14388 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 02 | 14390 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 02 | 14392 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 02 | 14394 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 02 | 14396 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 02 | 14398 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 02 | 14400 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 02 | 14402 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 02 | 14404 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 02 | 14406 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 02 | 14408 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 02 | 14410 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 02 | 14412 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 02 | 14414 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 02 | 14416 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 02 | 14418 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 02 | 14420 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 02 | 14422 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 02 | 14424 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 02 | 14426 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 02 | 14428 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 02 | 14430 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 02 | 14432 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 02 | 14434 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 02 | 14436 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 02 | 14438 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 02 | 14440 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 02 | 14442 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 02 | 14444 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 02 | 14446 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 02 | 14448 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 02 | 14450 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 02 | 14452 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 02 | 14454 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 02 | 14456 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 02 | 14458 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 02 | 14460 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 02 | 14462 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 02 | 14464 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 02 | 14466 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 02 | 14468 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 02 | 14470 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 02 | 14472 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 02 | 14474 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 02 | 14476 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 02 | 14478 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 02 | 14480 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 02 | 14482 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 02 | 14484 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 02 | 14486 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 02 | 14488 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 02 | 14490 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 02 | 14492 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 02 | 14494 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 02 | 14496 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 02 | 14498 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 02 | 14500 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 02 | 14502 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 02 | 14504 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 02 | 14506 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 02 | 14508 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 02 | 14510 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 02 | 14512 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 02 | 14514 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 02 | 14516 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 02 | 14518 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 02 | 14520 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 02 | 14522 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 02 | 14524 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 02 | 14526 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 02 | 14528 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 02 | 14530 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 02 | 14532 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 02 | 14534 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 02 | 14536 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 02 | 14538 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 02 | 14540 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 02 | 14542 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 02 | 14544 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 02 | 14546 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 02 | 14548 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 02 | 14550 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 02 | 14552 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 02 | 14554 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 02 | 14556 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 02 | 14558 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 02 | 14560 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 02 | 14562 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 02 | 14564 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 02 | 14566 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 02 | 14568 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 02 | 14570 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 02 | 14572 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 02 | 14574 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 02 | 14576 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 02 | 14578 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 02 | 14580 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 02 | 14582 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 03 | 14584 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 03 | 14586 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 03 | 14588 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 03 | 14590 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 03 | 14592 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 03 | 14594 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 03 | 14596 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 03 | 14598 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 03 | 14600 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 03 | 14602 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 03 | 14604 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 03 | 14606 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 03 | 14608 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 03 | 14610 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 03 | 14612 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 03 | 14614 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 03 | 14616 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 03 | 14618 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 03 | 14620 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 03 | 14622 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 03 | 14624 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 03 | 14626 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 03 | 14628 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 03 | 14630 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 03 | 14632 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 03 | 14634 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 03 | 14636 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 03 | 14638 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 03 | 14640 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 03 | 14642 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 03 | 14644 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 03 | 14646 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 03 | 14648 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 03 | 14650 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 03 | 14652 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 03 | 14654 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 03 | 14656 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 03 | 14658 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 03 | 14660 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 03 | 14662 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 03 | 14664 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 03 | 14666 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 03 | 14668 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 03 | 14670 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 03 | 14672 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 03 | 14674 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 03 | 14676 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 03 | 14678 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 03 | 14680 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 03 | 14682 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 03 | 14684 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 03 | 14686 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 03 | 14688 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 03 | 14690 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 03 | 14692 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 03 | 14694 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 03 | 14696 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 03 | 14698 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 03 | 14700 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 03 | 14702 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 03 | 14704 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 03 | 14706 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 03 | 14708 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 03 | 14710 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 03 | 14712 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 03 | 14714 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 03 | 14716 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 03 | 14718 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 03 | 14720 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 03 | 14722 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 03 | 14724 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 03 | 14726 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 03 | 14728 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 03 | 14730 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 03 | 14732 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 03 | 14734 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 03 | 14736 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 03 | 14738 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 03 | 14740 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 03 | 14742 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 03 | 14744 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 03 | 14746 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 03 | 14748 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 03 | 14750 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 03 | 14752 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 03 | 14754 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 03 | 14756 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 03 | 14758 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 03 | 14760 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 03 | 14762 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 03 | 14764 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 03 | 14766 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 03 | 14768 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 03 | 14770 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 03 | 14772 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 03 | 14774 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 03 | 14776 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 03 | 14778 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 03 | 14780 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 03 | 14782 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 03 | 14784 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 03 | 14786 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 03 | 14788 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 03 | 14790 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 03 | 14792 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 03 | 14794 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 03 | 14796 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 03 | 14798 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 03 | 14800 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 03 | 14802 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 03 | 14804 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 03 | 14806 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 03 | 14808 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 03 | 14810 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 03 | 14812 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 03 | 14814 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 03 | 14816 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 03 | 14818 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 03 | 14820 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 03 | 14822 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 04 | 14824 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 04 | 14826 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 04 | 14828 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 04 | 14830 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 04 | 14832 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 04 | 14834 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 04 | 14836 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 04 | 14838 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 04 | 14840 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 04 | 14842 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 04 | 14844 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 04 | 14846 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 04 | 14848 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 04 | 14850 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 04 | 14852 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 04 | 14854 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 04 | 14856 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 04 | 14858 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 04 | 14860 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 04 | 14862 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 04 | 14864 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 04 | 14866 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 04 | 14868 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 04 | 14870 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 04 | 14872 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 04 | 14874 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 04 | 14876 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 04 | 14878 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 04 | 14880 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 04 | 14882 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 04 | 14884 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 04 | 14886 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 04 | 14888 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 04 | 14890 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 04 | 14892 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 04 | 14894 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 04 | 14896 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 04 | 14898 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 04 | 14900 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 04 | 14902 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 04 | 14904 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 04 | 14906 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 04 | 14908 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 04 | 14910 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 04 | 14912 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 04 | 14914 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 04 | 14916 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 04 | 14918 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 04 | 14920 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 04 | 14922 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 04 | 14924 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 04 | 14926 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 04 | 14928 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 04 | 14930 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 04 | 14932 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 04 | 14934 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 04 | 14936 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 04 | 14938 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 04 | 14940 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 04 | 14942 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 04 | 14944 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 04 | 14946 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 04 | 14948 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 04 | 14950 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 04 | 14952 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 04 | 14954 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 04 | 14956 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 04 | 14958 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 04 | 14960 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 04 | 14962 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 04 | 14964 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 04 | 14966 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 04 | 14968 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 04 | 14970 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 04 | 14972 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 04 | 14974 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 04 | 14976 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 04 | 14978 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 04 | 14980 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 04 | 14982 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 04 | 14984 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 04 | 14986 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 04 | 14988 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 04 | 14990 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 04 | 14992 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 04 | 14994 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 04 | 14996 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 04 | 14998 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 04 | 15000 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 04 | 15002 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 04 | 15004 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 04 | 15006 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 04 | 15008 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 04 | 15010 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 04 | 15012 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 04 | 15014 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 04 | 15016 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 04 | 15018 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 04 | 15020 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 04 | 15022 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 04 | 15024 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 04 | 15026 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 04 | 15028 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 04 | 15030 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 04 | 15032 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 04 | 15034 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 04 | 15036 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 04 | 15038 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 04 | 15040 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 04 | 15042 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 04 | 15044 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 04 | 15046 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 04 | 15048 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 04 | 15050 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 04 | 15052 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 04 | 15054 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 04 | 15056 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 04 | 15058 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 04 | 15060 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 04 | 15062 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 05 | 15064 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 05 | 15066 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 05 | 15068 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 05 | 15070 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 05 | 15072 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 05 | 15074 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 05 | 15076 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 05 | 15078 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 05 | 15080 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 05 | 15082 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 05 | 15084 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 05 | 15086 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 05 | 15088 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 05 | 15090 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 05 | 15092 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 05 | 15094 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 05 | 15096 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 05 | 15098 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 05 | 15100 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 05 | 15102 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 05 | 15104 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 05 | 15106 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 05 | 15108 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 05 | 15110 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 05 | 15112 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 05 | 15114 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 05 | 15116 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 05 | 15118 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 05 | 15120 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 05 | 15122 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 05 | 15124 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 05 | 15126 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 05 | 15128 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 05 | 15130 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 05 | 15132 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 05 | 15134 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 05 | 15136 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 05 | 15138 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 05 | 15140 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 05 | 15142 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 05 | 15144 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 05 | 15146 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 05 | 15148 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 05 | 15150 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 05 | 15152 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 05 | 15154 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 05 | 15156 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 05 | 15158 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 05 | 15160 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 05 | 15162 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 05 | 15164 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 05 | 15166 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 05 | 15168 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 05 | 15170 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 05 | 15172 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 05 | 15174 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 05 | 15176 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 05 | 15178 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 05 | 15180 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 05 | 15182 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 05 | 15184 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 05 | 15186 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 05 | 15188 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 05 | 15190 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 05 | 15192 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 05 | 15194 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 05 | 15196 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 05 | 15198 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 05 | 15200 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 05 | 15202 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 05 | 15204 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 05 | 15206 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 05 | 15208 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 05 | 15210 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 05 | 15212 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 05 | 15214 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 05 | 15216 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 05 | 15218 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 05 | 15220 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 05 | 15222 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 05 | 15224 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 05 | 15226 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 05 | 15228 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 05 | 15230 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 05 | 15232 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 05 | 15234 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 05 | 15236 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 05 | 15238 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 05 | 15240 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 05 | 15242 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 05 | 15244 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 05 | 15246 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 05 | 15248 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 05 | 15250 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 05 | 15252 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 05 | 15254 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 05 | 15256 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 05 | 15258 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 05 | 15260 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 05 | 15262 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 05 | 15264 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 05 | 15266 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 05 | 15268 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 05 | 15270 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 05 | 15272 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 05 | 15274 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 05 | 15276 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 05 | 15278 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 05 | 15280 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 05 | 15282 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 05 | 15284 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 05 | 15286 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 05 | 15288 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 05 | 15290 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 05 | 15292 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 05 | 15294 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 05 | 15296 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 05 | 15298 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 05 | 15300 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 05 | 15302 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 06 | 15304 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 06 | 15306 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 06 | 15308 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 06 | 15310 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 06 | 15312 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 06 | 15314 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 06 | 15316 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 06 | 15318 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 06 | 15320 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 06 | 15322 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 06 | 15324 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 06 | 15326 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 06 | 15328 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 06 | 15330 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 06 | 15332 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 06 | 15334 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 06 | 15336 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 06 | 15338 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 06 | 15340 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 06 | 15342 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 06 | 15344 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 06 | 15346 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 06 | 15348 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 06 | 15350 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 06 | 15352 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 06 | 15354 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 06 | 15356 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 06 | 15358 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 06 | 15360 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 06 | 15362 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 06 | 15364 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 06 | 15366 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 06 | 15368 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 06 | 15370 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 06 | 15372 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 06 | 15374 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 06 | 15376 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 06 | 15378 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 06 | 15380 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 06 | 15382 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 06 | 15384 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 06 | 15386 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 06 | 15388 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 06 | 15390 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 06 | 15392 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 06 | 15394 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 06 | 15396 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 06 | 15398 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 06 | 15400 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 06 | 15402 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 06 | 15404 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 06 | 15406 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 06 | 15408 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 06 | 15410 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 06 | 15412 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 06 | 15414 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 06 | 15416 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 06 | 15418 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 06 | 15420 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 06 | 15422 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 06 | 15424 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 06 | 15426 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 06 | 15428 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 06 | 15430 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 06 | 15432 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 06 | 15434 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 06 | 15436 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 06 | 15438 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 06 | 15440 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 06 | 15442 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 06 | 15444 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 06 | 15446 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 06 | 15448 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 06 | 15450 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 06 | 15452 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 06 | 15454 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 06 | 15456 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 06 | 15458 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 06 | 15460 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 06 | 15462 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 06 | 15464 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 06 | 15466 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 06 | 15468 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 06 | 15470 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 06 | 15472 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 06 | 15474 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 06 | 15476 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 06 | 15478 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 06 | 15480 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 06 | 15482 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 06 | 15484 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 06 | 15486 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 06 | 15488 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 06 | 15490 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 06 | 15492 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 06 | 15494 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 06 | 15496 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 06 | 15498 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 06 | 15500 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 06 | 15502 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 06 | 15504 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 06 | 15506 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 06 | 15508 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 06 | 15510 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 06 | 15512 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 06 | 15514 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 06 | 15516 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 06 | 15518 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 06 | 15520 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 06 | 15522 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 06 | 15524 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 06 | 15526 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 06 | 15528 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 06 | 15530 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 06 | 15532 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 06 | 15534 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 06 | 15536 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 06 | 15538 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 06 | 15540 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 06 | 15542 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 07 | 15544 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 07 | 15546 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 07 | 15548 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 07 | 15550 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 07 | 15552 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 07 | 15554 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 07 | 15556 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 07 | 15558 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 07 | 15560 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 07 | 15562 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 07 | 15564 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 07 | 15566 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 07 | 15568 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 07 | 15570 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 07 | 15572 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 07 | 15574 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 07 | 15576 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 07 | 15578 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 07 | 15580 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 07 | 15582 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 07 | 15584 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 07 | 15586 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 07 | 15588 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 07 | 15590 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 07 | 15592 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 07 | 15594 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 07 | 15596 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 07 | 15598 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 07 | 15600 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 07 | 15602 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 07 | 15604 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 07 | 15606 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 07 | 15608 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 07 | 15610 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 07 | 15612 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 07 | 15614 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 07 | 15616 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 07 | 15618 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 07 | 15620 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 07 | 15622 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 07 | 15624 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 07 | 15626 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 07 | 15628 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 07 | 15630 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 07 | 15632 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 07 | 15634 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 07 | 15636 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 07 | 15638 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 07 | 15640 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 07 | 15642 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 07 | 15644 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 07 | 15646 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 07 | 15648 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 07 | 15650 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 07 | 15652 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 07 | 15654 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 07 | 15656 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 07 | 15658 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 07 | 15660 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 07 | 15662 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 07 | 15664 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 07 | 15666 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 07 | 15668 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 07 | 15670 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 07 | 15672 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 07 | 15674 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 07 | 15676 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 07 | 15678 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 07 | 15680 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 07 | 15682 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 07 | 15684 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 07 | 15686 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 07 | 15688 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 07 | 15690 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 07 | 15692 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 07 | 15694 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 07 | 15696 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 07 | 15698 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 07 | 15700 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 07 | 15702 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 07 | 15704 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 07 | 15706 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 07 | 15708 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 07 | 15710 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 07 | 15712 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 07 | 15714 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 07 | 15716 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 07 | 15718 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 07 | 15720 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 07 | 15722 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 07 | 15724 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 07 | 15726 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 07 | 15728 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 07 | 15730 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 07 | 15732 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 07 | 15734 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 07 | 15736 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 07 | 15738 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 07 | 15740 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 07 | 15742 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 07 | 15744 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 07 | 15746 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 07 | 15748 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 07 | 15750 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 07 | 15752 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 07 | 15754 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 07 | 15756 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 07 | 15758 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 07 | 15760 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 07 | 15762 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 07 | 15764 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 07 | 15766 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 07 | 15768 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 07 | 15770 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 07 | 15772 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 07 | 15774 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 07 | 15776 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 07 | 15778 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 07 | 15780 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 07 | 15782 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 08 | 15784 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 08 | 15786 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 08 | 15788 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 08 | 15790 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 08 | 15792 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 08 | 15794 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 08 | 15796 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 08 | 15798 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 08 | 15800 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 08 | 15802 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 08 | 15804 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 08 | 15806 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 08 | 15808 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 08 | 15810 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 08 | 15812 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 08 | 15814 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 08 | 15816 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 08 | 15818 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 08 | 15820 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 08 | 15822 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 08 | 15824 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 08 | 15826 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 08 | 15828 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 08 | 15830 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 08 | 15832 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 08 | 15834 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 08 | 15836 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 08 | 15838 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 08 | 15840 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 08 | 15842 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 08 | 15844 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 08 | 15846 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 08 | 15848 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 08 | 15850 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 08 | 15852 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 08 | 15854 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 08 | 15856 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 08 | 15858 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 08 | 15860 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 08 | 15862 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 08 | 15864 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 08 | 15866 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 08 | 15868 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 08 | 15870 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 08 | 15872 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 08 | 15874 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 08 | 15876 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 08 | 15878 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 08 | 15880 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 08 | 15882 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 08 | 15884 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 08 | 15886 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 08 | 15888 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 08 | 15890 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 08 | 15892 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 08 | 15894 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 08 | 15896 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 08 | 15898 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 08 | 15900 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 08 | 15902 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 08 | 15904 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 08 | 15906 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 08 | 15908 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 08 | 15910 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 08 | 15912 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 08 | 15914 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 08 | 15916 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 08 | 15918 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 08 | 15920 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 08 | 15922 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 08 | 15924 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 08 | 15926 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 08 | 15928 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 08 | 15930 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 08 | 15932 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 08 | 15934 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 08 | 15936 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 08 | 15938 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 08 | 15940 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 08 | 15942 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 08 | 15944 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 08 | 15946 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 08 | 15948 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 08 | 15950 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 08 | 15952 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 08 | 15954 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 08 | 15956 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 08 | 15958 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 08 | 15960 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 08 | 15962 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 08 | 15964 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 08 | 15966 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 08 | 15968 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 08 | 15970 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 08 | 15972 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 08 | 15974 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 08 | 15976 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 08 | 15978 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 08 | 15980 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 08 | 15982 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 08 | 15984 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 08 | 15986 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 08 | 15988 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 08 | 15990 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 08 | 15992 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 08 | 15994 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 08 | 15996 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 08 | 15998 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 08 | 16000 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 08 | 16002 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 08 | 16004 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 08 | 16006 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 08 | 16008 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 08 | 16010 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 08 | 16012 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 08 | 16014 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 08 | 16016 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 08 | 16018 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 08 | 16020 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 08 | 16022 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 09 | 16024 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 09 | 16026 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 09 | 16028 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 09 | 16030 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 09 | 16032 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 09 | 16034 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 09 | 16036 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 09 | 16038 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 09 | 16040 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 09 | 16042 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 09 | 16044 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 09 | 16046 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 09 | 16048 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 09 | 16050 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 09 | 16052 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 09 | 16054 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 09 | 16056 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 09 | 16058 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 09 | 16060 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 09 | 16062 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 09 | 16064 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 09 | 16066 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 09 | 16068 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 09 | 16070 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 09 | 16072 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 09 | 16074 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 09 | 16076 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 09 | 16078 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 09 | 16080 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 09 | 16082 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 09 | 16084 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 09 | 16086 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 09 | 16088 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 09 | 16090 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 09 | 16092 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 09 | 16094 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 09 | 16096 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 09 | 16098 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 09 | 16100 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 09 | 16102 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 09 | 16104 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 09 | 16106 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 09 | 16108 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 09 | 16110 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 09 | 16112 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 09 | 16114 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 09 | 16116 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 09 | 16118 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 09 | 16120 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 09 | 16122 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 09 | 16124 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 09 | 16126 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 09 | 16128 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 09 | 16130 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 09 | 16132 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 09 | 16134 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 09 | 16136 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 09 | 16138 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 09 | 16140 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 09 | 16142 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 09 | 16144 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 09 | 16146 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 09 | 16148 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 09 | 16150 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 09 | 16152 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 09 | 16154 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 09 | 16156 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 09 | 16158 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 09 | 16160 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 09 | 16162 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 09 | 16164 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 09 | 16166 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 09 | 16168 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 09 | 16170 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 09 | 16172 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 09 | 16174 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 09 | 16176 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 09 | 16178 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 09 | 16180 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 09 | 16182 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 09 | 16184 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 09 | 16186 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 09 | 16188 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 09 | 16190 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 09 | 16192 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 09 | 16194 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 09 | 16196 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 09 | 16198 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 09 | 16200 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 09 | 16202 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 09 | 16204 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 09 | 16206 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 09 | 16208 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 09 | 16210 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 09 | 16212 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 09 | 16214 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 09 | 16216 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 09 | 16218 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 09 | 16220 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 09 | 16222 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 09 | 16224 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 09 | 16226 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 09 | 16228 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 09 | 16230 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 09 | 16232 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 09 | 16234 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 09 | 16236 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 09 | 16238 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 09 | 16240 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 09 | 16242 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 09 | 16244 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 09 | 16246 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 09 | 16248 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 09 | 16250 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 09 | 16252 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 09 | 16254 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 09 | 16256 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 09 | 16258 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 09 | 16260 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 09 | 16262 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 10 | 16264 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 10 | 16266 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 10 | 16268 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 10 | 16270 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 10 | 16272 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 10 | 16274 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 10 | 16276 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 10 | 16278 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 10 | 16280 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 10 | 16282 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 10 | 16284 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 10 | 16286 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 10 | 16288 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 10 | 16290 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 10 | 16292 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 10 | 16294 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 10 | 16296 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 10 | 16298 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 10 | 16300 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 10 | 16302 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 10 | 16304 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 10 | 16306 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 10 | 16308 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 10 | 16310 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 10 | 16312 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 10 | 16314 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 10 | 16316 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 10 | 16318 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 10 | 16320 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 10 | 16322 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 10 | 16324 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 10 | 16326 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 10 | 16328 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 10 | 16330 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 10 | 16332 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 10 | 16334 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 10 | 16336 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 10 | 16338 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 10 | 16340 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 10 | 16342 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 10 | 16344 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 10 | 16346 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 10 | 16348 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 10 | 16350 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 10 | 16352 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 10 | 16354 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 10 | 16356 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 10 | 16358 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 10 | 16360 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 10 | 16362 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 10 | 16364 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 10 | 16366 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 10 | 16368 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 10 | 16370 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 10 | 16372 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 10 | 16374 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 10 | 16376 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 10 | 16378 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 10 | 16380 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 10 | 16382 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 10 | 16384 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 10 | 16386 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 10 | 16388 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 10 | 16390 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 10 | 16392 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 10 | 16394 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 10 | 16396 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 10 | 16398 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 10 | 16400 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 10 | 16402 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 10 | 16404 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 10 | 16406 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 10 | 16408 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 10 | 16410 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 10 | 16412 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 10 | 16414 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 10 | 16416 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 10 | 16418 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 10 | 16420 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 10 | 16422 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 10 | 16424 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 10 | 16426 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 10 | 16428 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 10 | 16430 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 10 | 16432 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 10 | 16434 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 10 | 16436 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 10 | 16438 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 10 | 16440 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 10 | 16442 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 10 | 16444 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 10 | 16446 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 10 | 16448 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 10 | 16450 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 10 | 16452 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 10 | 16454 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 10 | 16456 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 10 | 16458 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 10 | 16460 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 10 | 16462 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 10 | 16464 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 10 | 16466 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 10 | 16468 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 10 | 16470 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 10 | 16472 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 10 | 16474 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 10 | 16476 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 10 | 16478 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 10 | 16480 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 10 | 16482 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 10 | 16484 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 10 | 16486 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 10 | 16488 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 10 | 16490 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 10 | 16492 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 10 | 16494 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 10 | 16496 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 10 | 16498 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 10 | 16500 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 10 | 16502 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 11 | 16504 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 11 | 16506 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 11 | 16508 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 11 | 16510 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 11 | 16512 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 11 | 16514 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 11 | 16516 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 11 | 16518 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 11 | 16520 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 11 | 16522 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 11 | 16524 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 11 | 16526 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 11 | 16528 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 11 | 16530 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 11 | 16532 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 11 | 16534 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 11 | 16536 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 11 | 16538 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 11 | 16540 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 11 | 16542 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 11 | 16544 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 11 | 16546 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 11 | 16548 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 11 | 16550 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 11 | 16552 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 11 | 16554 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 11 | 16556 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 11 | 16558 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 11 | 16560 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 11 | 16562 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 11 | 16564 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 11 | 16566 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 11 | 16568 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 11 | 16570 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 11 | 16572 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 11 | 16574 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 11 | 16576 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 11 | 16578 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 11 | 16580 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 11 | 16582 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 11 | 16584 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 11 | 16586 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 11 | 16588 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 11 | 16590 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 11 | 16592 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 11 | 16594 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 11 | 16596 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 11 | 16598 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 11 | 16600 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 11 | 16602 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 11 | 16604 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 11 | 16606 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 11 | 16608 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 11 | 16610 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 11 | 16612 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 11 | 16614 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 11 | 16616 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 11 | 16618 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 11 | 16620 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 11 | 16622 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 11 | 16624 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 11 | 16626 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 11 | 16628 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 11 | 16630 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 11 | 16632 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 11 | 16634 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 11 | 16636 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 11 | 16638 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 11 | 16640 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 11 | 16642 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 11 | 16644 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 11 | 16646 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 11 | 16648 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 11 | 16650 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 11 | 16652 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 11 | 16654 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 11 | 16656 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 11 | 16658 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 11 | 16660 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 11 | 16662 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 11 | 16664 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 11 | 16666 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 11 | 16668 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 11 | 16670 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 11 | 16672 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 11 | 16674 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 11 | 16676 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 11 | 16678 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 11 | 16680 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 11 | 16682 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 11 | 16684 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 11 | 16686 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 11 | 16688 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 11 | 16690 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 11 | 16692 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 11 | 16694 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 11 | 16696 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 11 | 16698 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 11 | 16700 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 11 | 16702 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 11 | 16704 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 11 | 16706 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 11 | 16708 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 11 | 16710 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 11 | 16712 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 11 | 16714 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 11 | 16716 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 11 | 16718 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 11 | 16720 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 11 | 16722 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 11 | 16724 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 11 | 16726 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 11 | 16728 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 11 | 16730 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 11 | 16732 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 11 | 16734 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 11 | 16736 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 11 | 16738 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 11 | 16740 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 11 | 16742 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 12 | 16744 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 12 | 16746 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 12 | 16748 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 12 | 16750 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 12 | 16752 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 12 | 16754 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 12 | 16756 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 12 | 16758 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 12 | 16760 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 12 | 16762 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 12 | 16764 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 12 | 16766 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 12 | 16768 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 12 | 16770 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 12 | 16772 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 12 | 16774 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 12 | 16776 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 12 | 16778 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 12 | 16780 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 12 | 16782 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 12 | 16784 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 12 | 16786 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 12 | 16788 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 12 | 16790 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 12 | 16792 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 12 | 16794 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 12 | 16796 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 12 | 16798 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 12 | 16800 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 12 | 16802 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 12 | 16804 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 12 | 16806 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 12 | 16808 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 12 | 16810 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 12 | 16812 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 12 | 16814 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 12 | 16816 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 12 | 16818 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 12 | 16820 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 12 | 16822 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 12 | 16824 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 12 | 16826 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 12 | 16828 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 12 | 16830 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 12 | 16832 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 12 | 16834 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 12 | 16836 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 12 | 16838 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 12 | 16840 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 12 | 16842 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 12 | 16844 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 12 | 16846 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 12 | 16848 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 12 | 16850 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 12 | 16852 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 12 | 16854 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 12 | 16856 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 12 | 16858 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 12 | 16860 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 12 | 16862 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 12 | 16864 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 12 | 16866 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 12 | 16868 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 12 | 16870 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 12 | 16872 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 12 | 16874 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 12 | 16876 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 12 | 16878 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 12 | 16880 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 12 | 16882 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 12 | 16884 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 12 | 16886 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 12 | 16888 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 12 | 16890 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 12 | 16892 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 12 | 16894 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 12 | 16896 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 12 | 16898 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 12 | 16900 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 12 | 16902 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 12 | 16904 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 12 | 16906 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 12 | 16908 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 12 | 16910 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 12 | 16912 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 12 | 16914 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 12 | 16916 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 12 | 16918 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 12 | 16920 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 12 | 16922 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 12 | 16924 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 12 | 16926 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 12 | 16928 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 12 | 16930 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 12 | 16932 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 12 | 16934 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 12 | 16936 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 12 | 16938 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 12 | 16940 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 12 | 16942 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 12 | 16944 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 12 | 16946 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 12 | 16948 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 12 | 16950 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 12 | 16952 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 12 | 16954 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 12 | 16956 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 12 | 16958 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 12 | 16960 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 12 | 16962 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 12 | 16964 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 12 | 16966 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 12 | 16968 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 12 | 16970 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 12 | 16972 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 12 | 16974 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 12 | 16976 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 12 | 16978 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 12 | 16980 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 12 | 16982 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 13 | 16984 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 13 | 16986 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 13 | 16988 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 13 | 16990 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 13 | 16992 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 13 | 16994 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 13 | 16996 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 13 | 16998 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 13 | 17000 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 13 | 17002 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 13 | 17004 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 13 | 17006 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 13 | 17008 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 13 | 17010 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 13 | 17012 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 13 | 17014 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 13 | 17016 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 13 | 17018 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 13 | 17020 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 13 | 17022 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 13 | 17024 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 13 | 17026 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 13 | 17028 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 13 | 17030 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 13 | 17032 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 13 | 17034 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 13 | 17036 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 13 | 17038 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 13 | 17040 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 13 | 17042 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 13 | 17044 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 13 | 17046 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 13 | 17048 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 13 | 17050 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 13 | 17052 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 13 | 17054 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 13 | 17056 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 13 | 17058 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 13 | 17060 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 13 | 17062 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 13 | 17064 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 13 | 17066 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 13 | 17068 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 13 | 17070 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 13 | 17072 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 13 | 17074 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 13 | 17076 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 13 | 17078 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 13 | 17080 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 13 | 17082 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 13 | 17084 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 13 | 17086 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 13 | 17088 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 13 | 17090 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 13 | 17092 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 13 | 17094 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 13 | 17096 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 13 | 17098 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 13 | 17100 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 13 | 17102 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 13 | 17104 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 13 | 17106 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 13 | 17108 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 13 | 17110 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 13 | 17112 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 13 | 17114 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 13 | 17116 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 13 | 17118 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 13 | 17120 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 13 | 17122 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 13 | 17124 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 13 | 17126 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 13 | 17128 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 13 | 17130 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 13 | 17132 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 13 | 17134 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 13 | 17136 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 13 | 17138 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 13 | 17140 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 13 | 17142 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 13 | 17144 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 13 | 17146 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 13 | 17148 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 13 | 17150 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 13 | 17152 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 13 | 17154 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 13 | 17156 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 13 | 17158 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 13 | 17160 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 13 | 17162 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 13 | 17164 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 13 | 17166 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 13 | 17168 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 13 | 17170 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 13 | 17172 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 13 | 17174 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 13 | 17176 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 13 | 17178 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 13 | 17180 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 13 | 17182 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 13 | 17184 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 13 | 17186 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 13 | 17188 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 13 | 17190 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 13 | 17192 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 13 | 17194 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 13 | 17196 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 13 | 17198 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 13 | 17200 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 13 | 17202 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 13 | 17204 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 13 | 17206 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 13 | 17208 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 13 | 17210 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 13 | 17212 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 13 | 17214 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 13 | 17216 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 13 | 17218 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 13 | 17220 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 13 | 17222 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 14 | 17224 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 14 | 17226 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 14 | 17228 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 14 | 17230 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 14 | 17232 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 14 | 17234 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 14 | 17236 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 14 | 17238 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 14 | 17240 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 14 | 17242 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 14 | 17244 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 14 | 17246 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 14 | 17248 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 14 | 17250 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 14 | 17252 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 14 | 17254 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 14 | 17256 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 14 | 17258 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 14 | 17260 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 14 | 17262 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 14 | 17264 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 14 | 17266 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 14 | 17268 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 14 | 17270 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 14 | 17272 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 14 | 17274 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 14 | 17276 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 14 | 17278 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 14 | 17280 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 14 | 17282 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 14 | 17284 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 14 | 17286 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 14 | 17288 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 14 | 17290 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 14 | 17292 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 14 | 17294 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 14 | 17296 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 14 | 17298 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 14 | 17300 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 14 | 17302 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 14 | 17304 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 14 | 17306 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 14 | 17308 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 14 | 17310 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 14 | 17312 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 14 | 17314 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 14 | 17316 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 14 | 17318 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 14 | 17320 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 14 | 17322 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 14 | 17324 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 14 | 17326 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 14 | 17328 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 14 | 17330 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 14 | 17332 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 14 | 17334 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 14 | 17336 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 14 | 17338 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 14 | 17340 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 14 | 17342 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 14 | 17344 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 14 | 17346 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 14 | 17348 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 14 | 17350 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 14 | 17352 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 14 | 17354 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 14 | 17356 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 14 | 17358 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 14 | 17360 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 14 | 17362 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 14 | 17364 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 14 | 17366 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 14 | 17368 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 14 | 17370 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 14 | 17372 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 14 | 17374 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 14 | 17376 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 14 | 17378 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 14 | 17380 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 14 | 17382 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 14 | 17384 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 14 | 17386 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 14 | 17388 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 14 | 17390 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 14 | 17392 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 14 | 17394 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 14 | 17396 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 14 | 17398 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 14 | 17400 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 14 | 17402 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 14 | 17404 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 14 | 17406 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 14 | 17408 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 14 | 17410 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 14 | 17412 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 14 | 17414 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 14 | 17416 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 14 | 17418 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 14 | 17420 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 14 | 17422 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 14 | 17424 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 14 | 17426 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 14 | 17428 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 14 | 17430 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 14 | 17432 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 14 | 17434 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 14 | 17436 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 14 | 17438 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 14 | 17440 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 14 | 17442 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 14 | 17444 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 14 | 17446 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 14 | 17448 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 14 | 17450 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 14 | 17452 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 14 | 17454 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 14 | 17456 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 14 | 17458 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 14 | 17460 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 14 | 17462 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 15 | 17464 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 15 | 17466 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 15 | 17468 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 15 | 17470 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 15 | 17472 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 15 | 17474 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 15 | 17476 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 15 | 17478 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 15 | 17480 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 15 | 17482 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 15 | 17484 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 15 | 17486 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 15 | 17488 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 15 | 17490 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 15 | 17492 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 15 | 17494 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 15 | 17496 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 15 | 17498 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 15 | 17500 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 15 | 17502 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 15 | 17504 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 15 | 17506 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 15 | 17508 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 15 | 17510 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 15 | 17512 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 15 | 17514 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 15 | 17516 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 15 | 17518 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 15 | 17520 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 15 | 17522 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 15 | 17524 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 15 | 17526 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 15 | 17528 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 15 | 17530 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 15 | 17532 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 15 | 17534 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 15 | 17536 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 15 | 17538 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 15 | 17540 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 15 | 17542 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 15 | 17544 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 15 | 17546 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 15 | 17548 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 15 | 17550 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 15 | 17552 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 15 | 17554 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 15 | 17556 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 15 | 17558 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 15 | 17560 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 15 | 17562 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 15 | 17564 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 15 | 17566 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 15 | 17568 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 15 | 17570 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 15 | 17572 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 15 | 17574 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 15 | 17576 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 15 | 17578 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 15 | 17580 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 15 | 17582 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 15 | 17584 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 15 | 17586 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 15 | 17588 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 15 | 17590 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 15 | 17592 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 15 | 17594 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 15 | 17596 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 15 | 17598 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 15 | 17600 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 15 | 17602 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 15 | 17604 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 15 | 17606 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 15 | 17608 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 15 | 17610 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 15 | 17612 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 15 | 17614 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 15 | 17616 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 15 | 17618 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 15 | 17620 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 15 | 17622 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 15 | 17624 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 15 | 17626 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 15 | 17628 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 15 | 17630 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 15 | 17632 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 15 | 17634 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 15 | 17636 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 15 | 17638 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 15 | 17640 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 15 | 17642 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 15 | 17644 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 15 | 17646 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 15 | 17648 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 15 | 17650 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 15 | 17652 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 15 | 17654 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 15 | 17656 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 15 | 17658 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 15 | 17660 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 15 | 17662 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 15 | 17664 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 15 | 17666 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 15 | 17668 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 15 | 17670 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 15 | 17672 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 15 | 17674 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 15 | 17676 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 15 | 17678 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 15 | 17680 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 15 | 17682 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 15 | 17684 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 15 | 17686 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 15 | 17688 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 15 | 17690 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 15 | 17692 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 15 | 17694 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 15 | 17696 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 15 | 17698 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 15 | 17700 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 15 | 17702 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 16 | 17704 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 16 | 17706 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 16 | 17708 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 16 | 17710 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 16 | 17712 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 16 | 17714 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 16 | 17716 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 16 | 17718 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 16 | 17720 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 16 | 17722 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 16 | 17724 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 16 | 17726 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 16 | 17728 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 16 | 17730 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 16 | 17732 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 16 | 17734 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 16 | 17736 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 16 | 17738 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 16 | 17740 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 16 | 17742 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 16 | 17744 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 16 | 17746 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 16 | 17748 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 16 | 17750 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 16 | 17752 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 16 | 17754 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 16 | 17756 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 16 | 17758 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 16 | 17760 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 16 | 17762 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 16 | 17764 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 16 | 17766 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 16 | 17768 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 16 | 17770 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 16 | 17772 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 16 | 17774 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 16 | 17776 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 16 | 17778 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 16 | 17780 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 16 | 17782 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 16 | 17784 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 16 | 17786 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 16 | 17788 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 16 | 17790 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 16 | 17792 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 16 | 17794 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 16 | 17796 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 16 | 17798 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 16 | 17800 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 16 | 17802 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 16 | 17804 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 16 | 17806 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 16 | 17808 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 16 | 17810 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 16 | 17812 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 16 | 17814 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 16 | 17816 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 16 | 17818 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 16 | 17820 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 16 | 17822 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 16 | 17824 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 16 | 17826 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 16 | 17828 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 16 | 17830 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 16 | 17832 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 16 | 17834 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 16 | 17836 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 16 | 17838 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 16 | 17840 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 16 | 17842 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 16 | 17844 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 16 | 17846 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 16 | 17848 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 16 | 17850 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 16 | 17852 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 16 | 17854 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 16 | 17856 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 16 | 17858 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 16 | 17860 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 16 | 17862 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 16 | 17864 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 16 | 17866 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 16 | 17868 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 16 | 17870 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 16 | 17872 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 16 | 17874 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 16 | 17876 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 16 | 17878 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 16 | 17880 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 16 | 17882 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 16 | 17884 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 16 | 17886 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 16 | 17888 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 16 | 17890 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 16 | 17892 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 16 | 17894 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 16 | 17896 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 16 | 17898 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 16 | 17900 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 16 | 17902 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 16 | 17904 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 16 | 17906 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 16 | 17908 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 16 | 17910 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 16 | 17912 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 16 | 17914 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 16 | 17916 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 16 | 17918 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 16 | 17920 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 16 | 17922 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 16 | 17924 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 16 | 17926 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 16 | 17928 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 16 | 17930 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 16 | 17932 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 16 | 17934 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 16 | 17936 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 16 | 17938 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 16 | 17940 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 16 | 17942 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 17 | 17944 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 17 | 17946 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 17 | 17948 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 17 | 17950 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 17 | 17952 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 17 | 17954 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 17 | 17956 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 17 | 17958 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 17 | 17960 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 17 | 17962 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 17 | 17964 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 17 | 17966 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 17 | 17968 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 17 | 17970 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 17 | 17972 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 17 | 17974 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 17 | 17976 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 17 | 17978 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 17 | 17980 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 17 | 17982 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 17 | 17984 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 17 | 17986 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 17 | 17988 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 17 | 17990 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 17 | 17992 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 17 | 17994 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 17 | 17996 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 17 | 17998 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 17 | 18000 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 17 | 18002 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 17 | 18004 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 17 | 18006 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 17 | 18008 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 17 | 18010 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 17 | 18012 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 17 | 18014 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 17 | 18016 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 17 | 18018 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 17 | 18020 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 17 | 18022 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 17 | 18024 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 17 | 18026 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 17 | 18028 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 17 | 18030 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 17 | 18032 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 17 | 18034 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 17 | 18036 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 17 | 18038 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 17 | 18040 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 17 | 18042 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 17 | 18044 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 17 | 18046 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 17 | 18048 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 17 | 18050 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 17 | 18052 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 17 | 18054 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 17 | 18056 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 17 | 18058 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 17 | 18060 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 17 | 18062 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 17 | 18064 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 17 | 18066 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 17 | 18068 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 17 | 18070 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 17 | 18072 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 17 | 18074 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 17 | 18076 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 17 | 18078 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 17 | 18080 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 17 | 18082 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 17 | 18084 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 17 | 18086 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 17 | 18088 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 17 | 18090 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 17 | 18092 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 17 | 18094 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 17 | 18096 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 17 | 18098 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 17 | 18100 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 17 | 18102 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 17 | 18104 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 17 | 18106 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 17 | 18108 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 17 | 18110 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 17 | 18112 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 17 | 18114 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 17 | 18116 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 17 | 18118 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 17 | 18120 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 17 | 18122 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 17 | 18124 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 17 | 18126 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 17 | 18128 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 17 | 18130 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 17 | 18132 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 17 | 18134 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 17 | 18136 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 17 | 18138 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 17 | 18140 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 17 | 18142 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 17 | 18144 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 17 | 18146 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 17 | 18148 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 17 | 18150 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 17 | 18152 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 17 | 18154 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 17 | 18156 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 17 | 18158 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 17 | 18160 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 17 | 18162 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 17 | 18164 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 17 | 18166 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 17 | 18168 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 17 | 18170 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 17 | 18172 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 17 | 18174 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 17 | 18176 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 17 | 18178 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 17 | 18180 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 17 | 18182 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 18 | 18184 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 18 | 18186 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 18 | 18188 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 18 | 18190 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 18 | 18192 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 18 | 18194 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 18 | 18196 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 18 | 18198 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 18 | 18200 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 18 | 18202 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 18 | 18204 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 18 | 18206 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 18 | 18208 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 18 | 18210 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 18 | 18212 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 18 | 18214 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 18 | 18216 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 18 | 18218 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 18 | 18220 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 18 | 18222 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 18 | 18224 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 18 | 18226 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 18 | 18228 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 18 | 18230 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 18 | 18232 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 18 | 18234 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 18 | 18236 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 18 | 18238 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 18 | 18240 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 18 | 18242 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 18 | 18244 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 18 | 18246 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 18 | 18248 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 18 | 18250 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 18 | 18252 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 18 | 18254 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 18 | 18256 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 18 | 18258 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 18 | 18260 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 18 | 18262 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 18 | 18264 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 18 | 18266 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 18 | 18268 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 18 | 18270 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 18 | 18272 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 18 | 18274 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 18 | 18276 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 18 | 18278 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 18 | 18280 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 18 | 18282 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 18 | 18284 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 18 | 18286 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 18 | 18288 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 18 | 18290 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 18 | 18292 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 18 | 18294 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 18 | 18296 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 18 | 18298 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 18 | 18300 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 18 | 18302 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 18 | 18304 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 18 | 18306 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 18 | 18308 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 18 | 18310 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 18 | 18312 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 18 | 18314 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 18 | 18316 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 18 | 18318 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 18 | 18320 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 18 | 18322 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 18 | 18324 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 18 | 18326 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 18 | 18328 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 18 | 18330 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 18 | 18332 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 18 | 18334 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 18 | 18336 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 18 | 18338 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 18 | 18340 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 18 | 18342 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 18 | 18344 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 18 | 18346 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 18 | 18348 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 18 | 18350 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 18 | 18352 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 18 | 18354 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 18 | 18356 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 18 | 18358 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 18 | 18360 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 18 | 18362 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 18 | 18364 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 18 | 18366 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 18 | 18368 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 18 | 18370 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 18 | 18372 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 18 | 18374 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 18 | 18376 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 18 | 18378 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 18 | 18380 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 18 | 18382 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 18 | 18384 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 18 | 18386 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 18 | 18388 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 18 | 18390 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 18 | 18392 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 18 | 18394 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 18 | 18396 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 18 | 18398 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 18 | 18400 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 18 | 18402 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 18 | 18404 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 18 | 18406 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 18 | 18408 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 18 | 18410 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 18 | 18412 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 18 | 18414 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 18 | 18416 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 18 | 18418 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 18 | 18420 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 18 | 18422 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 19 | 18424 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 19 | 18426 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 19 | 18428 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 19 | 18430 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 19 | 18432 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 19 | 18434 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 19 | 18436 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 19 | 18438 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 19 | 18440 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 19 | 18442 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 19 | 18444 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 19 | 18446 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 19 | 18448 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 19 | 18450 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 19 | 18452 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 19 | 18454 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 19 | 18456 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 19 | 18458 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 19 | 18460 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 19 | 18462 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 19 | 18464 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 19 | 18466 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 19 | 18468 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 19 | 18470 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 19 | 18472 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 19 | 18474 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 19 | 18476 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 19 | 18478 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 19 | 18480 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 19 | 18482 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 19 | 18484 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 19 | 18486 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 19 | 18488 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 19 | 18490 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 19 | 18492 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 19 | 18494 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 19 | 18496 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 19 | 18498 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 19 | 18500 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 19 | 18502 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 19 | 18504 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 19 | 18506 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 19 | 18508 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 19 | 18510 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 19 | 18512 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 19 | 18514 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 19 | 18516 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 19 | 18518 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 19 | 18520 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 19 | 18522 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 19 | 18524 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 19 | 18526 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 19 | 18528 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 19 | 18530 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 19 | 18532 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 19 | 18534 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 19 | 18536 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 19 | 18538 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 19 | 18540 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 19 | 18542 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 19 | 18544 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 19 | 18546 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 19 | 18548 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 19 | 18550 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 19 | 18552 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 19 | 18554 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 19 | 18556 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 19 | 18558 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 19 | 18560 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 19 | 18562 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 19 | 18564 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 19 | 18566 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 19 | 18568 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 19 | 18570 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 19 | 18572 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 19 | 18574 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 19 | 18576 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 19 | 18578 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 19 | 18580 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 19 | 18582 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 19 | 18584 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 19 | 18586 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 19 | 18588 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 19 | 18590 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 19 | 18592 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 19 | 18594 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 19 | 18596 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 19 | 18598 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 19 | 18600 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 19 | 18602 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 19 | 18604 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 19 | 18606 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 19 | 18608 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 19 | 18610 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 19 | 18612 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 19 | 18614 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 19 | 18616 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 19 | 18618 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 19 | 18620 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 19 | 18622 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 19 | 18624 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 19 | 18626 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 19 | 18628 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 19 | 18630 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 19 | 18632 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 19 | 18634 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 19 | 18636 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 19 | 18638 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 19 | 18640 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 19 | 18642 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 19 | 18644 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 19 | 18646 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 19 | 18648 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 19 | 18650 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 19 | 18652 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 19 | 18654 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 19 | 18656 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 19 | 18658 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 19 | 18660 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 19 | 18662 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 20 | 18664 | Float | RD | A | Maximum Value L1 Current (RMS)
Current measurement module group 20 | 18666 | Float | RD | A | Maximum Value L2 Current (RMS)
Current measurement module group 20 | 18668 | Float | RD | A | Maximum Value L3 Current (RMS)
Current measurement module group 20 | 18670 | Float | RD | A | Maximum Value L4 Current (RMS)
Current measurement module group 20 | 18672 | Float | RD | A | Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 20 | 18674 | Float | RD | W | Maximum Value L1 Active Power
Current measurement module group 20 | 18676 | Float | RD | W | Maximum Value L2 Active Power
Current measurement module group 20 | 18678 | Float | RD | W | Maximum Value L3 Active Power
Current measurement module group 20 | 18680 | Float | RD | W | Maximum Value L4 Active Power
Current measurement module group 20 | 18682 | Float | RD | W | Maximum Value Sum L1..L4 Active Power
Current measurement module group 20 | 18684 | Float | RD | var | Maximum Value L1 Reactive Power
Current measurement module group 20 | 18686 | Float | RD | var | Maximum Value L2 Reactive Power
Current measurement module group 20 | 18688 | Float | RD | var | Maximum Value L3 Reactive Power
Current measurement module group 20 | 18690 | Float | RD | var | Maximum Value L4 Reactive Power
Current measurement module group 20 | 18692 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 20 | 18694 | Float | RD | VA | Maximum Value L1 Apparent Power
Current measurement module group 20 | 18696 | Float | RD | VA | Maximum Value L2 Apparent Power
Current measurement module group 20 | 18698 | Float | RD | VA | Maximum Value L3 Apparent Power
Current measurement module group 20 | 18700 | Float | RD | VA | Maximum Value L4 Apparent Power
Current measurement module group 20 | 18702 | Float | RD | VA | Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 20 | 18704 | Float | RD | - | Maximum Value L1 Cos(phi)
Current measurement module group 20 | 18706 | Float | RD | - | Maximum Value L2 Cos(phi)
Current measurement module group 20 | 18708 | Float | RD | - | Maximum Value L3 Cos(phi)
Current measurement module group 20 | 18710 | Float | RD | - | Maximum Value L4 Cos(phi)
Current measurement module group 20 | 18712 | Float | RD | - | Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 20 | 18714 | Float | RD | - | Maximum Value L1 Power Factor
Current measurement module group 20 | 18716 | Float | RD | - | Maximum Value L2 Power Factor
Current measurement module group 20 | 18718 | Float | RD | - | Maximum Value L3 Power Factor
Current measurement module group 20 | 18720 | Float | RD | - | Maximum Value L4 Power Factor
Current measurement module group 20 | 18722 | Float | RD | - | Maximum Value Sum L1..L4 Power Factor
Current measurement module group 20 | 18724 | Float | RD | W | Maximum Value L1 Active Power (fundamental)
Current measurement module group 20 | 18726 | Float | RD | W | Maximum Value L2 Active Power (fundamental)
Current measurement module group 20 | 18728 | Float | RD | W | Maximum Value L3 Active Power (fundamental)
Current measurement module group 20 | 18730 | Float | RD | W | Maximum Value L4 Active Power (fundamental)
Current measurement module group 20 | 18732 | Float | RD | W | Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 20 | 18734 | Float | RD | var | Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 20 | 18736 | Float | RD | var | Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 20 | 18738 | Float | RD | var | Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 20 | 18740 | Float | RD | var | Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 20 | 18742 | Float | RD | var | Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 20 | 18744 | Float | RD | var | Maximum Value L1 Distortion Reactive Power
Current measurement module group 20 | 18746 | Float | RD | var | Maximum Value L2 Distortion Reactive Power
Current measurement module group 20 | 18748 | Float | RD | var | Maximum Value L3 Distortion Reactive Power
Current measurement module group 20 | 18750 | Float | RD | var | Maximum Value L4 Distortion Reactive Power
Current measurement module group 20 | 18752 | Float | RD | var | Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 20 | 18754 | Float | RD | % | Maximum Value L1 THD Current
Current measurement module group 20 | 18756 | Float | RD | % | Maximum Value L2 THD Current
Current measurement module group 20 | 18758 | Float | RD | % | Maximum Value L3 THD Current
Current measurement module group 20 | 18760 | Float | RD | % | Maximum Value L4 THD Current
Current measurement module group 20 | 18762 | Float | RD | % | Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 20 | 18764 | Float | RD | % | Maximum Value L1 TDD Current
Current measurement module group 20 | 18766 | Float | RD | % | Maximum Value L2 TDD Current
Current measurement module group 20 | 18768 | Float | RD | % | Maximum Value L3 TDD Current
Current measurement module group 20 | 18770 | Float | RD | % | Maximum Value L4 TDD Current
Current measurement module group 20 | 18772 | Float | RD | % | Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 20 | 18774 | Float | RD | % | Maximum Value L1 Load Current (%)
Current measurement module group 20 | 18776 | Float | RD | % | Maximum Value L2 Load Current (%)
Current measurement module group 20 | 18778 | Float | RD | % | Maximum Value L3 Load Current (%)
Current measurement module group 20 | 18780 | Float | RD | % | Maximum Value L4 Load Current (%)
Current measurement module group 20 | 18782 | Float | RD | % | Maximum Value Sum L1..L4 Sum Load Current (%)
Current measurement module group 20 | 18784 | DateTime | RD | - | DateTime Maximum Value L1 Current (RMS)
Current measurement module group 20 | 18786 | DateTime | RD | - | DateTime Maximum Value L2 Current (RMS)
Current measurement module group 20 | 18788 | DateTime | RD | - | DateTime Maximum Value L3 Current (RMS)
Current measurement module group 20 | 18790 | DateTime | RD | - | DateTime Maximum Value L4 Current (RMS)
Current measurement module group 20 | 18792 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Current (RMS)
Current measurement module group 20 | 18794 | DateTime | RD | - | DateTime Maximum Value L1 Active Power
Current measurement module group 20 | 18796 | DateTime | RD | - | DateTime Maximum Value L2 Active Power
Current measurement module group 20 | 18798 | DateTime | RD | - | DateTime Maximum Value L3 Active Power
Current measurement module group 20 | 18800 | DateTime | RD | - | DateTime Maximum Value L4 Active Power
Current measurement module group 20 | 18802 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power
Current measurement module group 20 | 18804 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power
Current measurement module group 20 | 18806 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power
Current measurement module group 20 | 18808 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power
Current measurement module group 20 | 18810 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power
Current measurement module group 20 | 18812 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power
Current measurement module group 20 | 18814 | DateTime | RD | - | DateTime Maximum Value L1 Apparent Power
Current measurement module group 20 | 18816 | DateTime | RD | - | DateTime Maximum Value L2 Apparent Power
Current measurement module group 20 | 18818 | DateTime | RD | - | DateTime Maximum Value L3 Apparent Power
Current measurement module group 20 | 18820 | DateTime | RD | - | DateTime Maximum Value L4 Apparent Power
Current measurement module group 20 | 18822 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Apparent Power
Current measurement module group 20 | 18824 | DateTime | RD | - | DateTime Maximum Value L1 Cos(phi)
Current measurement module group 20 | 18826 | DateTime | RD | - | DateTime Maximum Value L2 Cos(phi)
Current measurement module group 20 | 18828 | DateTime | RD | - | DateTime Maximum Value L3 Cos(phi)
Current measurement module group 20 | 18830 | DateTime | RD | - | DateTime Maximum Value L4 Cos(phi)
Current measurement module group 20 | 18832 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Cos(phi)
Current measurement module group 20 | 18834 | DateTime | RD | - | DateTime Maximum Value L1 Power Factor
Current measurement module group 20 | 18836 | DateTime | RD | - | DateTime Maximum Value L2 Power Factor
Current measurement module group 20 | 18838 | DateTime | RD | - | DateTime Maximum Value L3 Power Factor
Current measurement module group 20 | 18840 | DateTime | RD | - | DateTime Maximum Value L4 Power Factor
Current measurement module group 20 | 18842 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Power Factor
Current measurement module group 20 | 18844 | DateTime | RD | - | DateTime Maximum Value L1 Active Power (fundamental)
Current measurement module group 20 | 18846 | DateTime | RD | - | DateTime Maximum Value L2 Active Power (fundamental)
Current measurement module group 20 | 18848 | DateTime | RD | - | DateTime Maximum Value L3 Active Power (fundamental)
Current measurement module group 20 | 18850 | DateTime | RD | - | DateTime Maximum Value L4 Active Power (fundamental)
Current measurement module group 20 | 18852 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Active Power (fundamental)
Current measurement module group 20 | 18854 | DateTime | RD | - | DateTime Maximum Value L1 Reactive Power (fundamental)
Current measurement module group 20 | 18856 | DateTime | RD | - | DateTime Maximum Value L2 Reactive Power (fundamental)
Current measurement module group 20 | 18858 | DateTime | RD | - | DateTime Maximum Value L3 Reactive Power (fundamental)
Current measurement module group 20 | 18860 | DateTime | RD | - | DateTime Maximum Value L4 Reactive Power (fundamental)
Current measurement module group 20 | 18862 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Reactive Power (fundamental)
Current measurement module group 20 | 18864 | DateTime | RD | - | DateTime Maximum Value L1 Distortion Reactive Power
Current measurement module group 20 | 18866 | DateTime | RD | - | DateTime Maximum Value L2 Distortion Reactive Power
Current measurement module group 20 | 18868 | DateTime | RD | - | DateTime Maximum Value L3 Distortion Reactive Power
Current measurement module group 20 | 18870 | DateTime | RD | - | DateTime Maximum Value L4 Distortion Reactive Power
Current measurement module group 20 | 18872 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Distortion Reactive Power
Current measurement module group 20 | 18874 | DateTime | RD | - | DateTime Maximum Value L1 THD Current
Current measurement module group 20 | 18876 | DateTime | RD | - | DateTime Maximum Value L2 THD Current
Current measurement module group 20 | 18878 | DateTime | RD | - | DateTime Maximum Value L3 THD Current
Current measurement module group 20 | 18880 | DateTime | RD | - | DateTime Maximum Value L4 THD Current
Current measurement module group 20 | 18882 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum THD Current
Current measurement module group 20 | 18884 | DateTime | RD | - | DateTime Maximum Value L1 TDD Current
Current measurement module group 20 | 18886 | DateTime | RD | - | DateTime Maximum Value L2 TDD Current
Current measurement module group 20 | 18888 | DateTime | RD | - | DateTime Maximum Value L3 TDD Current
Current measurement module group 20 | 18890 | DateTime | RD | - | DateTime Maximum Value L4 TDD Current
Current measurement module group 20 | 18892 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum TDD Current
Current measurement module group 20 | 18894 | DateTime | RD | - | DateTime Maximum Value L1 Load Current (%)
Current measurement module group 20 | 18896 | DateTime | RD | - | DateTime Maximum Value L2 Load Current (%)
Current measurement module group 20 | 18898 | DateTime | RD | - | DateTime Maximum Value L3 Load Current (%)
Current measurement module group 20 | 18900 | DateTime | RD | - | DateTime Maximum Value L4 Load Current (%)
Current measurement module group 20 | 18902 | DateTime | RD | - | DateTime Maximum Value Sum L1..L4 Sum Load Current (%)
basic device | 19000 | Float | RD | V | Voltage L1-N
basic device | 19002 | Float | RD | V | Voltage L2-N
basic device | 19004 | Float | RD | V | Voltage L3-N
basic device | 19006 | Float | RD | V | Voltage L1-L2
basic device | 19008 | Float | RD | V | Voltage L2-L3
basic device | 19010 | Float | RD | V | Voltage L3-L1
Basic device measuring group 01 | 19012 | Float | RD | A | Current, I1
Basic device measuring group 01 | 19014 | Float | RD | A | Current, I2
Basic device measuring group 01 | 19016 | Float | RD | A | Current, I3
Basic device measuring group 01 | 19018 | Float | RD | A | Current, I4
Basic device measuring group 01 | 19020 | Float | RD | W | Active power P1
Basic device measuring group 01 | 19022 | Float | RD | W | Active power P2
Basic device measuring group 01 | 19024 | Float | RD | W | Active power P3
Basic device measuring group 01 | 19026 | Float | RD | W | Sum; Psum3=P1+P2+P3
Basic device measuring group 01 | 19028 | Float | RD | VA | Apparent power S1
Basic device measuring group 01 | 19030 | Float | RD | VA | Apparent power S2
Basic device measuring group 01 | 19032 | Float | RD | VA | Apparent power S3
Basic device measuring group 01 | 19034 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Basic device measuring group 01 | 19036 | Float | RD | var | Reactive power (fundamental frequency) Q1
Basic device measuring group 01 | 19038 | Float | RD | var | Reactive power (fundamental frequency) Q2
Basic device measuring group 01 | 19040 | Float | RD | var | Reactive power (fundamental frequency) Q3
Basic device measuring group 01 | 19042 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Basic device measuring group 01 | 19044 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Basic device measuring group 01 | 19046 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Basic device measuring group 01 | 19048 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Basic device measuring group 01 | 19050 | Float | RD | Hz | Frequency
Basic device measuring group 01 | 19052 | RotationFieldKind | RD | - | Rotary field; 1=right, 0=none, -1=left
Basic device measuring group 01 | 19054 | Float | RD | Wh | Active energy W1
Basic device measuring group 01 | 19056 | Float | RD | Wh | Active energy W2
Basic device measuring group 01 | 19058 | Float | RD | Wh | Active energy W3
Basic device measuring group 01 | 19060 | Float | RD | Wh | Active energy W1..W3
Basic device measuring group 01 | 19062 | Float | RD | Wh | Active energy W1, related
Basic device measuring group 01 | 19064 | Float | RD | Wh | Active energy W2, drawn
Basic device measuring group 01 | 19066 | Float | RD | Wh | Active energy W3, purchased
Basic device measuring group 01 | 19068 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Basic device measuring group 01 | 19070 | Float | RD | Wh | Active energy W1, supplied
Basic device measuring group 01 | 19072 | Float | RD | Wh | Active energy W2, supplied
Basic device measuring group 01 | 19074 | Float | RD | Wh | Active energy W3, supplied
Basic device measuring group 01 | 19076 | Float | RD | Wh | Active energy W1..W3, supplied
Basic device measuring group 01 | 19078 | Float | RD | VAh | Apparent energy WS1
Basic device measuring group 01 | 19080 | Float | RD | VAh | Apparent energy WS2
Basic device measuring group 01 | 19082 | Float | RD | VAh | Apparent energy WS3
Basic device measuring group 01 | 19084 | Float | RD | VAh | Apparent energy WS1..WS3
Basic device measuring group 01 | 19086 | Float | RD | varh | Reactive energy WQ1
Basic device measuring group 01 | 19088 | Float | RD | varh | Reactive energy WQ2
Basic device measuring group 01 | 19090 | Float | RD | varh | Reactive energy WQ3
Basic device measuring group 01 | 19092 | Float | RD | varh | Reactive energy WQ1..WQ3
Basic device measuring group 01 | 19094 | Float | RD | varh | Reactive energy WQ1, inductive
Basic device measuring group 01 | 19096 | Float | RD | varh | Reactive energy WQ2, inductive
Basic device measuring group 01 | 19098 | Float | RD | varh | Reactive energy WQ3, inductive
Basic device measuring group 01 | 19100 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Basic device measuring group 01 | 19102 | Float | RD | varh | Reactive energy WQ1, capacitive
Basic device measuring group 01 | 19104 | Float | RD | varh | Reactive energy WQ2, capacitive
Basic device measuring group 01 | 19106 | Float | RD | varh | Reactive energy WQ3, capacitive
Basic device measuring group 01 | 19108 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Basic device measuring group 01 | 19110 | Float | RD | % | Harmonic, THD,U L1-N
Basic device measuring group 01 | 19112 | Float | RD | % | Harmonic, THD,U L2-N
Basic device measuring group 01 | 19114 | Float | RD | % | Harmonic, THD,U L3-N
Basic device measuring group 01 | 19116 | Float | RD | % | Harmonic, THD, I1
Basic device measuring group 01 | 19118 | Float | RD | % | Harmonic, THD, I2
Basic device measuring group 01 | 19120 | Float | RD | % | Harmonic, THD, I3
Basic device measuring group 02 | 19200 | Float | RD | A | Current, I1
Basic device measuring group 02 | 19202 | Float | RD | A | Current, I2
Basic device measuring group 02 | 19204 | Float | RD | A | Current, I3
Basic device measuring group 02 | 19206 | Float | RD | A | Current, I4
Basic device measuring group 02 | 19208 | Float | RD | W | Active power P1
Basic device measuring group 02 | 19210 | Float | RD | W | Active power P2
Basic device measuring group 02 | 19212 | Float | RD | W | Active power P3
Basic device measuring group 02 | 19214 | Float | RD | W | Sum; Psum3=P1+P2+P3
Basic device measuring group 02 | 19216 | Float | RD | VA | Apparent power S1
Basic device measuring group 02 | 19218 | Float | RD | VA | Apparent power S2
Basic device measuring group 02 | 19220 | Float | RD | VA | Apparent power S3
Basic device measuring group 02 | 19222 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Basic device measuring group 02 | 19224 | Float | RD | var | Reactive power (fundamental frequency) Q1
Basic device measuring group 02 | 19226 | Float | RD | var | Reactive power (fundamental frequency) Q2
Basic device measuring group 02 | 19228 | Float | RD | var | Reactive power (fundamental frequency) Q3
Basic device measuring group 02 | 19230 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Basic device measuring group 02 | 19232 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Basic device measuring group 02 | 19234 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Basic device measuring group 02 | 19236 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Basic device measuring group 02 | 19238 | Float | RD | Wh | Active energy W1
Basic device measuring group 02 | 19240 | Float | RD | Wh | Active energy W2
Basic device measuring group 02 | 19242 | Float | RD | Wh | Active energy W3
Basic device measuring group 02 | 19244 | Float | RD | Wh | Active energy W1..W3
Basic device measuring group 02 | 19246 | Float | RD | Wh | Active energy W1, related
Basic device measuring group 02 | 19248 | Float | RD | Wh | Active energy W2, drawn
Basic device measuring group 02 | 19250 | Float | RD | Wh | Active energy W3, purchased
Basic device measuring group 02 | 19252 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Basic device measuring group 02 | 19254 | Float | RD | Wh | Active energy W1, supplied
Basic device measuring group 02 | 19256 | Float | RD | Wh | Active energy W2, supplied
Basic device measuring group 02 | 19258 | Float | RD | Wh | Active energy W3, supplied
Basic device measuring group 02 | 19260 | Float | RD | Wh | Active energy W1..W3, supplied
Basic device measuring group 02 | 19262 | Float | RD | VAh | Apparent energy WS1
Basic device measuring group 02 | 19264 | Float | RD | VAh | Apparent energy WS2
Basic device measuring group 02 | 19266 | Float | RD | VAh | Apparent energy WS3
Basic device measuring group 02 | 19268 | Float | RD | VAh | Apparent energy WS1..WS3
Basic device measuring group 02 | 19270 | Float | RD | varh | Reactive energy WQ1
Basic device measuring group 02 | 19272 | Float | RD | varh | Reactive energy WQ2
Basic device measuring group 02 | 19274 | Float | RD | varh | Reactive energy WQ3
Basic device measuring group 02 | 19276 | Float | RD | varh | Reactive energy WQ1..WQ3
Basic device measuring group 02 | 19278 | Float | RD | varh | Reactive energy WQ1, inductive
Basic device measuring group 02 | 19280 | Float | RD | varh | Reactive energy WQ2, inductive
Basic device measuring group 02 | 19282 | Float | RD | varh | Reactive energy WQ3, inductive
Basic device measuring group 02 | 19284 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Basic device measuring group 02 | 19286 | Float | RD | varh | Reactive energy WQ1, capacitive
Basic device measuring group 02 | 19288 | Float | RD | varh | Reactive energy WQ2, capacitive
Basic device measuring group 02 | 19290 | Float | RD | varh | Reactive energy WQ3, capacitive
Basic device measuring group 02 | 19292 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Basic device measuring group 02 | 19294 | Float | RD | % | Harmonics, THD, I1
Basic device measuring group 02 | 19296 | Float | RD | % | Harmonic, THD, I2
Basic device measuring group 02 | 19298 | Float | RD | % | Harmonic, THD, I3
Basic device measuring group 03 | 19300 | Float | RD | A | Current, I1
Basic device measuring group 03 | 19302 | Float | RD | A | Current, I2
Basic device measuring group 03 | 19304 | Float | RD | A | Current, I3
Basic device measuring group 03 | 19306 | Float | RD | A | Current, I4
Basic device measuring group 03 | 19308 | Float | RD | W | Active power P1
Basic device measuring group 03 | 19310 | Float | RD | W | Active power P2
Basic device measuring group 03 | 19312 | Float | RD | W | Active power P3
Basic device measuring group 03 | 19314 | Float | RD | W | Sum; Psum3=P1+P2+P3
Basic device measuring group 03 | 19316 | Float | RD | VA | Apparent power S1
Basic device measuring group 03 | 19318 | Float | RD | VA | Apparent power S2
Basic device measuring group 03 | 19320 | Float | RD | VA | Apparent power S3
Basic device measuring group 03 | 19322 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Basic device measuring group 03 | 19324 | Float | RD | var | Reactive power (fundamental frequency) Q1
Basic device measuring group 03 | 19326 | Float | RD | var | Reactive power (fundamental frequency) Q2
Basic device measuring group 03 | 19328 | Float | RD | var | Reactive power (fundamental frequency) Q3
Basic device measuring group 03 | 19330 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Basic device measuring group 03 | 19332 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Basic device measuring group 03 | 19334 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Basic device measuring group 03 | 19336 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Basic device measuring group 03 | 19338 | Float | RD | Wh | Active energy W1
Basic device measuring group 03 | 19340 | Float | RD | Wh | Active energy W2
Basic device measuring group 03 | 19342 | Float | RD | Wh | Active energy W3
Basic device measuring group 03 | 19344 | Float | RD | Wh | Active energy W1..W3
Basic device measuring group 03 | 19346 | Float | RD | Wh | Active energy W1, related
Basic device measuring group 03 | 19348 | Float | RD | Wh | Active energy W2, drawn
Basic device measuring group 03 | 19350 | Float | RD | Wh | Active energy W3, purchased
Basic device measuring group 03 | 19352 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Basic device measuring group 03 | 19354 | Float | RD | Wh | Active energy W1, supplied
Basic device measuring group 03 | 19356 | Float | RD | Wh | Active energy W2, supplied
Basic device measuring group 03 | 19358 | Float | RD | Wh | Active energy W3, supplied
Basic device measuring group 03 | 19360 | Float | RD | Wh | Active energy W1..W3, supplied
Basic device measuring group 03 | 19362 | Float | RD | VAh | Apparent energy WS1
Basic device measuring group 03 | 19364 | Float | RD | VAh | Apparent energy WS2
Basic device measuring group 03 | 19366 | Float | RD | VAh | Apparent energy WS3
Basic device measuring group 03 | 19368 | Float | RD | VAh | Apparent energy WS1..WS3
Basic device measuring group 03 | 19370 | Float | RD | varh | Reactive energy WQ1
Basic device measuring group 03 | 19372 | Float | RD | varh | Reactive energy WQ2
Basic device measuring group 03 | 19374 | Float | RD | varh | Reactive energy WQ3
Basic device measuring group 03 | 19376 | Float | RD | varh | Reactive energy WQ1..WQ3
Basic device measuring group 03 | 19378 | Float | RD | varh | Reactive energy WQ1, inductive
Basic device measuring group 03 | 19380 | Float | RD | varh | Reactive energy WQ2, inductive
Basic device measuring group 03 | 19382 | Float | RD | varh | Reactive energy WQ3, inductive
Basic device measuring group 03 | 19384 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Basic device measuring group 03 | 19386 | Float | RD | varh | Reactive energy WQ1, capacitive
Basic device measuring group 03 | 19388 | Float | RD | varh | Reactive energy WQ2, capacitive
Basic device measuring group 03 | 19390 | Float | RD | varh | Reactive energy WQ3, capacitive
Basic device measuring group 03 | 19392 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Basic device measuring group 03 | 19394 | Float | RD | % | Harmonics, THD, I1
Basic device measuring group 03 | 19396 | Float | RD | % | Harmonic, THD, I2
Basic device measuring group 03 | 19398 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 01 | 19400 | Float | RD | A | Current, I1
Current measurement module group 01 | 19402 | Float | RD | A | Current, I2
Current measurement module group 01 | 19404 | Float | RD | A | Current, I3
Current measurement module group 01 | 19406 | Float | RD | A | Current, I4
Current measurement module group 01 | 19408 | Float | RD | W | Active power P1
Current measurement module group 01 | 19410 | Float | RD | W | Active power P2
Current measurement module group 01 | 19412 | Float | RD | W | Active power P3
Current measurement module group 01 | 19414 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 01 | 19416 | Float | RD | VA | Apparent power S1
Current measurement module group 01 | 19418 | Float | RD | VA | Apparent power S2
Current measurement module group 01 | 19420 | Float | RD | VA | Apparent power S3
Current measurement module group 01 | 19422 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 01 | 19424 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 01 | 19426 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 01 | 19428 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 01 | 19430 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 01 | 19432 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 01 | 19434 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 01 | 19436 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 01 | 19438 | Float | RD | Wh | Active energy W1
Current measurement module group 01 | 19440 | Float | RD | Wh | Active energy W2
Current measurement module group 01 | 19442 | Float | RD | Wh | Active energy W3
Current measurement module group 01 | 19444 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 01 | 19446 | Float | RD | Wh | Active energy W1, related
Current measurement module group 01 | 19448 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 01 | 19450 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 01 | 19452 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 01 | 19454 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 01 | 19456 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 01 | 19458 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 01 | 19460 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 01 | 19462 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 01 | 19464 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 01 | 19466 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 01 | 19468 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 01 | 19470 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 01 | 19472 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 01 | 19474 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 01 | 19476 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 01 | 19478 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 01 | 19480 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 01 | 19482 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 01 | 19484 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 01 | 19486 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 01 | 19488 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 01 | 19490 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 01 | 19492 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 01 | 19494 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 01 | 19496 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 01 | 19498 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 02 | 19500 | Float | RD | A | Current, I1
Current measurement module group 02 | 19502 | Float | RD | A | Current, I2
Current measurement module group 02 | 19504 | Float | RD | A | Current, I3
Current measurement module group 02 | 19506 | Float | RD | A | Current, I4
Current measurement module group 02 | 19508 | Float | RD | W | Active power P1
Current measurement module group 02 | 19510 | Float | RD | W | Active power P2
Current measurement module group 02 | 19512 | Float | RD | W | Active power P3
Current measurement module group 02 | 19514 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 02 | 19516 | Float | RD | VA | Apparent power S1
Current measurement module group 02 | 19518 | Float | RD | VA | Apparent power S2
Current measurement module group 02 | 19520 | Float | RD | VA | Apparent power S3
Current measurement module group 02 | 19522 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 02 | 19524 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 02 | 19526 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 02 | 19528 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 02 | 19530 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 02 | 19532 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 02 | 19534 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 02 | 19536 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 02 | 19538 | Float | RD | Wh | Active energy W1
Current measurement module group 02 | 19540 | Float | RD | Wh | Active energy W2
Current measurement module group 02 | 19542 | Float | RD | Wh | Active energy W3
Current measurement module group 02 | 19544 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 02 | 19546 | Float | RD | Wh | Active energy W1, related
Current measurement module group 02 | 19548 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 02 | 19550 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 02 | 19552 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 02 | 19554 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 02 | 19556 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 02 | 19558 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 02 | 19560 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 02 | 19562 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 02 | 19564 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 02 | 19566 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 02 | 19568 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 02 | 19570 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 02 | 19572 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 02 | 19574 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 02 | 19576 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 02 | 19578 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 02 | 19580 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 02 | 19582 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 02 | 19584 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 02 | 19586 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 02 | 19588 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 02 | 19590 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 02 | 19592 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 02 | 19594 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 02 | 19596 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 02 | 19598 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 03 | 19600 | Float | RD | A | Current, I1
Current measurement module group 03 | 19602 | Float | RD | A | Current, I2
Current measurement module group 03 | 19604 | Float | RD | A | Current, I3
Current measurement module group 03 | 19606 | Float | RD | A | Current, I4
Current measurement module group 03 | 19608 | Float | RD | W | Active power P1
Current measurement module group 03 | 19610 | Float | RD | W | Active power P2
Current measurement module group 03 | 19612 | Float | RD | W | Active power P3
Current measurement module group 03 | 19614 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 03 | 19616 | Float | RD | VA | Apparent power S1
Current measurement module group 03 | 19618 | Float | RD | VA | Apparent power S2
Current measurement module group 03 | 19620 | Float | RD | VA | Apparent power S3
Current measurement module group 03 | 19622 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 03 | 19624 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 03 | 19626 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 03 | 19628 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 03 | 19630 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 03 | 19632 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 03 | 19634 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 03 | 19636 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 03 | 19638 | Float | RD | Wh | Active energy W1
Current measurement module group 03 | 19640 | Float | RD | Wh | Active energy W2
Current measurement module group 03 | 19642 | Float | RD | Wh | Active energy W3
Current measurement module group 03 | 19644 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 03 | 19646 | Float | RD | Wh | Active energy W1, related
Current measurement module group 03 | 19648 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 03 | 19650 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 03 | 19652 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 03 | 19654 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 03 | 19656 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 03 | 19658 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 03 | 19660 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 03 | 19662 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 03 | 19664 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 03 | 19666 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 03 | 19668 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 03 | 19670 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 03 | 19672 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 03 | 19674 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 03 | 19676 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 03 | 19678 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 03 | 19680 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 03 | 19682 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 03 | 19684 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 03 | 19686 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 03 | 19688 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 03 | 19690 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 03 | 19692 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 03 | 19694 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 03 | 19696 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 03 | 19698 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 04 | 19700 | Float | RD | A | Current, I1
Current measurement module group 04 | 19702 | Float | RD | A | Current, I2
Current measurement module group 04 | 19704 | Float | RD | A | Current, I3
Current measurement module group 04 | 19706 | Float | RD | A | Current, I4
Current measurement module group 04 | 19708 | Float | RD | W | Active power P1
Current measurement module group 04 | 19710 | Float | RD | W | Active power P2
Current measurement module group 04 | 19712 | Float | RD | W | Active power P3
Current measurement module group 04 | 19714 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 04 | 19716 | Float | RD | VA | Apparent power S1
Current measurement module group 04 | 19718 | Float | RD | VA | Apparent power S2
Current measurement module group 04 | 19720 | Float | RD | VA | Apparent power S3
Current measurement module group 04 | 19722 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 04 | 19724 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 04 | 19726 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 04 | 19728 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 04 | 19730 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 04 | 19732 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 04 | 19734 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 04 | 19736 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 04 | 19738 | Float | RD | Wh | Active energy W1
Current measurement module group 04 | 19740 | Float | RD | Wh | Active energy W2
Current measurement module group 04 | 19742 | Float | RD | Wh | Active energy W3
Current measurement module group 04 | 19744 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 04 | 19746 | Float | RD | Wh | Active energy W1, related
Current measurement module group 04 | 19748 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 04 | 19750 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 04 | 19752 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 04 | 19754 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 04 | 19756 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 04 | 19758 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 04 | 19760 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 04 | 19762 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 04 | 19764 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 04 | 19766 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 04 | 19768 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 04 | 19770 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 04 | 19772 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 04 | 19774 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 04 | 19776 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 04 | 19778 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 04 | 19780 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 04 | 19782 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 04 | 19784 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 04 | 19786 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 04 | 19788 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 04 | 19790 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 04 | 19792 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 04 | 19794 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 04 | 19796 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 04 | 19798 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 05 | 19800 | Float | RD | A | Current, I1
Current measurement module group 05 | 19802 | Float | RD | A | Current, I2
Current measurement module group 05 | 19804 | Float | RD | A | Current, I3
Current measurement module group 05 | 19806 | Float | RD | A | Current, I4
Current measurement module group 05 | 19808 | Float | RD | W | Active power P1
Current measurement module group 05 | 19810 | Float | RD | W | Active power P2
Current measurement module group 05 | 19812 | Float | RD | W | Active power P3
Current measurement module group 05 | 19814 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 05 | 19816 | Float | RD | VA | Apparent power S1
Current measurement module group 05 | 19818 | Float | RD | VA | Apparent power S2
Current measurement module group 05 | 19820 | Float | RD | VA | Apparent power S3
Current measurement module group 05 | 19822 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 05 | 19824 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 05 | 19826 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 05 | 19828 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 05 | 19830 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 05 | 19832 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 05 | 19834 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 05 | 19836 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 05 | 19838 | Float | RD | Wh | Active energy W1
Current measurement module group 05 | 19840 | Float | RD | Wh | Active energy W2
Current measurement module group 05 | 19842 | Float | RD | Wh | Active energy W3
Current measurement module group 05 | 19844 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 05 | 19846 | Float | RD | Wh | Active energy W1, related
Current measurement module group 05 | 19848 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 05 | 19850 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 05 | 19852 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 05 | 19854 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 05 | 19856 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 05 | 19858 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 05 | 19860 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 05 | 19862 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 05 | 19864 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 05 | 19866 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 05 | 19868 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 05 | 19870 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 05 | 19872 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 05 | 19874 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 05 | 19876 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 05 | 19878 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 05 | 19880 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 05 | 19882 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 05 | 19884 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 05 | 19886 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 05 | 19888 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 05 | 19890 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 05 | 19892 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 05 | 19894 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 05 | 19896 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 05 | 19898 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 06 | 19900 | Float | RD | A | Current, I1
Current measurement module group 06 | 19902 | Float | RD | A | Current, I2
Current measurement module group 06 | 19904 | Float | RD | A | Current, I3
Current measurement module group 06 | 19906 | Float | RD | A | Current, I4
Current measurement module group 06 | 19908 | Float | RD | W | Active power P1
Current measurement module group 06 | 19910 | Float | RD | W | Active power P2
Current measurement module group 06 | 19912 | Float | RD | W | Active power P3
Current measurement module group 06 | 19914 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 06 | 19916 | Float | RD | VA | Apparent power S1
Current measurement module group 06 | 19918 | Float | RD | VA | Apparent power S2
Current measurement module group 06 | 19920 | Float | RD | VA | Apparent power S3
Current measurement module group 06 | 19922 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 06 | 19924 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 06 | 19926 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 06 | 19928 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 06 | 19930 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 06 | 19932 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 06 | 19934 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 06 | 19936 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 06 | 19938 | Float | RD | Wh | Active energy W1
Current measurement module group 06 | 19940 | Float | RD | Wh | Active energy W2
Current measurement module group 06 | 19942 | Float | RD | Wh | Active energy W3
Current measurement module group 06 | 19944 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 06 | 19946 | Float | RD | Wh | Active energy W1, related
Current measurement module group 06 | 19948 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 06 | 19950 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 06 | 19952 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 06 | 19954 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 06 | 19956 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 06 | 19958 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 06 | 19960 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 06 | 19962 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 06 | 19964 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 06 | 19966 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 06 | 19968 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 06 | 19970 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 06 | 19972 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 06 | 19974 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 06 | 19976 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 06 | 19978 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 06 | 19980 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 06 | 19982 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 06 | 19984 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 06 | 19986 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 06 | 19988 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 06 | 19990 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 06 | 19992 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 06 | 19994 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 06 | 19996 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 06 | 19998 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 07 | 20000 | Float | RD | A | Current, I1
Current measurement module group 07 | 20002 | Float | RD | A | Current, I2
Current measurement module group 07 | 20004 | Float | RD | A | Current, I3
Current measurement module group 07 | 20006 | Float | RD | A | Current, I4
Current measurement module group 07 | 20008 | Float | RD | W | Active power P1
Current measurement module group 07 | 20010 | Float | RD | W | Active power P2
Current measurement module group 07 | 20012 | Float | RD | W | Active power P3
Current measurement module group 07 | 20014 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 07 | 20016 | Float | RD | VA | Apparent power S1
Current measurement module group 07 | 20018 | Float | RD | VA | Apparent power S2
Current measurement module group 07 | 20020 | Float | RD | VA | Apparent power S3
Current measurement module group 07 | 20022 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 07 | 20024 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 07 | 20026 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 07 | 20028 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 07 | 20030 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 07 | 20032 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 07 | 20034 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 07 | 20036 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 07 | 20038 | Float | RD | Wh | Active energy W1
Current measurement module group 07 | 20040 | Float | RD | Wh | Active energy W2
Current measurement module group 07 | 20042 | Float | RD | Wh | Active energy W3
Current measurement module group 07 | 20044 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 07 | 20046 | Float | RD | Wh | Active energy W1, related
Current measurement module group 07 | 20048 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 07 | 20050 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 07 | 20052 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 07 | 20054 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 07 | 20056 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 07 | 20058 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 07 | 20060 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 07 | 20062 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 07 | 20064 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 07 | 20066 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 07 | 20068 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 07 | 20070 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 07 | 20072 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 07 | 20074 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 07 | 20076 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 07 | 20078 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 07 | 20080 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 07 | 20082 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 07 | 20084 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 07 | 20086 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 07 | 20088 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 07 | 20090 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 07 | 20092 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 07 | 20094 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 07 | 20096 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 07 | 20098 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 08 | 20100 | Float | RD | A | Current, I1
Current measurement module group 08 | 20102 | Float | RD | A | Current, I2
Current measurement module group 08 | 20104 | Float | RD | A | Current, I3
Current measurement module group 08 | 20106 | Float | RD | A | Current, I4
Current measurement module group 08 | 20108 | Float | RD | W | Active power P1
Current measurement module group 08 | 20110 | Float | RD | W | Active power P2
Current measurement module group 08 | 20112 | Float | RD | W | Active power P3
Current measurement module group 08 | 20114 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 08 | 20116 | Float | RD | VA | Apparent power S1
Current measurement module group 08 | 20118 | Float | RD | VA | Apparent power S2
Current measurement module group 08 | 20120 | Float | RD | VA | Apparent power S3
Current measurement module group 08 | 20122 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 08 | 20124 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 08 | 20126 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 08 | 20128 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 08 | 20130 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 08 | 20132 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 08 | 20134 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 08 | 20136 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 08 | 20138 | Float | RD | Wh | Active energy W1
Current measurement module group 08 | 20140 | Float | RD | Wh | Active energy W2
Current measurement module group 08 | 20142 | Float | RD | Wh | Active energy W3
Current measurement module group 08 | 20144 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 08 | 20146 | Float | RD | Wh | Active energy W1, related
Current measurement module group 08 | 20148 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 08 | 20150 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 08 | 20152 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 08 | 20154 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 08 | 20156 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 08 | 20158 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 08 | 20160 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 08 | 20162 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 08 | 20164 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 08 | 20166 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 08 | 20168 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 08 | 20170 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 08 | 20172 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 08 | 20174 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 08 | 20176 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 08 | 20178 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 08 | 20180 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 08 | 20182 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 08 | 20184 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 08 | 20186 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 08 | 20188 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 08 | 20190 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 08 | 20192 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 08 | 20194 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 08 | 20196 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 08 | 20198 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 09 | 20200 | Float | RD | A | Current, I1
Current measurement module group 09 | 20202 | Float | RD | A | Current, I2
Current measurement module group 09 | 20204 | Float | RD | A | Current, I3
Current measurement module group 09 | 20206 | Float | RD | A | Current, I4
Current measurement module group 09 | 20208 | Float | RD | W | Active power P1
Current measurement module group 09 | 20210 | Float | RD | W | Active power P2
Current measurement module group 09 | 20212 | Float | RD | W | Active power P3
Current measurement module group 09 | 20214 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 09 | 20216 | Float | RD | VA | Apparent power S1
Current measurement module group 09 | 20218 | Float | RD | VA | Apparent power S2
Current measurement module group 09 | 20220 | Float | RD | VA | Apparent power S3
Current measurement module group 09 | 20222 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 09 | 20224 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 09 | 20226 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 09 | 20228 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 09 | 20230 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 09 | 20232 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 09 | 20234 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 09 | 20236 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 09 | 20238 | Float | RD | Wh | Active energy W1
Current measurement module group 09 | 20240 | Float | RD | Wh | Active energy W2
Current measurement module group 09 | 20242 | Float | RD | Wh | Active energy W3
Current measurement module group 09 | 20244 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 09 | 20246 | Float | RD | Wh | Active energy W1, related
Current measurement module group 09 | 20248 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 09 | 20250 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 09 | 20252 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 09 | 20254 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 09 | 20256 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 09 | 20258 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 09 | 20260 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 09 | 20262 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 09 | 20264 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 09 | 20266 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 09 | 20268 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 09 | 20270 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 09 | 20272 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 09 | 20274 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 09 | 20276 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 09 | 20278 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 09 | 20280 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 09 | 20282 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 09 | 20284 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 09 | 20286 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 09 | 20288 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 09 | 20290 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 09 | 20292 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 09 | 20294 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 09 | 20296 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 09 | 20298 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 10 | 20300 | Float | RD | A | Current, I1
Current measurement module group 10 | 20302 | Float | RD | A | Current, I2
Current measurement module group 10 | 20304 | Float | RD | A | Current, I3
Current measurement module group 10 | 20306 | Float | RD | A | Current, I4
Current measurement module group 10 | 20308 | Float | RD | W | Active power P1
Current measurement module group 10 | 20310 | Float | RD | W | Active power P2
Current measurement module group 10 | 20312 | Float | RD | W | Active power P3
Current measurement module group 10 | 20314 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 10 | 20316 | Float | RD | VA | Apparent power S1
Current measurement module group 10 | 20318 | Float | RD | VA | Apparent power S2
Current measurement module group 10 | 20320 | Float | RD | VA | Apparent power S3
Current measurement module group 10 | 20322 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 10 | 20324 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 10 | 20326 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 10 | 20328 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 10 | 20330 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 10 | 20332 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 10 | 20334 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 10 | 20336 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 10 | 20338 | Float | RD | Wh | Active energy W1
Current measurement module group 10 | 20340 | Float | RD | Wh | Active energy W2
Current measurement module group 10 | 20342 | Float | RD | Wh | Active energy W3
Current measurement module group 10 | 20344 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 10 | 20346 | Float | RD | Wh | Active energy W1, related
Current measurement module group 10 | 20348 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 10 | 20350 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 10 | 20352 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 10 | 20354 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 10 | 20356 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 10 | 20358 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 10 | 20360 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 10 | 20362 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 10 | 20364 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 10 | 20366 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 10 | 20368 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 10 | 20370 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 10 | 20372 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 10 | 20374 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 10 | 20376 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 10 | 20378 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 10 | 20380 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 10 | 20382 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 10 | 20384 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 10 | 20386 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 10 | 20388 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 10 | 20390 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 10 | 20392 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 10 | 20394 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 10 | 20396 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 10 | 20398 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 11 | 20400 | Float | RD | A | Current, I1
Current measurement module group 11 | 20402 | Float | RD | A | Current, I2
Current measurement module group 11 | 20404 | Float | RD | A | Current, I3
Current measurement module group 11 | 20406 | Float | RD | A | Current, I4
Current measurement module group 11 | 20408 | Float | RD | W | Active power P1
Current measurement module group 11 | 20410 | Float | RD | W | Active power P2
Current measurement module group 11 | 20412 | Float | RD | W | Active power P3
Current measurement module group 11 | 20414 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 11 | 20416 | Float | RD | VA | Apparent power S1
Current measurement module group 11 | 20418 | Float | RD | VA | Apparent power S2
Current measurement module group 11 | 20420 | Float | RD | VA | Apparent power S3
Current measurement module group 11 | 20422 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 11 | 20424 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 11 | 20426 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 11 | 20428 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 11 | 20430 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 11 | 20432 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 11 | 20434 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 11 | 20436 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 11 | 20438 | Float | RD | Wh | Active energy W1
Current measurement module group 11 | 20440 | Float | RD | Wh | Active energy W2
Current measurement module group 11 | 20442 | Float | RD | Wh | Active energy W3
Current measurement module group 11 | 20444 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 11 | 20446 | Float | RD | Wh | Active energy W1, related
Current measurement module group 11 | 20448 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 11 | 20450 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 11 | 20452 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 11 | 20454 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 11 | 20456 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 11 | 20458 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 11 | 20460 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 11 | 20462 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 11 | 20464 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 11 | 20466 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 11 | 20468 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 11 | 20470 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 11 | 20472 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 11 | 20474 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 11 | 20476 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 11 | 20478 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 11 | 20480 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 11 | 20482 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 11 | 20484 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 11 | 20486 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 11 | 20488 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 11 | 20490 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 11 | 20492 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 11 | 20494 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 11 | 20496 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 11 | 20498 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 12 | 20500 | Float | RD | A | Current, I1
Current measurement module group 12 | 20502 | Float | RD | A | Current, I2
Current measurement module group 12 | 20504 | Float | RD | A | Current, I3
Current measurement module group 12 | 20506 | Float | RD | A | Current, I4
Current measurement module group 12 | 20508 | Float | RD | W | Active power P1
Current measurement module group 12 | 20510 | Float | RD | W | Active power P2
Current measurement module group 12 | 20512 | Float | RD | W | Active power P3
Current measurement module group 12 | 20514 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 12 | 20516 | Float | RD | VA | Apparent power S1
Current measurement module group 12 | 20518 | Float | RD | VA | Apparent power S2
Current measurement module group 12 | 20520 | Float | RD | VA | Apparent power S3
Current measurement module group 12 | 20522 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 12 | 20524 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 12 | 20526 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 12 | 20528 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 12 | 20530 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 12 | 20532 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 12 | 20534 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 12 | 20536 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 12 | 20538 | Float | RD | Wh | Active energy W1
Current measurement module group 12 | 20540 | Float | RD | Wh | Active energy W2
Current measurement module group 12 | 20542 | Float | RD | Wh | Active energy W3
Current measurement module group 12 | 20544 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 12 | 20546 | Float | RD | Wh | Active energy W1, related
Current measurement module group 12 | 20548 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 12 | 20550 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 12 | 20552 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 12 | 20554 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 12 | 20556 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 12 | 20558 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 12 | 20560 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 12 | 20562 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 12 | 20564 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 12 | 20566 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 12 | 20568 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 12 | 20570 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 12 | 20572 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 12 | 20574 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 12 | 20576 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 12 | 20578 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 12 | 20580 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 12 | 20582 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 12 | 20584 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 12 | 20586 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 12 | 20588 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 12 | 20590 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 12 | 20592 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 12 | 20594 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 12 | 20596 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 12 | 20598 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 13 | 20600 | Float | RD | A | Current, I1
Current measurement module group 13 | 20602 | Float | RD | A | Current, I2
Current measurement module group 13 | 20604 | Float | RD | A | Current, I3
Current measurement module group 13 | 20606 | Float | RD | A | Current, I4
Current measurement module group 13 | 20608 | Float | RD | W | Active power P1
Current measurement module group 13 | 20610 | Float | RD | W | Active power P2
Current measurement module group 13 | 20612 | Float | RD | W | Active power P3
Current measurement module group 13 | 20614 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 13 | 20616 | Float | RD | VA | Apparent power S1
Current measurement module group 13 | 20618 | Float | RD | VA | Apparent power S2
Current measurement module group 13 | 20620 | Float | RD | VA | Apparent power S3
Current measurement module group 13 | 20622 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 13 | 20624 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 13 | 20626 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 13 | 20628 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 13 | 20630 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 13 | 20632 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 13 | 20634 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 13 | 20636 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 13 | 20638 | Float | RD | Wh | Active energy W1
Current measurement module group 13 | 20640 | Float | RD | Wh | Active energy W2
Current measurement module group 13 | 20642 | Float | RD | Wh | Active energy W3
Current measurement module group 13 | 20644 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 13 | 20646 | Float | RD | Wh | Active energy W1, related
Current measurement module group 13 | 20648 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 13 | 20650 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 13 | 20652 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 13 | 20654 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 13 | 20656 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 13 | 20658 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 13 | 20660 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 13 | 20662 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 13 | 20664 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 13 | 20666 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 13 | 20668 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 13 | 20670 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 13 | 20672 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 13 | 20674 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 13 | 20676 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 13 | 20678 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 13 | 20680 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 13 | 20682 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 13 | 20684 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 13 | 20686 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 13 | 20688 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 13 | 20690 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 13 | 20692 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 13 | 20694 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 13 | 20696 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 13 | 20698 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 14 | 20700 | Float | RD | A | Current, I1
Current measurement module group 14 | 20702 | Float | RD | A | Current, I2
Current measurement module group 14 | 20704 | Float | RD | A | Current, I3
Current measurement module group 14 | 20706 | Float | RD | A | Current, I4
Current measurement module group 14 | 20708 | Float | RD | W | Active power P1
Current measurement module group 14 | 20710 | Float | RD | W | Active power P2
Current measurement module group 14 | 20712 | Float | RD | W | Active power P3
Current measurement module group 14 | 20714 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 14 | 20716 | Float | RD | VA | Apparent power S1
Current measurement module group 14 | 20718 | Float | RD | VA | Apparent power S2
Current measurement module group 14 | 20720 | Float | RD | VA | Apparent power S3
Current measurement module group 14 | 20722 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 14 | 20724 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 14 | 20726 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 14 | 20728 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 14 | 20730 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 14 | 20732 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 14 | 20734 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 14 | 20736 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 14 | 20738 | Float | RD | Wh | Active energy W1
Current measurement module group 14 | 20740 | Float | RD | Wh | Active energy W2
Current measurement module group 14 | 20742 | Float | RD | Wh | Active energy W3
Current measurement module group 14 | 20744 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 14 | 20746 | Float | RD | Wh | Active energy W1, related
Current measurement module group 14 | 20748 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 14 | 20750 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 14 | 20752 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 14 | 20754 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 14 | 20756 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 14 | 20758 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 14 | 20760 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 14 | 20762 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 14 | 20764 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 14 | 20766 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 14 | 20768 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 14 | 20770 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 14 | 20772 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 14 | 20774 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 14 | 20776 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 14 | 20778 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 14 | 20780 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 14 | 20782 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 14 | 20784 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 14 | 20786 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 14 | 20788 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 14 | 20790 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 14 | 20792 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 14 | 20794 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 14 | 20796 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 14 | 20798 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 15 | 20800 | Float | RD | A | Current, I1
Current measurement module group 15 | 20802 | Float | RD | A | Current, I2
Current measurement module group 15 | 20804 | Float | RD | A | Current, I3
Current measurement module group 15 | 20806 | Float | RD | A | Current, I4
Current measurement module group 15 | 20808 | Float | RD | W | Active power P1
Current measurement module group 15 | 20810 | Float | RD | W | Active power P2
Current measurement module group 15 | 20812 | Float | RD | W | Active power P3
Current measurement module group 15 | 20814 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 15 | 20816 | Float | RD | VA | Apparent power S1
Current measurement module group 15 | 20818 | Float | RD | VA | Apparent power S2
Current measurement module group 15 | 20820 | Float | RD | VA | Apparent power S3
Current measurement module group 15 | 20822 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 15 | 20824 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 15 | 20826 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 15 | 20828 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 15 | 20830 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 15 | 20832 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 15 | 20834 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 15 | 20836 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 15 | 20838 | Float | RD | Wh | Active energy W1
Current measurement module group 15 | 20840 | Float | RD | Wh | Active energy W2
Current measurement module group 15 | 20842 | Float | RD | Wh | Active energy W3
Current measurement module group 15 | 20844 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 15 | 20846 | Float | RD | Wh | Active energy W1, related
Current measurement module group 15 | 20848 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 15 | 20850 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 15 | 20852 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 15 | 20854 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 15 | 20856 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 15 | 20858 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 15 | 20860 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 15 | 20862 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 15 | 20864 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 15 | 20866 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 15 | 20868 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 15 | 20870 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 15 | 20872 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 15 | 20874 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 15 | 20876 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 15 | 20878 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 15 | 20880 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 15 | 20882 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 15 | 20884 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 15 | 20886 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 15 | 20888 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 15 | 20890 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 15 | 20892 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 15 | 20894 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 15 | 20896 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 15 | 20898 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 16 | 20900 | Float | RD | A | Current, I1
Current measurement module group 16 | 20902 | Float | RD | A | Current, I2
Current measurement module group 16 | 20904 | Float | RD | A | Current, I3
Current measurement module group 16 | 20906 | Float | RD | A | Current, I4
Current measurement module group 16 | 20908 | Float | RD | W | Active power P1
Current measurement module group 16 | 20910 | Float | RD | W | Active power P2
Current measurement module group 16 | 20912 | Float | RD | W | Active power P3
Current measurement module group 16 | 20914 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 16 | 20916 | Float | RD | VA | Apparent power S1
Current measurement module group 16 | 20918 | Float | RD | VA | Apparent power S2
Current measurement module group 16 | 20920 | Float | RD | VA | Apparent power S3
Current measurement module group 16 | 20922 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 16 | 20924 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 16 | 20926 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 16 | 20928 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 16 | 20930 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 16 | 20932 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 16 | 20934 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 16 | 20936 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 16 | 20938 | Float | RD | Wh | Active energy W1
Current measurement module group 16 | 20940 | Float | RD | Wh | Active energy W2
Current measurement module group 16 | 20942 | Float | RD | Wh | Active energy W3
Current measurement module group 16 | 20944 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 16 | 20946 | Float | RD | Wh | Active energy W1, related
Current measurement module group 16 | 20948 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 16 | 20950 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 16 | 20952 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 16 | 20954 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 16 | 20956 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 16 | 20958 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 16 | 20960 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 16 | 20962 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 16 | 20964 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 16 | 20966 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 16 | 20968 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 16 | 20970 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 16 | 20972 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 16 | 20974 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 16 | 20976 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 16 | 20978 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 16 | 20980 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 16 | 20982 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 16 | 20984 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 16 | 20986 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 16 | 20988 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 16 | 20990 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 16 | 20992 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 16 | 20994 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 16 | 20996 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 16 | 20998 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 18 | 21000 | Float | RD | A | Current, I1
Current measurement module group 18 | 21002 | Float | RD | A | Current, I2
Current measurement module group 18 | 21004 | Float | RD | A | Current, I3
Current measurement module group 18 | 21006 | Float | RD | A | Current, I4
Current measurement module group 18 | 21008 | Float | RD | W | Active power P1
Current measurement module group 18 | 21010 | Float | RD | W | Active power P2
Current measurement module group 18 | 21012 | Float | RD | W | Active power P3
Current measurement module group 18 | 21014 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 18 | 21016 | Float | RD | VA | Apparent power S1
Current measurement module group 18 | 21018 | Float | RD | VA | Apparent power S2
Current measurement module group 18 | 21020 | Float | RD | VA | Apparent power S3
Current measurement module group 18 | 21022 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 18 | 21024 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 18 | 21026 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 18 | 21028 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 18 | 21030 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 18 | 21032 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 18 | 21034 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 18 | 21036 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 18 | 21038 | Float | RD | Wh | Active energy W1
Current measurement module group 18 | 21040 | Float | RD | Wh | Active energy W2
Current measurement module group 18 | 21042 | Float | RD | Wh | Active energy W3
Current measurement module group 18 | 21044 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 18 | 21046 | Float | RD | Wh | Active energy W1, related
Current measurement module group 18 | 21048 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 18 | 21050 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 18 | 21052 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 18 | 21054 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 18 | 21056 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 18 | 21058 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 18 | 21060 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 18 | 21062 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 18 | 21064 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 18 | 21066 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 18 | 21068 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 18 | 21070 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 18 | 21072 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 18 | 21074 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 18 | 21076 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 18 | 21078 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 18 | 21080 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 18 | 21082 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 18 | 21084 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 18 | 21086 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 18 | 21088 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 18 | 21090 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 18 | 21092 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 18 | 21094 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 18 | 21096 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 18 | 21098 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 18 | 21100 | Float | RD | A | Current, I1
Current measurement module group 18 | 21102 | Float | RD | A | Current, I2
Current measurement module group 18 | 21104 | Float | RD | A | Current, I3
Current measurement module group 18 | 21106 | Float | RD | A | Current, I4
Current measurement module group 18 | 21108 | Float | RD | W | Active power P1
Current measurement module group 18 | 21110 | Float | RD | W | Active power P2
Current measurement module group 18 | 21112 | Float | RD | W | Active power P3
Current measurement module group 18 | 21114 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 18 | 21116 | Float | RD | VA | Apparent power S1
Current measurement module group 18 | 21118 | Float | RD | VA | Apparent power S2
Current measurement module group 18 | 21120 | Float | RD | VA | Apparent power S3
Current measurement module group 18 | 21122 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 18 | 21124 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 18 | 21126 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 18 | 21128 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 18 | 21130 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 18 | 21132 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 18 | 21134 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 18 | 21136 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 18 | 21138 | Float | RD | Wh | Active energy W1
Current measurement module group 18 | 21140 | Float | RD | Wh | Active energy W2
Current measurement module group 18 | 21142 | Float | RD | Wh | Active energy W3
Current measurement module group 18 | 21144 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 18 | 21146 | Float | RD | Wh | Active energy W1, related
Current measurement module group 18 | 21148 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 18 | 21150 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 18 | 21152 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 18 | 21154 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 18 | 21156 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 18 | 21158 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 18 | 21160 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 18 | 21162 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 18 | 21164 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 18 | 21166 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 18 | 21168 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 18 | 21170 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 18 | 21172 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 18 | 21174 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 18 | 21176 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 18 | 21178 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 18 | 21180 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 18 | 21182 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 18 | 21184 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 18 | 21186 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 18 | 21188 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 18 | 21190 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 18 | 21192 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 18 | 21194 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 18 | 21196 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 18 | 21198 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 19 | 21200 | Float | RD | A | Current, I1
Current measurement module group 19 | 21202 | Float | RD | A | Current, I2
Current measurement module group 19 | 21204 | Float | RD | A | Current, I3
Current measurement module group 19 | 21206 | Float | RD | A | Current, I4
Current measurement module group 19 | 21208 | Float | RD | W | Active power P1
Current measurement module group 19 | 21210 | Float | RD | W | Active power P2
Current measurement module group 19 | 21212 | Float | RD | W | Active power P3
Current measurement module group 19 | 21214 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 19 | 21216 | Float | RD | VA | Apparent power S1
Current measurement module group 19 | 21218 | Float | RD | VA | Apparent power S2
Current measurement module group 19 | 21220 | Float | RD | VA | Apparent power S3
Current measurement module group 19 | 21222 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 19 | 21224 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 19 | 21226 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 19 | 21228 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 19 | 21230 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 19 | 21232 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 19 | 21234 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 19 | 21236 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 19 | 21238 | Float | RD | Wh | Active energy W1
Current measurement module group 19 | 21240 | Float | RD | Wh | Active energy W2
Current measurement module group 19 | 21242 | Float | RD | Wh | Active energy W3
Current measurement module group 19 | 21244 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 19 | 21246 | Float | RD | Wh | Active energy W1, related
Current measurement module group 19 | 21248 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 19 | 21250 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 19 | 21252 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 19 | 21254 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 19 | 21256 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 19 | 21258 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 19 | 21260 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 19 | 21262 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 19 | 21264 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 19 | 21266 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 19 | 21268 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 19 | 21270 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 19 | 21272 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 19 | 21274 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 19 | 21276 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 19 | 21278 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 19 | 21280 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 19 | 21282 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 19 | 21284 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 19 | 21286 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 19 | 21288 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 19 | 21290 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 19 | 21292 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 19 | 21294 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 19 | 21296 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 19 | 21298 | Float | RD | % | Harmonic, THD, I3
Current measurement module group 20 | 21300 | Float | RD | A | Current, I1
Current measurement module group 20 | 21302 | Float | RD | A | Current, I2
Current measurement module group 20 | 21304 | Float | RD | A | Current, I3
Current measurement module group 20 | 21306 | Float | RD | A | Current, I4
Current measurement module group 20 | 21308 | Float | RD | W | Active power P1
Current measurement module group 20 | 21310 | Float | RD | W | Active power P2
Current measurement module group 20 | 21312 | Float | RD | W | Active power P3
Current measurement module group 20 | 21314 | Float | RD | W | Sum; Psum3=P1+P2+P3
Current measurement module group 20 | 21316 | Float | RD | VA | Apparent power S1
Current measurement module group 20 | 21318 | Float | RD | VA | Apparent power S2
Current measurement module group 20 | 21320 | Float | RD | VA | Apparent power S3
Current measurement module group 20 | 21322 | Float | RD | VA | Sum; Ssum3=S1+S2+S3
Current measurement module group 20 | 21324 | Float | RD | var | Reactive power (fundamental frequency) Q1
Current measurement module group 20 | 21326 | Float | RD | var | Reactive power (fundamental frequency) Q2
Current measurement module group 20 | 21328 | Float | RD | var | Reactive power (fundamental frequency) Q3
Current measurement module group 20 | 21330 | Float | RD | var | Sum; Qsum3=Q1+Q2+Q3
Current measurement module group 20 | 21332 | Float | RD | - | Fund.power factor, Cos(ϕ); UL1 IL1
Current measurement module group 20 | 21334 | Float | RD | - | Fund.power factor, Cos(ϕ); UL2 IL2
Current measurement module group 20 | 21336 | Float | RD | - | Fund.power factor, Cos(ϕ); UL3 IL3
Current measurement module group 20 | 21338 | Float | RD | Wh | Active energy W1
Current measurement module group 20 | 21340 | Float | RD | Wh | Active energy W2
Current measurement module group 20 | 21342 | Float | RD | Wh | Active energy W3
Current measurement module group 20 | 21344 | Float | RD | Wh | Active energy W1..W3
Current measurement module group 20 | 21346 | Float | RD | Wh | Active energy W1, related
Current measurement module group 20 | 21348 | Float | RD | Wh | Active energy W2, drawn
Current measurement module group 20 | 21350 | Float | RD | Wh | Active energy W3, purchased
Current measurement module group 20 | 21352 | Float | RD | Wh | Active energy W1..W3, purchased, tariff 1
Current measurement module group 20 | 21354 | Float | RD | Wh | Active energy W1, supplied
Current measurement module group 20 | 21356 | Float | RD | Wh | Active energy W2, supplied
Current measurement module group 20 | 21358 | Float | RD | Wh | Active energy W3, supplied
Current measurement module group 20 | 21360 | Float | RD | Wh | Active energy W1..W3, supplied
Current measurement module group 20 | 21362 | Float | RD | VAh | Apparent energy WS1
Current measurement module group 20 | 21364 | Float | RD | VAh | Apparent energy WS2
Current measurement module group 20 | 21366 | Float | RD | VAh | Apparent energy WS3
Current measurement module group 20 | 21368 | Float | RD | VAh | Apparent energy WS1..WS3
Current measurement module group 20 | 21370 | Float | RD | varh | Reactive energy WQ1
Current measurement module group 20 | 21372 | Float | RD | varh | Reactive energy WQ2
Current measurement module group 20 | 21374 | Float | RD | varh | Reactive energy WQ3
Current measurement module group 20 | 21376 | Float | RD | varh | Reactive energy WQ1..WQ3
Current measurement module group 20 | 21378 | Float | RD | varh | Reactive energy WQ1, inductive
Current measurement module group 20 | 21380 | Float | RD | varh | Reactive energy WQ2, inductive
Current measurement module group 20 | 21382 | Float | RD | varh | Reactive energy WQ3, inductive
Current measurement module group 20 | 21384 | Float | RD | varh | Reactive energy WQ1..WQ3, inductive
Current measurement module group 20 | 21386 | Float | RD | varh | Reactive energy WQ1, capacitive
Current measurement module group 20 | 21388 | Float | RD | varh | Reactive energy WQ2, capacitive
Current measurement module group 20 | 21390 | Float | RD | varh | Reactive energy WQ3, capacitive
Current measurement module group 20 | 21392 | Float | RD | varh | Reactive energy WQ1..WQ3, capacitive
Current measurement module group 20 | 21394 | Float | RD | % | Harmonics, THD, I1
Current measurement module group 20 | 21396 | Float | RD | % | Harmonic, THD, I2
Current measurement module group 20 | 21398 | Float | RD | % | Harmonic, THD, I3
basic device | 21400 | Boolean | RD | - | Digital-In1, status
basic device | 21401 | Boolean | RD | - | Digital-In2, status
basic device | 21402 | Boolean | RD | - | Digital-In3, status
basic device | 21403 | Boolean | RD | - | Digital-In4, status
basic device | 21404 | Float | RD | - | Digital In1, S0 counter status
basic device | 21406 | Float | RD | - | Digital In2, S0 counter status
basic device | 21408 | Float | RD | - | Digital In3, S0 counter reading
basic device | 21410 | Float | RD | - | Digital In4, S0 counter reading
basic device | 21412 | UInt32 | RD | - | Digital In1, pulse counter status
basic device | 21414 | UInt32 | RD | - | Digital In2, pulse counter reading
basic device | 21416 | UInt32 | RD | - | Digital-In3, pulse counter reading
basic device | 21418 | UInt32 | RD | - | Digital-In4, pulse counter reading
basic device | 21420 | float | RD | °C | Temperature multifunction input1
basic device | 21422 | float | RD | °C | Temperature multifunction input2
basic device | 21424 | float | RD | °C | Temperature multifunction input3
basic device | 21426 | float | RD | °C | Temperature multifunction input4
basic device | 21427 | short | RD | A | RCM limit value warning I1
basic device | 21428 | short | RD | A | RCM limit warning I2
basic device | 21429 | short | RD | A | RCM limit value warning I3
basic device | 21430 | short | RD | A | RCM limit value warning I4
basic device | 21431 | short | RD | A | RCM limit value exceeded I1
basic device | 21432 | short | RD | A | RCM limit value exceeded I2
basic device | 21433 | short | RD | A | RCM limit value exceeded I3
basic device | 21434 | short | RD | A | RCM limit value exceeded I4
basic device | 21435 | short | RD | - | Alarm bit, limit value exceeded I1, manually resettable
basic device | 21436 | short | RD | - | Alarm bit, limit value exceeded I2, manually resettable
basic device | 21437 | short | RD | - | Alarm bit, limit value exceeded I3, manually resettable
basic device | 21438 | short | RD | - | Alarm bit, limit value exceeded I4, manually resettable
basic device | 21439 | short | RD | - | Wire break detection I1
basic device | 21440 | short | RD | - | Wire break detection I2
basic device | 21441 | short | RD | - | Wire break detection I3
basic device | 21442 | short | RD | - | Wire break detection I4
basic device | 21444 | UInt16 | WR | - | The external synchronisation signal, write to non zero to trigger new MaxAvg interval
basic device | 21445 | Float | RD | % | Voltage Symmetric Unbalance in percent.
basic device | 21447 | Float | RD | V | Average Voltage ULN
basic device | 21449 | Float | RD | V | Average Voltage ULL
basic device | 21451 | Float | RD | var | Reactive power measuring group 1 basic device
basic device | 21453 | Float | RD | var | Reactive power measuring group 2 basic device
basic device | 21455 | Float | RD | var | Reactive power measuring group 3 basic device
basic device | 21457 | Float | RD | W | Fundamental real power measuring group 1 basic device
basic device | 21459 | Float | RD | W | Fundamental real power measuring group 2 basic device
basic device | 21461 | Float | RD | W | Fundamental real power measuring group 3 basic device
basic device | 21463 | Float | RD | - | Power factor measuring group 1 basic device
basic device | 21465 | Float | RD | - | Power factor measuring group 2 basic device
basic device | 21467 | Float | RD | - | Power factor measuring group 3 basic device
basic device | 21469 | Float | RD | - | Fundamental power factor (CosPhi) measuring group 1 basic device
basic device | 21471 | Float | RD | - | Fundamental power factor (CosPhi) measuring group 2 basic device
basic device | 21473 | Float | RD | - | Fundamental power factor (CosPhi) measuring group 3 basic device
basic device | 21475 | Float | RD | % | Current Unbalance measuring group 1 basic device
basic device | 21477 | Float | RD | % | Current Unbalance measuring group 2 basic device
basic device | 21479 | Float | RD | % | Current Unbalance measuring group 3 basic device
Basic device measuring group 01 | 21500 | Float | RD | W | Active power P4
Basic device measuring group 01 | 21502 | Float | RD | VA | Apparent power S4
Basic device measuring group 01 | 21504 | Float | RD | var | Reactive power (fundamental frequency) Q4
Basic device measuring group 01 | 21506 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Basic device measuring group 01 | 21508 | Float | RD | Wh | Active energy W4
Basic device measuring group 01 | 21510 | Float | RD | Wh | Active energy W4, purchased
Basic device measuring group 01 | 21512 | Float | RD | Wh | Active energy W4, supplied
Basic device measuring group 01 | 21514 | Float | RD | VAh | Apparent energy WS4
Basic device measuring group 01 | 21516 | Float | RD | varh | Reactive energy WQ4
Basic device measuring group 01 | 21518 | Float | RD | varh | Reactive energy WQ4, inductive
Basic device measuring group 01 | 21520 | Float | RD | varh | Reactive energy WQ4, capacitive
Basic device measuring group 01 | 21522 | Float | RD | % | Harmonics, THD, I4
Basic device measuring group 02 | 21524 | Float | RD | W | Active power P4
Basic device measuring group 02 | 21526 | Float | RD | VA | Apparent power S4
Basic device measuring group 02 | 21528 | Float | RD | var | Reactive power (fundamental frequency) Q4
Basic device measuring group 02 | 21530 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Basic device measuring group 02 | 21532 | Float | RD | Wh | Active energy W4
Basic device measuring group 02 | 21534 | Float | RD | Wh | Active energy W4, purchased
Basic device measuring group 02 | 21536 | Float | RD | Wh | Active energy W4, supplied
Basic device measuring group 02 | 21538 | Float | RD | VAh | Apparent energy WS4
Basic device measuring group 02 | 21540 | Float | RD | varh | Reactive energy WQ4
Basic device measuring group 02 | 21542 | Float | RD | varh | Reactive energy WQ4, inductive
Basic device measuring group 02 | 21544 | Float | RD | varh | Reactive energy WQ4, capacitive
Basic device measuring group 02 | 21546 | Float | RD | % | Harmonics, THD, I4
Basic device measuring group 03 | 21548 | Float | RD | W | Active power P4
Basic device measuring group 03 | 21550 | Float | RD | VA | Apparent power S4
Basic device measuring group 03 | 21552 | Float | RD | var | Reactive power (fundamental frequency) Q4
Basic device measuring group 03 | 21554 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Basic device measuring group 03 | 21556 | Float | RD | Wh | Active energy W4
Basic device measuring group 03 | 21558 | Float | RD | Wh | Active energy W4, purchased
Basic device measuring group 03 | 21560 | Float | RD | Wh | Active energy W4, supplied
Basic device measuring group 03 | 21562 | Float | RD | VAh | Apparent energy WS4
Basic device measuring group 03 | 21564 | Float | RD | varh | Reactive energy WQ4
Basic device measuring group 03 | 21566 | Float | RD | varh | Reactive energy WQ4, inductive
Basic device measuring group 03 | 21568 | Float | RD | varh | Reactive energy WQ4, capacitive
Basic device measuring group 03 | 21570 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 01 | 21572 | Float | RD | W | Active power P4
Current measurement module group 01 | 21574 | Float | RD | VA | Apparent power S4
Current measurement module group 01 | 21576 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 01 | 21578 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 01 | 21580 | Float | RD | Wh | Active energy W4
Current measurement module group 01 | 21582 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 01 | 21584 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 01 | 21586 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 01 | 21588 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 01 | 21590 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 01 | 21592 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 01 | 21594 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 02 | 21596 | Float | RD | W | Active power P4
Current measurement module group 02 | 21598 | Float | RD | VA | Apparent power S4
Current measurement module group 02 | 21600 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 02 | 21602 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 02 | 21604 | Float | RD | Wh | Active energy W4
Current measurement module group 02 | 21606 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 02 | 21608 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 02 | 21610 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 02 | 21612 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 02 | 21614 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 02 | 21616 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 02 | 21618 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 03 | 21620 | Float | RD | W | Active power P4
Current measurement module group 03 | 21622 | Float | RD | VA | Apparent power S4
Current measurement module group 03 | 21624 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 03 | 21626 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 03 | 21628 | Float | RD | Wh | Active energy W4
Current measurement module group 03 | 21630 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 03 | 21632 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 03 | 21634 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 03 | 21636 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 03 | 21638 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 03 | 21640 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 03 | 21642 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 04 | 21644 | Float | RD | W | Active power P4
Current measurement module group 04 | 21646 | Float | RD | VA | Apparent power S4
Current measurement module group 04 | 21648 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 04 | 21650 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 04 | 21652 | Float | RD | Wh | Active energy W4
Current measurement module group 04 | 21654 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 04 | 21656 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 04 | 21658 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 04 | 21660 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 04 | 21662 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 04 | 21664 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 04 | 21666 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 05 | 21668 | Float | RD | W | Active power P4
Current measurement module group 05 | 21670 | Float | RD | VA | Apparent power S4
Current measurement module group 05 | 21672 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 05 | 21674 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 05 | 21676 | Float | RD | Wh | Active energy W4
Current measurement module group 05 | 21678 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 05 | 21680 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 05 | 21682 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 05 | 21684 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 05 | 21686 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 05 | 21688 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 05 | 21690 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 06 | 21692 | Float | RD | W | Active power P4
Current measurement module group 06 | 21694 | Float | RD | VA | Apparent power S4
Current measurement module group 06 | 21696 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 06 | 21698 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 06 | 21700 | Float | RD | Wh | Active energy W4
Current measurement module group 06 | 21702 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 06 | 21704 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 06 | 21706 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 06 | 21708 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 06 | 21710 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 06 | 21712 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 06 | 21714 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 07 | 21716 | Float | RD | W | Active power P4
Current measurement module group 07 | 21718 | Float | RD | VA | Apparent power S4
Current measurement module group 07 | 21720 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 07 | 21722 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 07 | 21724 | Float | RD | Wh | Active energy W4
Current measurement module group 07 | 21726 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 07 | 21728 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 07 | 21730 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 07 | 21732 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 07 | 21734 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 07 | 21736 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 07 | 21738 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 08 | 21740 | Float | RD | W | Active power P4
Current measurement module group 08 | 21742 | Float | RD | VA | Apparent power S4
Current measurement module group 08 | 21744 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 08 | 21746 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 08 | 21748 | Float | RD | Wh | Active energy W4
Current measurement module group 08 | 21750 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 08 | 21752 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 08 | 21754 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 08 | 21756 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 08 | 21758 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 08 | 21760 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 08 | 21762 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 09 | 21764 | Float | RD | W | Active power P4
Current measurement module group 09 | 21766 | Float | RD | VA | Apparent power S4
Current measurement module group 09 | 21768 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 09 | 21770 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 09 | 21772 | Float | RD | Wh | Active energy W4
Current measurement module group 09 | 21774 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 09 | 21776 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 09 | 21778 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 09 | 21780 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 09 | 21782 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 09 | 21784 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 09 | 21786 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 10 | 21788 | Float | RD | W | Active power P4
Current measurement module group 10 | 21790 | Float | RD | VA | Apparent power S4
Current measurement module group 10 | 21792 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 10 | 21794 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 10 | 21796 | Float | RD | Wh | Active energy W4
Current measurement module group 10 | 21798 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 10 | 21800 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 10 | 21802 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 10 | 21804 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 10 | 21806 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 10 | 21808 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 10 | 21810 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 11 | 21812 | Float | RD | W | Active power P4
Current measurement module group 11 | 21814 | Float | RD | VA | Apparent power S4
Current measurement module group 11 | 21816 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 11 | 21818 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 11 | 21820 | Float | RD | Wh | Active energy W4
Current measurement module group 11 | 21822 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 11 | 21824 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 11 | 21826 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 11 | 21828 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 11 | 21830 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 11 | 21832 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 11 | 21834 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 12 | 21836 | Float | RD | W | Active power P4
Current measurement module group 12 | 21838 | Float | RD | VA | Apparent power S4
Current measurement module group 12 | 21840 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 12 | 21842 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 12 | 21844 | Float | RD | Wh | Active energy W4
Current measurement module group 12 | 21846 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 12 | 21848 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 12 | 21850 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 12 | 21852 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 12 | 21854 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 12 | 21856 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 12 | 21858 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 13 | 21860 | Float | RD | W | Active power P4
Current measurement module group 13 | 21862 | Float | RD | VA | Apparent power S4
Current measurement module group 13 | 21864 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 13 | 21866 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 13 | 21868 | Float | RD | Wh | Active energy W4
Current measurement module group 13 | 21870 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 13 | 21872 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 13 | 21874 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 13 | 21876 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 13 | 21878 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 13 | 21880 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 13 | 21882 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 14 | 21884 | Float | RD | W | Active power P4
Current measurement module group 14 | 21886 | Float | RD | VA | Apparent power S4
Current measurement module group 14 | 21888 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 14 | 21890 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 14 | 21892 | Float | RD | Wh | Active energy W4
Current measurement module group 14 | 21894 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 14 | 21896 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 14 | 21898 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 14 | 21900 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 14 | 21902 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 14 | 21904 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 14 | 21906 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 15 | 21908 | Float | RD | W | Active power P4
Current measurement module group 15 | 21910 | Float | RD | VA | Apparent power S4
Current measurement module group 15 | 21912 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 15 | 21914 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 15 | 21916 | Float | RD | Wh | Active energy W4
Current measurement module group 15 | 21918 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 15 | 21920 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 15 | 21922 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 15 | 21924 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 15 | 21926 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 15 | 21928 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 15 | 21930 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 16 | 21932 | Float | RD | W | Active power P4
Current measurement module group 16 | 21934 | Float | RD | VA | Apparent power S4
Current measurement module group 16 | 21936 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 16 | 21938 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 16 | 21940 | Float | RD | Wh | Active energy W4
Current measurement module group 16 | 21942 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 16 | 21944 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 16 | 21946 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 16 | 21948 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 16 | 21950 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 16 | 21952 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 16 | 21954 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 17 | 21956 | Float | RD | W | Active power P4
Current measurement module group 17 | 21958 | Float | RD | VA | Apparent power S4
Current measurement module group 17 | 21960 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 17 | 21962 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 17 | 21964 | Float | RD | Wh | Active energy W4
Current measurement module group 17 | 21966 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 17 | 21968 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 17 | 21970 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 17 | 21972 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 17 | 21974 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 17 | 21976 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 17 | 21978 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 18 | 21980 | Float | RD | W | Active power P4
Current measurement module group 18 | 21982 | Float | RD | VA | Apparent power S4
Current measurement module group 18 | 21984 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 18 | 21986 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 18 | 21988 | Float | RD | Wh | Active energy W4
Current measurement module group 18 | 21990 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 18 | 21992 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 18 | 21994 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 18 | 21996 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 18 | 21998 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 18 | 22000 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 18 | 22002 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 19 | 22004 | Float | RD | W | Active power P4
Current measurement module group 19 | 22006 | Float | RD | VA | Apparent power S4
Current measurement module group 19 | 22008 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 19 | 22010 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 19 | 22012 | Float | RD | Wh | Active energy W4
Current measurement module group 19 | 22014 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 19 | 22016 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 19 | 22018 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 19 | 22020 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 19 | 22022 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 19 | 22024 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 19 | 22026 | Float | RD | % | Harmonics, THD, I4
Current measurement module group 20 | 22028 | Float | RD | W | Active power P4
Current measurement module group 20 | 22030 | Float | RD | VA | Apparent power S4
Current measurement module group 20 | 22032 | Float | RD | var | Reactive power (fundamental frequency) Q4
Current measurement module group 20 | 22034 | Float | RD | - | Fund.power factor, Cos(ϕ); UL4 IL4
Current measurement module group 20 | 22036 | Float | RD | Wh | Active energy W4
Current measurement module group 20 | 22038 | Float | RD | Wh | Active energy W4, purchased
Current measurement module group 20 | 22040 | Float | RD | Wh | Active energy W4, supplied
Current measurement module group 20 | 22042 | Float | RD | VAh | Apparent energy WS4
Current measurement module group 20 | 22044 | Float | RD | varh | Reactive energy WQ4
Current measurement module group 20 | 22046 | Float | RD | varh | Reactive energy WQ4, inductive
Current measurement module group 20 | 22048 | Float | RD | varh | Reactive energy WQ4, capacitive
Current measurement module group 20 | 22050 | Float | RD | % | Harmonics, THD, I4
Basic device measuring group 01 | 22200 | Float | RD | var | harmonic distortion power L1
Basic device measuring group 01 | 22202 | Float | RD | var | harmonic distortion power L2
Basic device measuring group 01 | 22204 | Float | RD | var | harmonic distortion power L3
Basic device measuring group 01 | 22208 | Float | RD | - | crest factor L1
Basic device measuring group 01 | 22210 | Float | RD | - | crest factor L2
Basic device measuring group 01 | 22212 | Float | RD | - | crest factor L3
Basic device measuring group 01 | 22214 | Float | RD | - | crest factor L4
Basic device measuring group 01 | 22216 | Float | RD | % | load current percentages of the rated current I1
Basic device measuring group 01 | 22218 | Float | RD | % | load current percentages of the rated current I2
Basic device measuring group 01 | 22220 | Float | RD | % | load current percentages of the rated current I3
Basic device measuring group 01 | 22222 | Float | RD | % | load current percentages of the rated current I4
Basic device measuring group 01 | 22224 | Float | RD | % | The total demand distortion TDD IL1
Basic device measuring group 01 | 22226 | Float | RD | % | The total demand distortion TDD IL2
Basic device measuring group 01 | 22228 | Float | RD | % | The total demand distortion TDD IL3
Basic device measuring group 01 | 22230 | Float | RD | % | The total demand distortion TDD IL4
Basic device measuring group 01 | 22232 | Float | RD | W | The fundamental real power L1
Basic device measuring group 01 | 22234 | Float | RD | W | The fundamental real power L2
Basic device measuring group 01 | 22236 | Float | RD | W | The fundamental real power L3
Basic device measuring group 01 | 22240 | Float | RD | - | The power factor L1
Basic device measuring group 01 | 22242 | Float | RD | - | The power factor L2
Basic device measuring group 01 | 22244 | Float | RD | - | The power factor L3
Basic device measuring group 01 | 22248 | Float | RD | var | The reactive power L1
Basic device measuring group 01 | 22250 | Float | RD | var | The reactive power L2
Basic device measuring group 01 | 22252 | Float | RD | var | The reactive power L3
Basic device measuring group 01 | 22256 | Float | RD | A | The imaginary part of current IL1
Basic device measuring group 01 | 22258 | Float | RD | A | The imaginary part of current IL2
Basic device measuring group 01 | 22260 | Float | RD | A | The imaginary part of current IL3
Basic device measuring group 01 | 22262 | Float | RD | A | The imaginary part of current IL4
Basic device measuring group 01 | 22264 | Float | RD | A | The real part of current IL1
Basic device measuring group 01 | 22266 | Float | RD | A | The real part of current IL2
Basic device measuring group 01 | 22268 | Float | RD | A | The real part of current IL3
Basic device measuring group 01 | 22270 | Float | RD | A | The real part of current IL4
Basic device measuring group 02 | 22272 | Float | RD | var | harmonic distortion power L1
Basic device measuring group 02 | 22274 | Float | RD | var | harmonic distortion power L2
Basic device measuring group 02 | 22276 | Float | RD | var | harmonic distortion power L3
Basic device measuring group 02 | 22280 | Float | RD | - | crest factor L1
Basic device measuring group 02 | 22282 | Float | RD | - | crest factor L2
Basic device measuring group 02 | 22284 | Float | RD | - | crest factor L3
Basic device measuring group 02 | 22286 | Float | RD | - | crest factor L4
Basic device measuring group 02 | 22288 | Float | RD | % | load current percentages of the rated current I1
Basic device measuring group 02 | 22290 | Float | RD | % | load current percentages of the rated current I2
Basic device measuring group 02 | 22292 | Float | RD | % | load current percentages of the rated current I3
Basic device measuring group 02 | 22294 | Float | RD | % | load current percentages of the rated current I4
Basic device measuring group 02 | 22296 | Float | RD | % | The total demand distortion TDD IL1
Basic device measuring group 02 | 22298 | Float | RD | % | The total demand distortion TDD IL2
Basic device measuring group 02 | 22300 | Float | RD | % | The total demand distortion TDD IL3
Basic device measuring group 02 | 22302 | Float | RD | % | The total demand distortion TDD IL4
Basic device measuring group 02 | 22304 | Float | RD | W | The fundamental real power L1
Basic device measuring group 02 | 22306 | Float | RD | W | The fundamental real power L2
Basic device measuring group 02 | 22308 | Float | RD | W | The fundamental real power L3
Basic device measuring group 02 | 22312 | Float | RD | - | The power factor L1
Basic device measuring group 02 | 22314 | Float | RD | - | The power factor L2
Basic device measuring group 02 | 22316 | Float | RD | - | The power factor L3
Basic device measuring group 02 | 22320 | Float | RD | var | The reactive power L1
Basic device measuring group 02 | 22322 | Float | RD | var | The reactive power L2
Basic device measuring group 02 | 22324 | Float | RD | var | The reactive power L3
Basic device measuring group 02 | 22328 | Float | RD | A | The imaginary part of current IL1
Basic device measuring group 02 | 22330 | Float | RD | A | The imaginary part of current IL2
Basic device measuring group 02 | 22332 | Float | RD | A | The imaginary part of current IL3
Basic device measuring group 02 | 22334 | Float | RD | A | The imaginary part of current IL4
Basic device measuring group 02 | 22336 | Float | RD | A | The real part of current IL1
Basic device measuring group 02 | 22338 | Float | RD | A | The real part of current IL2
Basic device measuring group 02 | 22340 | Float | RD | A | The real part of current IL3
Basic device measuring group 02 | 22342 | Float | RD | A | The real part of current IL4
Basic device measuring group 03 | 22344 | Float | RD | var | harmonic distortion power L1
Basic device measuring group 03 | 22346 | Float | RD | var | harmonic distortion power L2
Basic device measuring group 03 | 22348 | Float | RD | var | harmonic distortion power L3
Basic device measuring group 03 | 22352 | Float | RD | - | crest factor L1
Basic device measuring group 03 | 22354 | Float | RD | - | crest factor L2
Basic device measuring group 03 | 22356 | Float | RD | - | crest factor L3
Basic device measuring group 03 | 22358 | Float | RD | - | crest factor L4
Basic device measuring group 03 | 22360 | Float | RD | % | load current percentages of the rated current I1
Basic device measuring group 03 | 22362 | Float | RD | % | load current percentages of the rated current I2
Basic device measuring group 03 | 22364 | Float | RD | % | load current percentages of the rated current I3
Basic device measuring group 03 | 22366 | Float | RD | % | load current percentages of the rated current I4
Basic device measuring group 03 | 22368 | Float | RD | % | The total demand distortion TDD IL1
Basic device measuring group 03 | 22370 | Float | RD | % | The total demand distortion TDD IL2
Basic device measuring group 03 | 22372 | Float | RD | % | The total demand distortion TDD IL3
Basic device measuring group 03 | 22374 | Float | RD | % | The total demand distortion TDD IL4
Basic device measuring group 03 | 22376 | Float | RD | W | The fundamental real power L1
Basic device measuring group 03 | 22378 | Float | RD | W | The fundamental real power L2
Basic device measuring group 03 | 22380 | Float | RD | W | The fundamental real power L3
Basic device measuring group 03 | 22384 | Float | RD | - | The power factor L1
Basic device measuring group 03 | 22386 | Float | RD | - | The power factor L2
Basic device measuring group 03 | 22388 | Float | RD | - | The power factor L3
Basic device measuring group 03 | 22392 | Float | RD | var | The reactive power L1
Basic device measuring group 03 | 22394 | Float | RD | var | The reactive power L2
Basic device measuring group 03 | 22396 | Float | RD | var | The reactive power L3
Basic device measuring group 03 | 22400 | Float | RD | A | The imaginary part of current IL1
Basic device measuring group 03 | 22402 | Float | RD | A | The imaginary part of current IL2
Basic device measuring group 03 | 22404 | Float | RD | A | The imaginary part of current IL3
Basic device measuring group 03 | 22406 | Float | RD | A | The imaginary part of current IL4
Basic device measuring group 03 | 22408 | Float | RD | A | The real part of current IL1
Basic device measuring group 03 | 22410 | Float | RD | A | The real part of current IL2
Basic device measuring group 03 | 22412 | Float | RD | A | The real part of current IL3
Basic device measuring group 03 | 22414 | Float | RD | A | The real part of current IL4
Current measurement module group 01 | 22416 | Float | RD | var | harmonic distortion power L1
Current measurement module group 01 | 22418 | Float | RD | var | harmonic distortion power L2
Current measurement module group 01 | 22420 | Float | RD | var | harmonic distortion power L3
Current measurement module group 01 | 22424 | Float | RD | - | crest factor L1
Current measurement module group 01 | 22426 | Float | RD | - | crest factor L2
Current measurement module group 01 | 22428 | Float | RD | - | crest factor L3
Current measurement module group 01 | 22430 | Float | RD | - | crest factor L4
Current measurement module group 01 | 22432 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 01 | 22434 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 01 | 22436 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 01 | 22438 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 01 | 22440 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 01 | 22442 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 01 | 22444 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 01 | 22446 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 01 | 22448 | Float | RD | W | The fundamental real power L1
Current measurement module group 01 | 22450 | Float | RD | W | The fundamental real power L2
Current measurement module group 01 | 22452 | Float | RD | W | The fundamental real power L3
Current measurement module group 01 | 22456 | Float | RD | - | The power factor L1
Current measurement module group 01 | 22458 | Float | RD | - | The power factor L2
Current measurement module group 01 | 22460 | Float | RD | - | The power factor L3
Current measurement module group 01 | 22464 | Float | RD | var | The reactive power L1
Current measurement module group 01 | 22466 | Float | RD | var | The reactive power L2
Current measurement module group 01 | 22468 | Float | RD | var | The reactive power L3
Current measurement module group 01 | 22472 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 01 | 22474 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 01 | 22476 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 01 | 22478 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 01 | 22480 | Float | RD | A | The real part of current IL1
Current measurement module group 01 | 22482 | Float | RD | A | The real part of current IL2
Current measurement module group 01 | 22484 | Float | RD | A | The real part of current IL3
Current measurement module group 01 | 22486 | Float | RD | A | The real part of current IL4
Current measurement module group 02 | 22488 | Float | RD | var | harmonic distortion power L1
Current measurement module group 02 | 22490 | Float | RD | var | harmonic distortion power L2
Current measurement module group 02 | 22492 | Float | RD | var | harmonic distortion power L3
Current measurement module group 02 | 22496 | Float | RD | - | crest factor L1
Current measurement module group 02 | 22498 | Float | RD | - | crest factor L2
Current measurement module group 02 | 22500 | Float | RD | - | crest factor L3
Current measurement module group 02 | 22502 | Float | RD | - | crest factor L4
Current measurement module group 02 | 22504 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 02 | 22506 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 02 | 22508 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 02 | 22510 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 02 | 22512 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 02 | 22514 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 02 | 22516 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 02 | 22518 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 02 | 22520 | Float | RD | W | The fundamental real power L1
Current measurement module group 02 | 22522 | Float | RD | W | The fundamental real power L2
Current measurement module group 02 | 22524 | Float | RD | W | The fundamental real power L3
Current measurement module group 02 | 22528 | Float | RD | - | The power factor L1
Current measurement module group 02 | 22530 | Float | RD | - | The power factor L2
Current measurement module group 02 | 22532 | Float | RD | - | The power factor L3
Current measurement module group 02 | 22536 | Float | RD | var | The reactive power L1
Current measurement module group 02 | 22538 | Float | RD | var | The reactive power L2
Current measurement module group 02 | 22540 | Float | RD | var | The reactive power L3
Current measurement module group 02 | 22544 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 02 | 22546 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 02 | 22548 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 02 | 22550 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 02 | 22552 | Float | RD | A | The real part of current IL1
Current measurement module group 02 | 22554 | Float | RD | A | The real part of current IL2
Current measurement module group 02 | 22556 | Float | RD | A | The real part of current IL3
Current measurement module group 02 | 22558 | Float | RD | A | The real part of current IL4
Current measurement module group 03 | 22560 | Float | RD | var | harmonic distortion power L1
Current measurement module group 03 | 22562 | Float | RD | var | harmonic distortion power L2
Current measurement module group 03 | 22564 | Float | RD | var | harmonic distortion power L3
Current measurement module group 03 | 22568 | Float | RD | - | crest factor L1
Current measurement module group 03 | 22570 | Float | RD | - | crest factor L2
Current measurement module group 03 | 22572 | Float | RD | - | crest factor L3
Current measurement module group 03 | 22574 | Float | RD | - | crest factor L4
Current measurement module group 03 | 22576 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 03 | 22578 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 03 | 22580 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 03 | 22582 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 03 | 22584 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 03 | 22586 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 03 | 22588 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 03 | 22590 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 03 | 22592 | Float | RD | W | The fundamental real power L1
Current measurement module group 03 | 22594 | Float | RD | W | The fundamental real power L2
Current measurement module group 03 | 22596 | Float | RD | W | The fundamental real power L3
Current measurement module group 03 | 22600 | Float | RD | - | The power factor L1
Current measurement module group 03 | 22602 | Float | RD | - | The power factor L2
Current measurement module group 03 | 22604 | Float | RD | - | The power factor L3
Current measurement module group 03 | 22608 | Float | RD | var | The reactive power L1
Current measurement module group 03 | 22610 | Float | RD | var | The reactive power L2
Current measurement module group 03 | 22612 | Float | RD | var | The reactive power L3
Current measurement module group 03 | 22616 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 03 | 22618 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 03 | 22620 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 03 | 22622 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 03 | 22624 | Float | RD | A | The real part of current IL1
Current measurement module group 03 | 22626 | Float | RD | A | The real part of current IL2
Current measurement module group 03 | 22628 | Float | RD | A | The real part of current IL3
Current measurement module group 03 | 22630 | Float | RD | A | The real part of current IL4
Current measurement module group 04 | 22632 | Float | RD | var | harmonic distortion power L1
Current measurement module group 04 | 22634 | Float | RD | var | harmonic distortion power L2
Current measurement module group 04 | 22636 | Float | RD | var | harmonic distortion power L3
Current measurement module group 04 | 22640 | Float | RD | - | crest factor L1
Current measurement module group 04 | 22642 | Float | RD | - | crest factor L2
Current measurement module group 04 | 22644 | Float | RD | - | crest factor L3
Current measurement module group 04 | 22646 | Float | RD | - | crest factor L4
Current measurement module group 04 | 22648 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 04 | 22650 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 04 | 22652 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 04 | 22654 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 04 | 22656 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 04 | 22658 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 04 | 22660 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 04 | 22662 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 04 | 22664 | Float | RD | W | The fundamental real power L1
Current measurement module group 04 | 22666 | Float | RD | W | The fundamental real power L2
Current measurement module group 04 | 22668 | Float | RD | W | The fundamental real power L3
Current measurement module group 04 | 22672 | Float | RD | - | The power factor L1
Current measurement module group 04 | 22674 | Float | RD | - | The power factor L2
Current measurement module group 04 | 22676 | Float | RD | - | The power factor L3
Current measurement module group 04 | 22680 | Float | RD | var | The reactive power L1
Current measurement module group 04 | 22682 | Float | RD | var | The reactive power L2
Current measurement module group 04 | 22684 | Float | RD | var | The reactive power L3
Current measurement module group 04 | 22688 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 04 | 22690 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 04 | 22692 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 04 | 22694 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 04 | 22696 | Float | RD | A | The real part of current IL1
Current measurement module group 04 | 22698 | Float | RD | A | The real part of current IL2
Current measurement module group 04 | 22700 | Float | RD | A | The real part of current IL3
Current measurement module group 04 | 22702 | Float | RD | A | The real part of current IL4
Current measurement module group 05 | 22704 | Float | RD | var | harmonic distortion power L1
Current measurement module group 05 | 22706 | Float | RD | var | harmonic distortion power L2
Current measurement module group 05 | 22708 | Float | RD | var | harmonic distortion power L3
Current measurement module group 05 | 22712 | Float | RD | - | crest factor L1
Current measurement module group 05 | 22714 | Float | RD | - | crest factor L2
Current measurement module group 05 | 22716 | Float | RD | - | crest factor L3
Current measurement module group 05 | 22718 | Float | RD | - | crest factor L4
Current measurement module group 05 | 22720 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 05 | 22722 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 05 | 22724 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 05 | 22726 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 05 | 22728 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 05 | 22730 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 05 | 22732 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 05 | 22734 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 05 | 22736 | Float | RD | W | The fundamental real power L1
Current measurement module group 05 | 22738 | Float | RD | W | The fundamental real power L2
Current measurement module group 05 | 22740 | Float | RD | W | The fundamental real power L3
Current measurement module group 05 | 22744 | Float | RD | - | The power factor L1
Current measurement module group 05 | 22746 | Float | RD | - | The power factor L2
Current measurement module group 05 | 22748 | Float | RD | - | The power factor L3
Current measurement module group 05 | 22752 | Float | RD | var | The reactive power L1
Current measurement module group 05 | 22754 | Float | RD | var | The reactive power L2
Current measurement module group 05 | 22756 | Float | RD | var | The reactive power L3
Current measurement module group 05 | 22760 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 05 | 22762 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 05 | 22764 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 05 | 22766 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 05 | 22768 | Float | RD | A | The real part of current IL1
Current measurement module group 05 | 22770 | Float | RD | A | The real part of current IL2
Current measurement module group 05 | 22772 | Float | RD | A | The real part of current IL3
Current measurement module group 05 | 22774 | Float | RD | A | The real part of current IL4
Current measurement module group 06 | 22776 | Float | RD | var | harmonic distortion power L1
Current measurement module group 06 | 22778 | Float | RD | var | harmonic distortion power L2
Current measurement module group 06 | 22780 | Float | RD | var | harmonic distortion power L3
Current measurement module group 06 | 22784 | Float | RD | - | crest factor L1
Current measurement module group 06 | 22786 | Float | RD | - | crest factor L2
Current measurement module group 06 | 22788 | Float | RD | - | crest factor L3
Current measurement module group 06 | 22790 | Float | RD | - | crest factor L4
Current measurement module group 06 | 22792 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 06 | 22794 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 06 | 22796 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 06 | 22798 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 06 | 22800 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 06 | 22802 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 06 | 22804 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 06 | 22806 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 06 | 22808 | Float | RD | W | The fundamental real power L1
Current measurement module group 06 | 22810 | Float | RD | W | The fundamental real power L2
Current measurement module group 06 | 22812 | Float | RD | W | The fundamental real power L3
Current measurement module group 06 | 22816 | Float | RD | - | The power factor L1
Current measurement module group 06 | 22818 | Float | RD | - | The power factor L2
Current measurement module group 06 | 22820 | Float | RD | - | The power factor L3
Current measurement module group 06 | 22824 | Float | RD | var | The reactive power L1
Current measurement module group 06 | 22826 | Float | RD | var | The reactive power L2
Current measurement module group 06 | 22828 | Float | RD | var | The reactive power L3
Current measurement module group 06 | 22832 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 06 | 22834 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 06 | 22836 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 06 | 22838 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 06 | 22840 | Float | RD | A | The real part of current IL1
Current measurement module group 06 | 22842 | Float | RD | A | The real part of current IL2
Current measurement module group 06 | 22844 | Float | RD | A | The real part of current IL3
Current measurement module group 06 | 22846 | Float | RD | A | The real part of current IL4
Current measurement module group 07 | 22848 | Float | RD | var | harmonic distortion power L1
Current measurement module group 07 | 22850 | Float | RD | var | harmonic distortion power L2
Current measurement module group 07 | 22852 | Float | RD | var | harmonic distortion power L3
Current measurement module group 07 | 22856 | Float | RD | - | crest factor L1
Current measurement module group 07 | 22858 | Float | RD | - | crest factor L2
Current measurement module group 07 | 22860 | Float | RD | - | crest factor L3
Current measurement module group 07 | 22862 | Float | RD | - | crest factor L4
Current measurement module group 07 | 22864 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 07 | 22866 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 07 | 22868 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 07 | 22870 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 07 | 22872 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 07 | 22874 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 07 | 22876 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 07 | 22878 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 07 | 22880 | Float | RD | W | The fundamental real power L1
Current measurement module group 07 | 22882 | Float | RD | W | The fundamental real power L2
Current measurement module group 07 | 22884 | Float | RD | W | The fundamental real power L3
Current measurement module group 07 | 22888 | Float | RD | - | The power factor L1
Current measurement module group 07 | 22890 | Float | RD | - | The power factor L2
Current measurement module group 07 | 22892 | Float | RD | - | The power factor L3
Current measurement module group 07 | 22896 | Float | RD | var | The reactive power L1
Current measurement module group 07 | 22898 | Float | RD | var | The reactive power L2
Current measurement module group 07 | 22900 | Float | RD | var | The reactive power L3
Current measurement module group 07 | 22904 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 07 | 22906 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 07 | 22908 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 07 | 22910 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 07 | 22912 | Float | RD | A | The real part of current IL1
Current measurement module group 07 | 22914 | Float | RD | A | The real part of current IL2
Current measurement module group 07 | 22916 | Float | RD | A | The real part of current IL3
Current measurement module group 07 | 22918 | Float | RD | A | The real part of current IL4
Current measurement module group 08 | 22920 | Float | RD | var | harmonic distortion power L1
Current measurement module group 08 | 22922 | Float | RD | var | harmonic distortion power L2
Current measurement module group 08 | 22924 | Float | RD | var | harmonic distortion power L3
Current measurement module group 08 | 22928 | Float | RD | - | crest factor L1
Current measurement module group 08 | 22930 | Float | RD | - | crest factor L2
Current measurement module group 08 | 22932 | Float | RD | - | crest factor L3
Current measurement module group 08 | 22934 | Float | RD | - | crest factor L4
Current measurement module group 08 | 22936 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 08 | 22938 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 08 | 22940 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 08 | 22942 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 08 | 22944 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 08 | 22946 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 08 | 22948 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 08 | 22950 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 08 | 22952 | Float | RD | W | The fundamental real power L1
Current measurement module group 08 | 22954 | Float | RD | W | The fundamental real power L2
Current measurement module group 08 | 22956 | Float | RD | W | The fundamental real power L3
Current measurement module group 08 | 22960 | Float | RD | - | The power factor L1
Current measurement module group 08 | 22962 | Float | RD | - | The power factor L2
Current measurement module group 08 | 22964 | Float | RD | - | The power factor L3
Current measurement module group 08 | 22968 | Float | RD | var | The reactive power L1
Current measurement module group 08 | 22970 | Float | RD | var | The reactive power L2
Current measurement module group 08 | 22972 | Float | RD | var | The reactive power L3
Current measurement module group 08 | 22976 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 08 | 22978 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 08 | 22980 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 08 | 22982 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 08 | 22984 | Float | RD | A | The real part of current IL1
Current measurement module group 08 | 22986 | Float | RD | A | The real part of current IL2
Current measurement module group 08 | 22988 | Float | RD | A | The real part of current IL3
Current measurement module group 08 | 22990 | Float | RD | A | The real part of current IL4
Current measurement module group 09 | 22992 | Float | RD | var | harmonic distortion power L1
Current measurement module group 09 | 22994 | Float | RD | var | harmonic distortion power L2
Current measurement module group 09 | 22996 | Float | RD | var | harmonic distortion power L3
Current measurement module group 09 | 23000 | Float | RD | - | crest factor L1
Current measurement module group 09 | 23002 | Float | RD | - | crest factor L2
Current measurement module group 09 | 23004 | Float | RD | - | crest factor L3
Current measurement module group 09 | 23006 | Float | RD | - | crest factor L4
Current measurement module group 09 | 23008 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 09 | 23010 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 09 | 23012 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 09 | 23014 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 09 | 23016 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 09 | 23018 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 09 | 23020 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 09 | 23022 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 09 | 23024 | Float | RD | W | The fundamental real power L1
Current measurement module group 09 | 23026 | Float | RD | W | The fundamental real power L2
Current measurement module group 09 | 23028 | Float | RD | W | The fundamental real power L3
Current measurement module group 09 | 23032 | Float | RD | - | The power factor L1
Current measurement module group 09 | 23034 | Float | RD | - | The power factor L2
Current measurement module group 09 | 23036 | Float | RD | - | The power factor L3
Current measurement module group 09 | 23040 | Float | RD | var | The reactive power L1
Current measurement module group 09 | 23042 | Float | RD | var | The reactive power L2
Current measurement module group 09 | 23044 | Float | RD | var | The reactive power L3
Current measurement module group 09 | 23048 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 09 | 23050 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 09 | 23052 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 09 | 23054 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 09 | 23056 | Float | RD | A | The real part of current IL1
Current measurement module group 09 | 23058 | Float | RD | A | The real part of current IL2
Current measurement module group 09 | 23060 | Float | RD | A | The real part of current IL3
Current measurement module group 09 | 23062 | Float | RD | A | The real part of current IL4
Current measurement module group 10 | 23064 | Float | RD | var | harmonic distortion power L1
Current measurement module group 10 | 23066 | Float | RD | var | harmonic distortion power L2
Current measurement module group 10 | 23068 | Float | RD | var | harmonic distortion power L3
Current measurement module group 10 | 23072 | Float | RD | - | crest factor L1
Current measurement module group 10 | 23074 | Float | RD | - | crest factor L2
Current measurement module group 10 | 23076 | Float | RD | - | crest factor L3
Current measurement module group 10 | 23078 | Float | RD | - | crest factor L4
Current measurement module group 10 | 23080 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 10 | 23082 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 10 | 23084 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 10 | 23086 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 10 | 23088 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 10 | 23090 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 10 | 23092 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 10 | 23094 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 10 | 23096 | Float | RD | W | The fundamental real power L1
Current measurement module group 10 | 23098 | Float | RD | W | The fundamental real power L2
Current measurement module group 10 | 23100 | Float | RD | W | The fundamental real power L3
Current measurement module group 10 | 23104 | Float | RD | - | The power factor L1
Current measurement module group 10 | 23106 | Float | RD | - | The power factor L2
Current measurement module group 10 | 23108 | Float | RD | - | The power factor L3
Current measurement module group 10 | 23112 | Float | RD | var | The reactive power L1
Current measurement module group 10 | 23114 | Float | RD | var | The reactive power L2
Current measurement module group 10 | 23116 | Float | RD | var | The reactive power L3
Current measurement module group 10 | 23120 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 10 | 23122 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 10 | 23124 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 10 | 23126 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 10 | 23128 | Float | RD | A | The real part of current IL1
Current measurement module group 10 | 23130 | Float | RD | A | The real part of current IL2
Current measurement module group 10 | 23132 | Float | RD | A | The real part of current IL3
Current measurement module group 10 | 23134 | Float | RD | A | The real part of current IL4
Current measurement module group 11 | 23136 | Float | RD | var | harmonic distortion power L1
Current measurement module group 11 | 23138 | Float | RD | var | harmonic distortion power L2
Current measurement module group 11 | 23140 | Float | RD | var | harmonic distortion power L3
Current measurement module group 11 | 23144 | Float | RD | - | crest factor L1
Current measurement module group 11 | 23146 | Float | RD | - | crest factor L2
Current measurement module group 11 | 23148 | Float | RD | - | crest factor L3
Current measurement module group 11 | 23150 | Float | RD | - | crest factor L4
Current measurement module group 11 | 23152 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 11 | 23154 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 11 | 23156 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 11 | 23158 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 11 | 23160 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 11 | 23162 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 11 | 23164 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 11 | 23166 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 11 | 23168 | Float | RD | W | The fundamental real power L1
Current measurement module group 11 | 23170 | Float | RD | W | The fundamental real power L2
Current measurement module group 11 | 23172 | Float | RD | W | The fundamental real power L3
Current measurement module group 11 | 23176 | Float | RD | - | The power factor L1
Current measurement module group 11 | 23178 | Float | RD | - | The power factor L2
Current measurement module group 11 | 23180 | Float | RD | - | The power factor L3
Current measurement module group 11 | 23184 | Float | RD | var | The reactive power L1
Current measurement module group 11 | 23186 | Float | RD | var | The reactive power L2
Current measurement module group 11 | 23188 | Float | RD | var | The reactive power L3
Current measurement module group 11 | 23192 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 11 | 23194 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 11 | 23196 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 11 | 23198 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 11 | 23200 | Float | RD | A | The real part of current IL1
Current measurement module group 11 | 23202 | Float | RD | A | The real part of current IL2
Current measurement module group 11 | 23204 | Float | RD | A | The real part of current IL3
Current measurement module group 11 | 23206 | Float | RD | A | The real part of current IL4
Current measurement module group 12 | 23208 | Float | RD | var | harmonic distortion power L1
Current measurement module group 12 | 23210 | Float | RD | var | harmonic distortion power L2
Current measurement module group 12 | 23212 | Float | RD | var | harmonic distortion power L3
Current measurement module group 12 | 23216 | Float | RD | - | crest factor L1
Current measurement module group 12 | 23218 | Float | RD | - | crest factor L2
Current measurement module group 12 | 23220 | Float | RD | - | crest factor L3
Current measurement module group 12 | 23222 | Float | RD | - | crest factor L4
Current measurement module group 12 | 23224 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 12 | 23226 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 12 | 23228 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 12 | 23230 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 12 | 23232 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 12 | 23234 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 12 | 23236 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 12 | 23238 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 12 | 23240 | Float | RD | W | The fundamental real power L1
Current measurement module group 12 | 23242 | Float | RD | W | The fundamental real power L2
Current measurement module group 12 | 23244 | Float | RD | W | The fundamental real power L3
Current measurement module group 12 | 23248 | Float | RD | - | The power factor L1
Current measurement module group 12 | 23250 | Float | RD | - | The power factor L2
Current measurement module group 12 | 23252 | Float | RD | - | The power factor L3
Current measurement module group 12 | 23256 | Float | RD | var | The reactive power L1
Current measurement module group 12 | 23258 | Float | RD | var | The reactive power L2
Current measurement module group 12 | 23260 | Float | RD | var | The reactive power L3
Current measurement module group 12 | 23264 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 12 | 23266 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 12 | 23268 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 12 | 23270 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 12 | 23272 | Float | RD | A | The real part of current IL1
Current measurement module group 12 | 23274 | Float | RD | A | The real part of current IL2
Current measurement module group 12 | 23276 | Float | RD | A | The real part of current IL3
Current measurement module group 12 | 23278 | Float | RD | A | The real part of current IL4
Current measurement module group 13 | 23280 | Float | RD | var | harmonic distortion power L1
Current measurement module group 13 | 23282 | Float | RD | var | harmonic distortion power L2
Current measurement module group 13 | 23284 | Float | RD | var | harmonic distortion power L3
Current measurement module group 13 | 23288 | Float | RD | - | crest factor L1
Current measurement module group 13 | 23290 | Float | RD | - | crest factor L2
Current measurement module group 13 | 23292 | Float | RD | - | crest factor L3
Current measurement module group 13 | 23294 | Float | RD | - | crest factor L4
Current measurement module group 13 | 23296 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 13 | 23298 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 13 | 23300 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 13 | 23302 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 13 | 23304 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 13 | 23306 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 13 | 23308 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 13 | 23310 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 13 | 23312 | Float | RD | W | The fundamental real power L1
Current measurement module group 13 | 23314 | Float | RD | W | The fundamental real power L2
Current measurement module group 13 | 23316 | Float | RD | W | The fundamental real power L3
Current measurement module group 13 | 23320 | Float | RD | - | The power factor L1
Current measurement module group 13 | 23322 | Float | RD | - | The power factor L2
Current measurement module group 13 | 23324 | Float | RD | - | The power factor L3
Current measurement module group 13 | 23328 | Float | RD | var | The reactive power L1
Current measurement module group 13 | 23330 | Float | RD | var | The reactive power L2
Current measurement module group 13 | 23332 | Float | RD | var | The reactive power L3
Current measurement module group 13 | 23336 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 13 | 23338 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 13 | 23340 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 13 | 23342 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 13 | 23344 | Float | RD | A | The real part of current IL1
Current measurement module group 13 | 23346 | Float | RD | A | The real part of current IL2
Current measurement module group 13 | 23348 | Float | RD | A | The real part of current IL3
Current measurement module group 13 | 23350 | Float | RD | A | The real part of current IL4
Current measurement module group 14 | 23352 | Float | RD | var | harmonic distortion power L1
Current measurement module group 14 | 23354 | Float | RD | var | harmonic distortion power L2
Current measurement module group 14 | 23356 | Float | RD | var | harmonic distortion power L3
Current measurement module group 14 | 23360 | Float | RD | - | crest factor L1
Current measurement module group 14 | 23362 | Float | RD | - | crest factor L2
Current measurement module group 14 | 23364 | Float | RD | - | crest factor L3
Current measurement module group 14 | 23366 | Float | RD | - | crest factor L4
Current measurement module group 14 | 23368 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 14 | 23370 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 14 | 23372 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 14 | 23374 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 14 | 23376 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 14 | 23378 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 14 | 23380 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 14 | 23382 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 14 | 23384 | Float | RD | W | The fundamental real power L1
Current measurement module group 14 | 23386 | Float | RD | W | The fundamental real power L2
Current measurement module group 14 | 23388 | Float | RD | W | The fundamental real power L3
Current measurement module group 14 | 23392 | Float | RD | - | The power factor L1
Current measurement module group 14 | 23394 | Float | RD | - | The power factor L2
Current measurement module group 14 | 23396 | Float | RD | - | The power factor L3
Current measurement module group 14 | 23400 | Float | RD | var | The reactive power L1
Current measurement module group 14 | 23402 | Float | RD | var | The reactive power L2
Current measurement module group 14 | 23404 | Float | RD | var | The reactive power L3
Current measurement module group 14 | 23408 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 14 | 23410 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 14 | 23412 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 14 | 23414 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 14 | 23416 | Float | RD | A | The real part of current IL1
Current measurement module group 14 | 23418 | Float | RD | A | The real part of current IL2
Current measurement module group 14 | 23420 | Float | RD | A | The real part of current IL3
Current measurement module group 14 | 23422 | Float | RD | A | The real part of current IL4
Current measurement module group 15 | 23424 | Float | RD | var | harmonic distortion power L1
Current measurement module group 15 | 23426 | Float | RD | var | harmonic distortion power L2
Current measurement module group 15 | 23428 | Float | RD | var | harmonic distortion power L3
Current measurement module group 15 | 23432 | Float | RD | - | crest factor L1
Current measurement module group 15 | 23434 | Float | RD | - | crest factor L2
Current measurement module group 15 | 23436 | Float | RD | - | crest factor L3
Current measurement module group 15 | 23438 | Float | RD | - | crest factor L4
Current measurement module group 15 | 23440 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 15 | 23442 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 15 | 23444 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 15 | 23446 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 15 | 23448 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 15 | 23450 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 15 | 23452 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 15 | 23454 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 15 | 23456 | Float | RD | W | The fundamental real power L1
Current measurement module group 15 | 23458 | Float | RD | W | The fundamental real power L2
Current measurement module group 15 | 23460 | Float | RD | W | The fundamental real power L3
Current measurement module group 15 | 23464 | Float | RD | - | The power factor L1
Current measurement module group 15 | 23466 | Float | RD | - | The power factor L2
Current measurement module group 15 | 23468 | Float | RD | - | The power factor L3
Current measurement module group 15 | 23472 | Float | RD | var | The reactive power L1
Current measurement module group 15 | 23474 | Float | RD | var | The reactive power L2
Current measurement module group 15 | 23476 | Float | RD | var | The reactive power L3
Current measurement module group 15 | 23480 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 15 | 23482 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 15 | 23484 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 15 | 23486 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 15 | 23488 | Float | RD | A | The real part of current IL1
Current measurement module group 15 | 23490 | Float | RD | A | The real part of current IL2
Current measurement module group 15 | 23492 | Float | RD | A | The real part of current IL3
Current measurement module group 15 | 23494 | Float | RD | A | The real part of current IL4
Current measurement module group 16 | 23496 | Float | RD | var | harmonic distortion power L1
Current measurement module group 16 | 23498 | Float | RD | var | harmonic distortion power L2
Current measurement module group 16 | 23500 | Float | RD | var | harmonic distortion power L3
Current measurement module group 16 | 23504 | Float | RD | - | crest factor L1
Current measurement module group 16 | 23506 | Float | RD | - | crest factor L2
Current measurement module group 16 | 23508 | Float | RD | - | crest factor L3
Current measurement module group 16 | 23510 | Float | RD | - | crest factor L4
Current measurement module group 16 | 23512 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 16 | 23514 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 16 | 23516 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 16 | 23518 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 16 | 23520 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 16 | 23522 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 16 | 23524 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 16 | 23526 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 16 | 23528 | Float | RD | W | The fundamental real power L1
Current measurement module group 16 | 23530 | Float | RD | W | The fundamental real power L2
Current measurement module group 16 | 23532 | Float | RD | W | The fundamental real power L3
Current measurement module group 16 | 23536 | Float | RD | - | The power factor L1
Current measurement module group 16 | 23538 | Float | RD | - | The power factor L2
Current measurement module group 16 | 23540 | Float | RD | - | The power factor L3
Current measurement module group 16 | 23544 | Float | RD | var | The reactive power L1
Current measurement module group 16 | 23546 | Float | RD | var | The reactive power L2
Current measurement module group 16 | 23548 | Float | RD | var | The reactive power L3
Current measurement module group 16 | 23552 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 16 | 23554 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 16 | 23556 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 16 | 23558 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 16 | 23560 | Float | RD | A | The real part of current IL1
Current measurement module group 16 | 23562 | Float | RD | A | The real part of current IL2
Current measurement module group 16 | 23564 | Float | RD | A | The real part of current IL3
Current measurement module group 16 | 23566 | Float | RD | A | The real part of current IL4
Current measurement module group 17 | 23568 | Float | RD | var | harmonic distortion power L1
Current measurement module group 17 | 23570 | Float | RD | var | harmonic distortion power L2
Current measurement module group 17 | 23572 | Float | RD | var | harmonic distortion power L3
Current measurement module group 17 | 23576 | Float | RD | - | crest factor L1
Current measurement module group 17 | 23578 | Float | RD | - | crest factor L2
Current measurement module group 17 | 23580 | Float | RD | - | crest factor L3
Current measurement module group 17 | 23582 | Float | RD | - | crest factor L4
Current measurement module group 17 | 23584 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 17 | 23586 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 17 | 23588 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 17 | 23590 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 17 | 23592 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 17 | 23594 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 17 | 23596 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 17 | 23598 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 17 | 23600 | Float | RD | W | The fundamental real power L1
Current measurement module group 17 | 23602 | Float | RD | W | The fundamental real power L2
Current measurement module group 17 | 23604 | Float | RD | W | The fundamental real power L3
Current measurement module group 17 | 23608 | Float | RD | - | The power factor L1
Current measurement module group 17 | 23610 | Float | RD | - | The power factor L2
Current measurement module group 17 | 23612 | Float | RD | - | The power factor L3
Current measurement module group 17 | 23616 | Float | RD | var | The reactive power L1
Current measurement module group 17 | 23618 | Float | RD | var | The reactive power L2
Current measurement module group 17 | 23620 | Float | RD | var | The reactive power L3
Current measurement module group 17 | 23624 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 17 | 23626 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 17 | 23628 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 17 | 23630 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 17 | 23632 | Float | RD | A | The real part of current IL1
Current measurement module group 17 | 23634 | Float | RD | A | The real part of current IL2
Current measurement module group 17 | 23636 | Float | RD | A | The real part of current IL3
Current measurement module group 17 | 23638 | Float | RD | A | The real part of current IL4
Current measurement module group 18 | 23640 | Float | RD | var | harmonic distortion power L1
Current measurement module group 18 | 23642 | Float | RD | var | harmonic distortion power L2
Current measurement module group 18 | 23644 | Float | RD | var | harmonic distortion power L3
Current measurement module group 18 | 23648 | Float | RD | - | crest factor L1
Current measurement module group 18 | 23650 | Float | RD | - | crest factor L2
Current measurement module group 18 | 23652 | Float | RD | - | crest factor L3
Current measurement module group 18 | 23654 | Float | RD | - | crest factor L4
Current measurement module group 18 | 23656 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 18 | 23658 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 18 | 23660 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 18 | 23662 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 18 | 23664 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 18 | 23666 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 18 | 23668 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 18 | 23670 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 18 | 23672 | Float | RD | W | The fundamental real power L1
Current measurement module group 18 | 23674 | Float | RD | W | The fundamental real power L2
Current measurement module group 18 | 23676 | Float | RD | W | The fundamental real power L3
Current measurement module group 18 | 23680 | Float | RD | - | The power factor L1
Current measurement module group 18 | 23682 | Float | RD | - | The power factor L2
Current measurement module group 18 | 23684 | Float | RD | - | The power factor L3
Current measurement module group 18 | 23688 | Float | RD | var | The reactive power L1
Current measurement module group 18 | 23690 | Float | RD | var | The reactive power L2
Current measurement module group 18 | 23692 | Float | RD | var | The reactive power L3
Current measurement module group 18 | 23696 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 18 | 23698 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 18 | 23700 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 18 | 23702 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 18 | 23704 | Float | RD | A | The real part of current IL1
Current measurement module group 18 | 23706 | Float | RD | A | The real part of current IL2
Current measurement module group 18 | 23708 | Float | RD | A | The real part of current IL3
Current measurement module group 18 | 23710 | Float | RD | A | The real part of current IL4
Current measurement module group 19 | 23712 | Float | RD | var | harmonic distortion power L1
Current measurement module group 19 | 23714 | Float | RD | var | harmonic distortion power L2
Current measurement module group 19 | 23716 | Float | RD | var | harmonic distortion power L3
Current measurement module group 19 | 23720 | Float | RD | - | crest factor L1
Current measurement module group 19 | 23722 | Float | RD | - | crest factor L2
Current measurement module group 19 | 23724 | Float | RD | - | crest factor L3
Current measurement module group 19 | 23726 | Float | RD | - | crest factor L4
Current measurement module group 19 | 23728 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 19 | 23730 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 19 | 23732 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 19 | 23734 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 19 | 23736 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 19 | 23738 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 19 | 23740 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 19 | 23742 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 19 | 23744 | Float | RD | W | The fundamental real power L1
Current measurement module group 19 | 23746 | Float | RD | W | The fundamental real power L2
Current measurement module group 19 | 23748 | Float | RD | W | The fundamental real power L3
Current measurement module group 19 | 23752 | Float | RD | - | The power factor L1
Current measurement module group 19 | 23754 | Float | RD | - | The power factor L2
Current measurement module group 19 | 23756 | Float | RD | - | The power factor L3
Current measurement module group 19 | 23760 | Float | RD | var | The reactive power L1
Current measurement module group 19 | 23762 | Float | RD | var | The reactive power L2
Current measurement module group 19 | 23764 | Float | RD | var | The reactive power L3
Current measurement module group 19 | 23768 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 19 | 23770 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 19 | 23772 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 19 | 23774 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 19 | 23776 | Float | RD | A | The real part of current IL1
Current measurement module group 19 | 23778 | Float | RD | A | The real part of current IL2
Current measurement module group 19 | 23780 | Float | RD | A | The real part of current IL3
Current measurement module group 19 | 23782 | Float | RD | A | The real part of current IL4
Current measurement module group 20 | 23784 | Float | RD | var | harmonic distortion power L1
Current measurement module group 20 | 23786 | Float | RD | var | harmonic distortion power L2
Current measurement module group 20 | 23788 | Float | RD | var | harmonic distortion power L3
Current measurement module group 20 | 23792 | Float | RD | - | crest factor L1
Current measurement module group 20 | 23794 | Float | RD | - | crest factor L2
Current measurement module group 20 | 23796 | Float | RD | - | crest factor L3
Current measurement module group 20 | 23798 | Float | RD | - | crest factor L4
Current measurement module group 20 | 23800 | Float | RD | % | load current percentages of the rated current I1
Current measurement module group 20 | 23802 | Float | RD | % | load current percentages of the rated current I2
Current measurement module group 20 | 23804 | Float | RD | % | load current percentages of the rated current I3
Current measurement module group 20 | 23806 | Float | RD | % | load current percentages of the rated current I4
Current measurement module group 20 | 23808 | Float | RD | % | The total demand distortion TDD IL1
Current measurement module group 20 | 23810 | Float | RD | % | The total demand distortion TDD IL2
Current measurement module group 20 | 23812 | Float | RD | % | The total demand distortion TDD IL3
Current measurement module group 20 | 23814 | Float | RD | % | The total demand distortion TDD IL4
Current measurement module group 20 | 23816 | Float | RD | W | The fundamental real power L1
Current measurement module group 20 | 23818 | Float | RD | W | The fundamental real power L2
Current measurement module group 20 | 23820 | Float | RD | W | The fundamental real power L3
Current measurement module group 20 | 23824 | Float | RD | - | The power factor L1
Current measurement module group 20 | 23826 | Float | RD | - | The power factor L2
Current measurement module group 20 | 23828 | Float | RD | - | The power factor L3
Current measurement module group 20 | 23832 | Float | RD | var | The reactive power L1
Current measurement module group 20 | 23834 | Float | RD | var | The reactive power L2
Current measurement module group 20 | 23836 | Float | RD | var | The reactive power L3
Current measurement module group 20 | 23840 | Float | RD | A | The imaginary part of current IL1
Current measurement module group 20 | 23842 | Float | RD | A | The imaginary part of current IL2
Current measurement module group 20 | 23844 | Float | RD | A | The imaginary part of current IL3
Current measurement module group 20 | 23846 | Float | RD | A | The imaginary part of current IL4
Current measurement module group 20 | 23848 | Float | RD | A | The real part of current IL1
Current measurement module group 20 | 23850 | Float | RD | A | The real part of current IL2
Current measurement module group 20 | 23852 | Float | RD | A | The real part of current IL3
Current measurement module group 20 | 23854 | Float | RD | A | The real part of current IL4
basic device | 24200 | Float | RD | - | Voltage transformer, primary
basic device | 24202 | Float | RD | - | Voltage transformer, secondary
Basic device measuring group 01 | 24204 | Float | RD | - | Current transformer, primary - L1
Basic device measuring group 01 | 24206 | Float | RD | - | Current transformer, secondary - L1
Basic device measuring group 01 | 24208 | Float | RD | - | Current transformer, primary - L2
Basic device measuring group 01 | 24210 | Float | RD | - | Current transformer, secondary - L2
Basic device measuring group 01 | 24212 | Float | RD | - | Current transformer, primary - L3
Basic device measuring group 01 | 24214 | Float | RD | - | Current transformer, secondary - L3
Basic device measuring group 01 | 24216 | Float | RD | - | Current transformer, primary - L4
Basic device measuring group 01 | 24218 | Float | RD | - | Current transformer, secondary - L4
Basic device measuring group 02 | 24220 | Float | RD | - | Current transformer, primary - L1
Basic device measuring group 02 | 24222 | Float | RD | - | Current transformer, secondary - L1
Basic device measuring group 02 | 24224 | Float | RD | - | Current transformer, primary - L2
Basic device measuring group 02 | 24226 | Float | RD | - | Current transformer, secondary - L2
Basic device measuring group 02 | 24228 | Float | RD | - | Current transformer, primary - L3
Basic device measuring group 02 | 24230 | Float | RD | - | Current transformer, secondary - L3
Basic device measuring group 02 | 24232 | Float | RD | - | Current transformer, primary - L4
Basic device measuring group 02 | 24234 | Float | RD | - | Current transformer, secondary - L4
Basic device measuring group 03 | 24236 | Float | RD | - | Current transformer, primary - L1
Basic device measuring group 03 | 24238 | Float | RD | - | Current transformer, secondary - L1
Basic device measuring group 03 | 24240 | Float | RD | - | Current transformer, primary - L2
Basic device measuring group 03 | 24242 | Float | RD | - | Current transformer, secondary - L2
Basic device measuring group 03 | 24244 | Float | RD | - | Current transformer, primary - L3
Basic device measuring group 03 | 24246 | Float | RD | - | Current transformer, secondary - L3
Basic device measuring group 03 | 24248 | Float | RD | - | Current transformer, primary - L4
Basic device measuring group 03 | 24250 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 01 | 24252 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 01 | 24254 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 01 | 24256 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 01 | 24258 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 01 | 24260 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 01 | 24262 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 01 | 24264 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 01 | 24266 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 02 | 24268 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 02 | 24270 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 02 | 24272 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 02 | 24274 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 02 | 24276 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 02 | 24278 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 02 | 24280 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 02 | 24282 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 03 | 24284 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 03 | 24286 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 03 | 24288 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 03 | 24290 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 03 | 24292 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 03 | 24294 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 03 | 24296 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 03 | 24298 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 04 | 24300 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 04 | 24302 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 04 | 24304 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 04 | 24306 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 04 | 24308 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 04 | 24310 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 04 | 24312 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 04 | 24314 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 05 | 24316 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 05 | 24318 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 05 | 24320 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 05 | 24322 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 05 | 24324 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 05 | 24326 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 05 | 24328 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 05 | 24330 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 06 | 24332 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 06 | 24334 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 06 | 24336 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 06 | 24338 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 06 | 24340 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 06 | 24342 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 06 | 24344 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 06 | 24346 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 07 | 24348 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 07 | 24350 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 07 | 24352 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 07 | 24354 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 07 | 24356 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 07 | 24358 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 07 | 24360 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 07 | 24362 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 08 | 24364 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 08 | 24366 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 08 | 24368 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 08 | 24370 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 08 | 24372 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 08 | 24374 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 08 | 24376 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 08 | 24378 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 09 | 24380 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 09 | 24382 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 09 | 24384 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 09 | 24386 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 09 | 24388 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 09 | 24390 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 09 | 24392 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 09 | 24394 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 10 | 24396 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 10 | 24398 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 10 | 24400 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 10 | 24402 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 10 | 24404 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 10 | 24406 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 10 | 24408 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 10 | 24410 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 11 | 24412 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 11 | 24414 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 11 | 24416 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 11 | 24418 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 11 | 24420 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 11 | 24422 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 11 | 24424 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 11 | 24426 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 12 | 24428 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 12 | 24430 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 12 | 24432 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 12 | 24434 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 12 | 24436 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 12 | 24438 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 12 | 24440 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 12 | 24442 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 13 | 24444 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 13 | 24446 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 13 | 24448 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 13 | 24450 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 13 | 24452 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 13 | 24454 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 13 | 24456 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 13 | 24458 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 14 | 24460 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 14 | 24462 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 14 | 24464 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 14 | 24466 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 14 | 24468 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 14 | 24470 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 14 | 24472 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 14 | 24474 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 15 | 24476 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 15 | 24478 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 15 | 24480 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 15 | 24482 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 15 | 24484 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 15 | 24486 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 15 | 24488 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 15 | 24490 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 16 | 24492 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 16 | 24494 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 16 | 24496 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 16 | 24498 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 16 | 24500 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 16 | 24502 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 16 | 24504 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 16 | 24506 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 17 | 24508 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 17 | 24510 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 17 | 24512 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 17 | 24514 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 17 | 24516 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 17 | 24518 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 17 | 24520 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 17 | 24522 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 18 | 24524 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 18 | 24526 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 18 | 24528 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 18 | 24530 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 18 | 24532 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 18 | 24534 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 18 | 24536 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 18 | 24538 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 19 | 24540 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 19 | 24542 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 19 | 24544 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 19 | 24546 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 19 | 24548 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 19 | 24550 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 19 | 24552 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 19 | 24554 | Float | RD | - | Current transformer, secondary - L4
Current measurement module group 20 | 24556 | Float | RD | - | Current transformer, primary - L1
Current measurement module group 20 | 24558 | Float | RD | - | Current transformer, secondary - L1
Current measurement module group 20 | 24560 | Float | RD | - | Current transformer, primary - L2
Current measurement module group 20 | 24562 | Float | RD | - | Current transformer, secondary - L2
Current measurement module group 20 | 24564 | Float | RD | - | Current transformer, primary - L3
Current measurement module group 20 | 24566 | Float | RD | - | Current transformer, secondary - L3
Current measurement module group 20 | 24568 | Float | RD | - | Current transformer, primary - L4
Current measurement module group 20 | 24570 | Float | RD | - | Current transformer, secondary - L4
Basic device measuring group 01 | 24650 | Float | RD | A | average of the current of the measurement group 1
Basic device measuring group 02 | 24652 | Float | RD | A | average of the current of the measurement group 2
Basic device measuring group 03 | 24654 | Float | RD | A | average of the current of the measurement group 3
Basic device measuring group 01 | 24656 | Float | RD | % | average of the load percentages of rated current of the measurement group 1
Basic device measuring group 02 | 24658 | Float | RD | % | average of the load percentages of rated current of the measurement group 2
Basic device measuring group 03 | 24660 | Float | RD | % | average of the load percentages of rated current of the measurement group 3
Basic device measuring group 01 | 24662 | Float | RD | A | The calculated neutral line current of the measurement group 1
Basic device measuring group 02 | 24664 | Float | RD | A | The calculated neutral line current of the measurement group 2
Basic device measuring group 03 | 24666 | Float | RD | A | The calculated neutral line current of the measurement group 3
Basic device measuring group 01 | 24668 | Float | RD | A | The RMS Current sum of the measurement group 1
Basic device measuring group 02 | 24670 | Float | RD | A | The RMS Current sum of the measurement group 2
Basic device measuring group 03 | 24672 | Float | RD | A | The RMS Current sum of the measurement group 3
Basic device measuring group 01 | 24674 | Float | RD | % | The average total demand distortion current of the measurement group 1
Basic device measuring group 02 | 24676 | Float | RD | % | The average total demand distortion current of the measurement group 2
Basic device measuring group 03 | 24678 | Float | RD | % | The average total demand distortion current of the measurement group 3
Basic device measuring group 01 | 24680 | Float | RD | % | The average total harmonic distortion current of the measurement group 1
Basic device measuring group 02 | 24682 | Float | RD | % | The average total harmonic distortion current of the measurement group 2
Basic device measuring group 03 | 24684 | Float | RD | % | The average total harmonic distortion current of the measurement group 3
Basic device measuring group 01 | 27000 | Float | RD | A | I1(L1) - max. average current value - highest value drag indicator
Basic device measuring group 01 | 27002 | Float | RD | A | I2(L2) - max. average current value - highest value drag indicator
Basic device measuring group 01 | 27004 | Float | RD | A | I3(L3) - max. average current value - highest value drag indicator
Basic device measuring group 01 | 27006 | Float | RD | A | I4(L4) - max. average current value - highest value drag indicator
Basic device measuring group 01 | 27008 | DateTime | RD | - | I1(L1) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 01 | 27010 | DateTime | RD | - | I2(L2) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 01 | 27012 | DateTime | RD | - | I3(L3) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 01 | 27014 | DateTime | RD | - | I4(L4) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 01 | 27016 | Float | RD | VA | I1(L1)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27018 | Float | RD | VA | I2(L2)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27020 | Float | RD | VA | I3(L3)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27022 | Float | RD | VA | I4(L4)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27024 | DateTime | RD | - | I1(L1)/S - date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27026 | DateTime | RD | - | I2(L2)/S - date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27028 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27030 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 01 | 27032 | Float | RD | W | I1(L1)/P - max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27034 | Float | RD | W | I2(L2)/P - max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27036 | Float | RD | W | I3(L3)/P - max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27038 | Float | RD | W | I4(L4)/P - max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27040 | DateTime | RD | - | I1(L1)/P - date/time, max. average active power value - highest value drag indicator
Basic device measuring group 01 | 27042 | DateTime | RD | - | I2(L2)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27044 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27046 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 01 | 27048 | Float | RD | W | I1(L1)/P- min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27050 | Float | RD | W | I2(L2)/P- min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27052 | Float | RD | W | I3(L3)/P- min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27054 | Float | RD | W | I4(L4)/P- min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27056 | DateTime | RD | - | I1(L1)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27058 | DateTime | RD | - | I2(L2)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27060 | DateTime | RD | - | I3(L3)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27062 | DateTime | RD | - | I4(L4)/P- date/time, min. active power average value - highest value drag indicator
Basic device measuring group 01 | 27064 | Float | RD | A | I1(L1) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27066 | Float | RD | A | I2(L2) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27068 | Float | RD | A | I3(L3) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27070 | Float | RD | A | I4(L4) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27072 | DateTime | RD | - | I1(L1) - date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27074 | DateTime | RD | - | I2(L2) - date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27076 | DateTime | RD | - | I3(L3) - Date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27078 | DateTime | RD | - | I4(L4) - Date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 01 | 27080 | Float | RD | VA | I1(L1)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27082 | Float | RD | VA | I2(L2)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27084 | Float | RD | VA | I3(L3)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27086 | Float | RD | VA | I4(L4)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27088 | DateTime | RD | - | I1(L1)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27090 | DateTime | RD | - | I2(L2)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27092 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27094 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27096 | Float | RD | W | I1(L1)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27098 | Float | RD | W | I2(L2)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27100 | Float | RD | W | I3(L3)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27102 | Float | RD | W | I4(L4)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27104 | DateTime | RD | - | I1(L1)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27106 | DateTime | RD | - | I2(L2)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27108 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27110 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27112 | Float | RD | W | I1(L1)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27114 | Float | RD | W | I2(L2)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27116 | Float | RD | W | I3(L3)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27118 | Float | RD | W | I4(L4)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27120 | DateTime | RD | - | I1(L1)/P- Date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27122 | DateTime | RD | - | I2(L2)/P- Date/time, min. average active power value - 2nd highest value drag indicator
Basic device measuring group 01 | 27124 | DateTime | RD | - | I3(L3)/P- Date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27126 | DateTime | RD | - | I4(L4)/P- date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 01 | 27128 | Float | RD | A | I1(L1) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27130 | Float | RD | A | I2(L2) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27132 | Float | RD | A | I3(L3) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27134 | Float | RD | A | I4(L4) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27136 | DateTime | RD | - | I1(L1) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27138 | DateTime | RD | - | I2(L2) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27140 | DateTime | RD | - | I3(L3) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27142 | DateTime | RD | - | I4(L4) - Date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 01 | 27144 | Float | RD | VA | I1(L1)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27146 | Float | RD | VA | I2(L2)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27148 | Float | RD | VA | I3(L3)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27150 | Float | RD | VA | I4(L4)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27152 | DateTime | RD | - | I1(L1)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27154 | DateTime | RD | - | I2(L2)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27156 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27158 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27160 | Float | RD | W | I1(L1)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27162 | Float | RD | W | I2(L2)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27164 | Float | RD | W | I3(L3)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27166 | Float | RD | W | I4(L4)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27168 | DateTime | RD | - | I1(L1)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27170 | DateTime | RD | - | I2(L2)/P - Date/time, max. average active power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27172 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27174 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27176 | Float | RD | W | I1(L1)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27178 | Float | RD | W | I2(L2)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27180 | Float | RD | W | I3(L3)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27182 | Float | RD | W | I4(L4)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 01 | 27184 | DateTime | RD | - | I1(L1)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27186 | DateTime | RD | - | I2(L2)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27188 | DateTime | RD | - | I3(L3)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27190 | DateTime | RD | - | I4(L4)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 01 | 27192 | Float | RD | A | Sum. I1(L1)-I4(L4) of the max. average current values - highest value drag indicator
Basic device measuring group 01 | 27194 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - highest value drag indicator
Basic device measuring group 01 | 27196 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - highest value drag indicator
Basic device measuring group 01 | 27198 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - highest value drag indicator
Basic device measuring group 01 | 27200 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - highest value drag indicator
Basic device measuring group 01 | 27202 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. active power mean values - highest value drag indicator
Basic device measuring group 01 | 27204 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - highest value drag indicator
Basic device measuring group 01 | 27206 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power mean values - highest value drag indicator
Basic device measuring group 01 | 27208 | Float | RD | A | Sum. I1(L1)-I4(L4) of max. average current values - 2nd highest value drag indicator
Basic device measuring group 01 | 27210 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - 2nd highest value drag indicator
Basic device measuring group 01 | 27212 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - 2nd highest value drag indicator
Basic device measuring group 01 | 27214 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - 2nd highest value drag indicator
Basic device measuring group 01 | 27216 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - 2nd highest value drag indicator
Basic device measuring group 01 | 27218 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. average active power values - 2nd highest value drag indicator
Basic device measuring group 01 | 27220 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - 2nd highest value drag indicator
Basic device measuring group 01 | 27222 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power average values - 2nd highest value drag indicator
Basic device measuring group 01 | 27224 | Float | RD | A | Sum. I1(L1)-I4(L4) of max. average current values - 3rd highest value drag indicator
Basic device measuring group 01 | 27226 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - 3rd highest value drag indicator
Basic device measuring group 01 | 27228 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - 3rd highest value drag indicator
Basic device measuring group 01 | 27230 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - 3rd highest value drag indicator
Basic device measuring group 01 | 27232 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - 3rd highest value drag indicator
Basic device measuring group 01 | 27234 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. active power mean values - 3rd highest value drag indicator
Basic device measuring group 01 | 27236 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - 3rd highest value drag indicator
Basic device measuring group 01 | 27238 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power mean values - 3rd highest value drag indicator
Basic device measuring group 02 | 27240 | Float | RD | A | I1(L1) - max. average current value - highest value drag indicator
Basic device measuring group 02 | 27242 | Float | RD | A | I2(L2) - max. average current value - highest value drag indicator
Basic device measuring group 02 | 27244 | Float | RD | A | I3(L3) - max. average current value - highest value drag indicator
Basic device measuring group 02 | 27246 | Float | RD | A | I4(L4) - max. average current value - highest value drag indicator
Basic device measuring group 02 | 27248 | DateTime | RD | - | I1(L1) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 02 | 27250 | DateTime | RD | - | I2(L2) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 02 | 27252 | DateTime | RD | - | I3(L3) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 02 | 27254 | DateTime | RD | - | I4(L4) - date/time, max. average current value - highest value drag indicator
Basic device measuring group 02 | 27256 | Float | RD | VA | I1(L1)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27258 | Float | RD | VA | I2(L2)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27260 | Float | RD | VA | I3(L3)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27262 | Float | RD | VA | I4(L4)/S - max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27264 | DateTime | RD | - | I1(L1)/S - date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27266 | DateTime | RD | - | I2(L2)/S - date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27268 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27270 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - highest value drag indicator
Basic device measuring group 02 | 27272 | Float | RD | W | I1(L1)/P - max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27274 | Float | RD | W | I2(L2)/P - max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27276 | Float | RD | W | I3(L3)/P - max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27278 | Float | RD | W | I4(L4)/P - max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27280 | DateTime | RD | - | I1(L1)/P - date/time, max. average active power value - highest value drag indicator
Basic device measuring group 02 | 27282 | DateTime | RD | - | I2(L2)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27284 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27286 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - highest value drag indicator
Basic device measuring group 02 | 27288 | Float | RD | W | I1(L1)/P- min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27290 | Float | RD | W | I2(L2)/P- min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27292 | Float | RD | W | I3(L3)/P- min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27294 | Float | RD | W | I4(L4)/P- min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27296 | DateTime | RD | - | I1(L1)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27298 | DateTime | RD | - | I2(L2)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27300 | DateTime | RD | - | I3(L3)/P- Date/time, min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27302 | DateTime | RD | - | I4(L4)/P- date/time, min. active power average value - highest value drag indicator
Basic device measuring group 02 | 27304 | Float | RD | A | I1(L1) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27306 | Float | RD | A | I2(L2) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27308 | Float | RD | A | I3(L3) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27310 | Float | RD | A | I4(L4) - max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27312 | DateTime | RD | - | I1(L1) - date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27314 | DateTime | RD | - | I2(L2) - date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27316 | DateTime | RD | - | I3(L3) - Date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27318 | DateTime | RD | - | I4(L4) - Date/time, max. average current value - 2nd highest value drag indicator
Basic device measuring group 02 | 27320 | Float | RD | VA | I1(L1)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27322 | Float | RD | VA | I2(L2)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27324 | Float | RD | VA | I3(L3)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27326 | Float | RD | VA | I4(L4)/S - max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27328 | DateTime | RD | - | I1(L1)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27330 | DateTime | RD | - | I2(L2)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27332 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27334 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27336 | Float | RD | W | I1(L1)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27338 | Float | RD | W | I2(L2)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27340 | Float | RD | W | I3(L3)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27342 | Float | RD | W | I4(L4)/P - max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27344 | DateTime | RD | - | I1(L1)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27346 | DateTime | RD | - | I2(L2)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27348 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27350 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27352 | Float | RD | W | I1(L1)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27354 | Float | RD | W | I2(L2)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27356 | Float | RD | W | I3(L3)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27358 | Float | RD | W | I4(L4)/P- min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27360 | DateTime | RD | - | I1(L1)/P- Date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27362 | DateTime | RD | - | I2(L2)/P- Date/time, min. average active power value - 2nd highest value drag indicator
Basic device measuring group 02 | 27364 | DateTime | RD | - | I3(L3)/P- Date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27366 | DateTime | RD | - | I4(L4)/P- date/time, min. active power average value - 2nd highest value drag indicator
Basic device measuring group 02 | 27368 | Float | RD | A | I1(L1) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27370 | Float | RD | A | I2(L2) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27372 | Float | RD | A | I3(L3) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27374 | Float | RD | A | I4(L4) - max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27376 | DateTime | RD | - | I1(L1) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27378 | DateTime | RD | - | I2(L2) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27380 | DateTime | RD | - | I3(L3) - date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27382 | DateTime | RD | - | I4(L4) - Date/time, max. average current value - 3rd highest value drag indicator
Basic device measuring group 02 | 27384 | Float | RD | VA | I1(L1)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27386 | Float | RD | VA | I2(L2)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27388 | Float | RD | VA | I3(L3)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27390 | Float | RD | VA | I4(L4)/S - max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27392 | DateTime | RD | - | I1(L1)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27394 | DateTime | RD | - | I2(L2)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27396 | DateTime | RD | - | I3(L3)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27398 | DateTime | RD | - | I4(L4)/S - Date/time, max. average apparent power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27400 | Float | RD | W | I1(L1)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27402 | Float | RD | W | I2(L2)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27404 | Float | RD | W | I3(L3)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27406 | Float | RD | W | I4(L4)/P - max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27408 | DateTime | RD | - | I1(L1)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27410 | DateTime | RD | - | I2(L2)/P - Date/time, max. average active power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27412 | DateTime | RD | - | I3(L3)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27414 | DateTime | RD | - | I4(L4)/P - Date/time, max. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27416 | Float | RD | W | I1(L1)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27418 | Float | RD | W | I2(L2)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27420 | Float | RD | W | I3(L3)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27422 | Float | RD | W | I4(L4)/P- min. active power average value - 3rd highest value drag indicator
Basic device measuring group 02 | 27424 | DateTime | RD | - | I1(L1)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27426 | DateTime | RD | - | I2(L2)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27428 | DateTime | RD | - | I3(L3)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27430 | DateTime | RD | - | I4(L4)/P- Date/time, min. average active power value - 3rd highest value drag indicator
Basic device measuring group 02 | 27432 | Float | RD | A | Sum. I1(L1)-I4(L4) of the max. average current values - highest value drag indicator
Basic device measuring group 02 | 27434 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - highest value drag indicator
Basic device measuring group 02 | 27436 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - highest value drag indicator
Basic device measuring group 02 | 27438 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - highest value drag indicator
Basic device measuring group 02 | 27440 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - highest value drag indicator
Basic device measuring group 02 | 27442 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. active power mean values - highest value drag indicator
Basic device measuring group 02 | 27444 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - highest value drag indicator
Basic device measuring group 02 | 27446 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power mean values - highest value drag indicator
Basic device measuring group 02 | 27448 | Float | RD | A | Sum. I1(L1)-I4(L4) of max. average current values - 2nd highest value drag indicator
Basic device measuring group 02 | 27450 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - 2nd highest value drag indicator
Basic device measuring group 02 | 27452 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - 2nd highest value drag indicator
Basic device measuring group 02 | 27454 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - 2nd highest value drag indicator
Basic device measuring group 02 | 27456 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - 2nd highest value drag indicator
Basic device measuring group 02 | 27458 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. average active power values - 2nd highest value drag indicator
Basic device measuring group 02 | 27460 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - 2nd highest value drag indicator
Basic device measuring group 02 | 27462 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power average values - 2nd highest value drag indicator
Basic device measuring group 02 | 27464 | Float | RD | A | Sum. I1(L1)-I4(L4) of max. average current values - 3rd highest value drag indicator
Basic device measuring group 02 | 27466 | DateTime | RD | - | Sum. I1(L1)-I4(L4) - date/time, max. average current values - 3rd highest value drag indicator
Basic device measuring group 02 | 27468 | Float | RD | VA | Sum. I1(L1)-I4(L4)/S of max. average apparent power values - 3rd highest value drag indicator
Basic device measuring group 02 | 27470 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/S - date/time, max. average apparent power values - 3rd highest value drag indicator
Basic device measuring group 02 | 27472 | Float | RD | W | Sum. I1(L1)-I4(L4)/P of max. average active power values - 3rd highest value drag indicator
Basic device measuring group 02 | 27474 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P - date/time, max. active power mean values - 3rd highest value drag indicator
Basic device measuring group 02 | 27476 | Float | RD | W | Sum. I1(L1)-I4(L4)/P- of min. active power mean values - 3rd highest value drag indicator
Basic device measuring group 02 | 27478 | DateTime | RD | - | Sum. I1(L1)-I4(L4)/P- date/time, min. active power mean values - 3rd highest value drag indicator
Module 01 DI 14 | 37000 | Boolean | RD | - | Digital-In1, status
Module 01 DI 14 | 37001 | Boolean | RD | - | Digital-In2, status
Module 01 DI 14 | 37002 | Boolean | RD | - | Digital-In3, status
Module 01 DI 14 | 37003 | Boolean | RD | - | Digital-In4, status
Module 01 DI 14 | 37004 | Boolean | RD | - | Digital-In5, status
Module 01 DI 14 | 37005 | Boolean | RD | - | Digital-In6, status
Module 01 DI 14 | 37006 | Boolean | RD | - | Digital-In7, status
Module 01 DI 14 | 37007 | Boolean | RD | - | Digital-In8, status
Module 01 DI 14 | 37008 | Boolean | RD | - | Digital-In9, status
Module 01 DI 14 | 37009 | Boolean | RD | - | Digital-In10, status
Module 01 DI 14 | 37010 | Boolean | RD | - | Digital-In11, status
Module 01 DI 14 | 37011 | Boolean | RD | - | Digital-In12, status
Module 01 DI 14 | 37012 | Boolean | RD | - | Digital-In13, status
Module 01 DI 14 | 37013 | Boolean | RD | - | Digital-In14, status
Module 01 DI 14 | 37014 | UInt16 | RD | - | Bit field of the digital inputs
Module 02 DI 14 | 37071 | Boolean | RD | - | Digital-In1, status
Module 02 DI 14 | 37072 | Boolean | RD | - | Digital-In2, status
Module 02 DI 14 | 37073 | Boolean | RD | - | Digital-In3, status
Module 02 DI 14 | 37074 | Boolean | RD | - | Digital-In4, status
Module 02 DI 14 | 37075 | Boolean | RD | - | Digital-In5, status
Module 02 DI 14 | 37076 | Boolean | RD | - | Digital-In6, status
Module 02 DI 14 | 37077 | Boolean | RD | - | Digital-In7, status
Module 02 DI 14 | 37078 | Boolean | RD | - | Digital-In8, status
Module 02 DI 14 | 37079 | Boolean | RD | - | Digital-In9, status
Module 02 DI 14 | 37080 | Boolean | RD | - | Digital-In10, status
Module 02 DI 14 | 37081 | Boolean | RD | - | Digital-In11, status
Module 02 DI 14 | 37082 | Boolean | RD | - | Digital-In12, status
Module 02 DI 14 | 37083 | Boolean | RD | - | Digital-In13, status
Module 02 DI 14 | 37084 | Boolean | RD | - | Digital-In14, status
Module 02 DI 14 | 37085 | UInt16 | RD | - | Bit field of the digital inputs
Module 03 DI 14 | 37142 | Boolean | RD | - | Digital-In1, status
Module 03 DI 14 | 37143 | Boolean | RD | - | Digital-In2, status
Module 03 DI 14 | 37144 | Boolean | RD | - | Digital-In3, status
Module 03 DI 14 | 37145 | Boolean | RD | - | Digital-In4, status
Module 03 DI 14 | 37146 | Boolean | RD | - | Digital-In5, status
Module 03 DI 14 | 37147 | Boolean | RD | - | Digital-In6, status
Module 03 DI 14 | 37148 | Boolean | RD | - | Digital-In7, status
Module 03 DI 14 | 37149 | Boolean | RD | - | Digital-In8, status
Module 03 DI 14 | 37150 | Boolean | RD | - | Digital-In9, status
Module 03 DI 14 | 37151 | Boolean | RD | - | Digital-In10, status
Module 03 DI 14 | 37152 | Boolean | RD | - | Digital-In11, status
Module 03 DI 14 | 37153 | Boolean | RD | - | Digital-In12, status
Module 03 DI 14 | 37154 | Boolean | RD | - | Digital-In13, status
Module 03 DI 14 | 37155 | Boolean | RD | - | Digital-In14, status
Module 03 DI 14 | 37156 | UInt16 | RD | - | Bit field of the digital inputs
Module 04 DI 14 | 37213 | Boolean | RD | - | Digital-In1, status
Module 04 DI 14 | 37214 | Boolean | RD | - | Digital-In2, status
Module 04 DI 14 | 37215 | Boolean | RD | - | Digital-In3, status
Module 04 DI 14 | 37216 | Boolean | RD | - | Digital-In4, status
Module 04 DI 14 | 37217 | Boolean | RD | - | Digital-In5, status
Module 04 DI 14 | 37218 | Boolean | RD | - | Digital-In6, status
Module 04 DI 14 | 37219 | Boolean | RD | - | Digital-In7, status
Module 04 DI 14 | 37220 | Boolean | RD | - | Digital-In8, status
Module 04 DI 14 | 37221 | Boolean | RD | - | Digital-In9, status
Module 04 DI 14 | 37222 | Boolean | RD | - | Digital-In10, status
Module 04 DI 14 | 37223 | Boolean | RD | - | Digital-In11, status
Module 04 DI 14 | 37224 | Boolean | RD | - | Digital-In12, status
Module 04 DI 14 | 37225 | Boolean | RD | - | Digital-In13, status
Module 04 DI 14 | 37226 | Boolean | RD | - | Digital-In14, status
Module 04 DI 14 | 37227 | UInt16 | RD | - | Bit field of the digital inputs
Module 05 DI 14 | 37284 | Boolean | RD | - | Digital-In1, status
Module 05 DI 14 | 37285 | Boolean | RD | - | Digital-In2, status
Module 05 DI 14 | 37286 | Boolean | RD | - | Digital-In3, status
Module 05 DI 14 | 37287 | Boolean | RD | - | Digital-In4, status
Module 05 DI 14 | 37288 | Boolean | RD | - | Digital-In5, status
Module 05 DI 14 | 37289 | Boolean | RD | - | Digital-In6, status
Module 05 DI 14 | 37290 | Boolean | RD | - | Digital-In7, status
Module 05 DI 14 | 37291 | Boolean | RD | - | Digital-In8, status
Module 05 DI 14 | 37292 | Boolean | RD | - | Digital-In9, status
Module 05 DI 14 | 37293 | Boolean | RD | - | Digital-In10, status
Module 05 DI 14 | 37294 | Boolean | RD | - | Digital-In11, status
Module 05 DI 14 | 37295 | Boolean | RD | - | Digital-In12, status
Module 05 DI 14 | 37296 | Boolean | RD | - | Digital-In13, status
Module 05 DI 14 | 37297 | Boolean | RD | - | Digital-In14, status
Module 05 DI 14 | 37298 | UInt16 | RD | - | Bit field of the digital inputs
Module 06 DI 14 | 37355 | Boolean | RD | - | Digital-In1, status
Module 06 DI 14 | 37356 | Boolean | RD | - | Digital-In2, status
Module 06 DI 14 | 37357 | Boolean | RD | - | Digital-In3, status
Module 06 DI 14 | 37358 | Boolean | RD | - | Digital-In4, status
Module 06 DI 14 | 37359 | Boolean | RD | - | Digital-In5, status
Module 06 DI 14 | 37360 | Boolean | RD | - | Digital-In6, status
Module 06 DI 14 | 37361 | Boolean | RD | - | Digital-In7, status
Module 06 DI 14 | 37362 | Boolean | RD | - | Digital-In8, status
Module 06 DI 14 | 37363 | Boolean | RD | - | Digital-In9, status
Module 06 DI 14 | 37364 | Boolean | RD | - | Digital-In10, status
Module 06 DI 14 | 37365 | Boolean | RD | - | Digital-In11, status
Module 06 DI 14 | 37366 | Boolean | RD | - | Digital-In12, status
Module 06 DI 14 | 37367 | Boolean | RD | - | Digital-In13, status
Module 06 DI 14 | 37368 | Boolean | RD | - | Digital-In14, status
Module 06 DI 14 | 37369 | UInt16 | RD | - | Bit field of the digital inputs
Module 07 DI 14 | 37426 | Boolean | RD | - | Digital-In1, status
Module 07 DI 14 | 37427 | Boolean | RD | - | Digital-In2, status
Module 07 DI 14 | 37428 | Boolean | RD | - | Digital-In3, status
Module 07 DI 14 | 37429 | Boolean | RD | - | Digital-In4, status
Module 07 DI 14 | 37430 | Boolean | RD | - | Digital-In5, status
Module 07 DI 14 | 37431 | Boolean | RD | - | Digital-In6, status
Module 07 DI 14 | 37432 | Boolean | RD | - | Digital-In7, status
Module 07 DI 14 | 37433 | Boolean | RD | - | Digital-In8, status
Module 07 DI 14 | 37434 | Boolean | RD | - | Digital-In9, status
Module 07 DI 14 | 37435 | Boolean | RD | - | Digital-In10, status
Module 07 DI 14 | 37436 | Boolean | RD | - | Digital-In11, status
Module 07 DI 14 | 37437 | Boolean | RD | - | Digital-In12, status
Module 07 DI 14 | 37438 | Boolean | RD | - | Digital-In13, status
Module 07 DI 14 | 37439 | Boolean | RD | - | Digital-In14, status
Module 07 DI 14 | 37440 | UInt16 | RD | - | Bit field of the digital inputs
Module 08 DI 14 | 37497 | Boolean | RD | - | Digital-In1, status
Module 08 DI 14 | 37498 | Boolean | RD | - | Digital-In2, status
Module 08 DI 14 | 37499 | Boolean | RD | - | Digital-In3, status
Module 08 DI 14 | 37500 | Boolean | RD | - | Digital-In4, status
Module 08 DI 14 | 37501 | Boolean | RD | - | Digital-In5, status
Module 08 DI 14 | 37502 | Boolean | RD | - | Digital-In6, status
Module 08 DI 14 | 37503 | Boolean | RD | - | Digital-In7, status
Module 08 DI 14 | 37504 | Boolean | RD | - | Digital-In8, status
Module 08 DI 14 | 37505 | Boolean | RD | - | Digital-In9, status
Module 08 DI 14 | 37506 | Boolean | RD | - | Digital-In10, status
Module 08 DI 14 | 37507 | Boolean | RD | - | Digital-In11, status
Module 08 DI 14 | 37508 | Boolean | RD | - | Digital-In12, status
Module 08 DI 14 | 37509 | Boolean | RD | - | Digital-In13, status
Module 08 DI 14 | 37510 | Boolean | RD | - | Digital-In14, status
Module 08 DI 14 | 37511 | UInt16 | RD | - | Bit field of the digital inputs
Module 09 DI 14 | 37568 | Boolean | RD | - | Digital-In1, status
Module 09 DI 14 | 37569 | Boolean | RD | - | Digital-In2, status
Module 09 DI 14 | 37570 | Boolean | RD | - | Digital-In3, status
Module 09 DI 14 | 37571 | Boolean | RD | - | Digital-In4, status
Module 09 DI 14 | 37572 | Boolean | RD | - | Digital-In5, status
Module 09 DI 14 | 37573 | Boolean | RD | - | Digital-In6, status
Module 09 DI 14 | 37574 | Boolean | RD | - | Digital-In7, status
Module 09 DI 14 | 37575 | Boolean | RD | - | Digital-In8, status
Module 09 DI 14 | 37576 | Boolean | RD | - | Digital-In9, status
Module 09 DI 14 | 37577 | Boolean | RD | - | Digital-In10, status
Module 09 DI 14 | 37578 | Boolean | RD | - | Digital-In11, status
Module 09 DI 14 | 37579 | Boolean | RD | - | Digital-In12, status
Module 09 DI 14 | 37580 | Boolean | RD | - | Digital-In13, status
Module 09 DI 14 | 37581 | Boolean | RD | - | Digital-In14, status
Module 09 DI 14 | 37582 | UInt16 | RD | - | Bit field of the digital inputs
Module 10 DI 14 | 37639 | Boolean | RD | - | Digital-In1, status
Module 10 DI 14 | 37640 | Boolean | RD | - | Digital-In2, status
Module 10 DI 14 | 37641 | Boolean | RD | - | Digital-In3, status
Module 10 DI 14 | 37642 | Boolean | RD | - | Digital-In4, status
Module 10 DI 14 | 37643 | Boolean | RD | - | Digital-In5, status
Module 10 DI 14 | 37644 | Boolean | RD | - | Digital-In6, status
Module 10 DI 14 | 37645 | Boolean | RD | - | Digital-In7, status
Module 10 DI 14 | 37646 | Boolean | RD | - | Digital-In8, status
Module 10 DI 14 | 37647 | Boolean | RD | - | Digital-In9, status
Module 10 DI 14 | 37648 | Boolean | RD | - | Digital-In10, status
Module 10 DI 14 | 37649 | Boolean | RD | - | Digital-In11, status
Module 10 DI 14 | 37650 | Boolean | RD | - | Digital-In12, status
Module 10 DI 14 | 37651 | Boolean | RD | - | Digital-In13, status
Module 10 DI 14 | 37652 | Boolean | RD | - | Digital-In14, status
Module 10 DI 14 | 37653 | UInt16 | RD | - | Bit field of the digital inputs
basic device | 41000 | Boolean | RD | - | Comparator group 1
basic device | 41001 | Boolean | RD | - | Comparator group 2
basic device | 41002 | Boolean | RD | - | Comparator group 3
basic device | 41003 | Boolean | RD | - | Comparator group 4
basic device | 41004 | Boolean | RD | - | Comparator group 5
basic device | 41005 | Boolean | RD | - | Comparator group 6
basic device | 41006 | Boolean | RD | - | Comparator group 7
basic device | 41007 | Boolean | RD | - | Comparator group 8
basic device | 41008 | Boolean | RD | - | Comparator group 9
basic device | 41009 | Boolean | RD | - | Comparator group 10
basic device | 41010 | Boolean | RD | - | Comparator group 11
basic device | 41011 | Boolean | RD | - | Comparator group 12
basic device | 41012 | Boolean | RD | - | Comparator group 13
basic device | 41013 | Boolean | RD | - | Comparator group 14
basic device | 41014 | Boolean | RD | - | Comparator group 15
basic device | 41015 | Boolean | RD | - | Comparator group 16
basic device | 41016 | Boolean | RD | - | Comparator group 17
basic device | 41017 | Boolean | RD | - | Comparator group 18
basic device | 41018 | Boolean | RD | - | Comparator group 19
basic device | 41019 | Boolean | RD | - | Comparator group 20
basic device | 41020 | Boolean | RD | - | Comparator group 21
basic device | 41021 | Boolean | RD | - | Comparator group 22
basic device | 41022 | Boolean | RD | - | Comparator group 23
basic device | 41023 | Boolean | RD | - | Comparator group 24
basic device | 41024 | Boolean | RD | - | Comparator group 25
basic device | 41025 | Boolean | RD | - | Comparator group 26
basic device | 41026 | Boolean | RD | - | Comparator group 27
basic device | 41027 | Boolean | RD | - | Comparator group 28
basic device | 41028 | Boolean | RD | - | Comparator group 29
basic device | 41029 | Boolean | RD | - | Comparator group 30
basic device | 41030 | Boolean | RD | - | Comparator group 31
basic device | 41031 | Boolean | RD | - | Comparator group 32
basic device | 41032 | Boolean | RD | - | Value comparator 1
basic device | 41033 | Float | RD | - | Constant Comparator 1
basic device | 41035 | Boolean | RD | - | Value comparator 2
basic device | 41036 | Float | RD | - | Constant Comparator 2
basic device | 41038 | Boolean | RD | - | Value comparator 3
basic device | 41039 | Float | RD | - | Constant Comparator 3
basic device | 41041 | Boolean | RD | - | Value comparator 4
basic device | 41042 | Float | RD | - | Constant Comparator 4
basic device | 41044 | Boolean | RD | - | Value comparator 5
basic device | 41045 | Float | RD | - | Constant Comparator 5
basic device | 41047 | Boolean | RD | - | Value Comparator 6
basic device | 41048 | Float | RD | - | Constant Comparator 6
basic device | 41050 | Boolean | RD | - | Value Comparator 7
basic device | 41051 | Float | RD | - | Constant Comparator 7
basic device | 41053 | Boolean | RD | - | Value Comparator 8
basic device | 41054 | Float | RD | - | Constant Comparator 8
basic device | 41056 | Boolean | RD | - | Value Comparator 9
basic device | 41057 | Float | RD | - | Constant Comparator 9
basic device | 41059 | Boolean | RD | - | Value Comparator 10
basic device | 41060 | Float | RD | - | Constant Comparator 10
basic device | 41062 | Boolean | RD | - | Value Comparator 11
basic device | 41063 | Float | RD | - | Constant Comparator 11
basic device | 41065 | Boolean | RD | - | Value Comparator 12
basic device | 41066 | Float | RD | - | Constant Comparator 12
basic device | 41068 | Boolean | RD | - | Value comparator 13
basic device | 41069 | Float | RD | - | Constant Comparator 13
basic device | 41071 | Boolean | RD | - | Value Comparator 14
basic device | 41072 | Float | RD | - | Constant Comparator 14
basic device | 41074 | Boolean | RD | - | Value Comparator 15
basic device | 41075 | Float | RD | - | Constant Comparator 15
basic device | 41077 | Boolean | RD | - | Value Comparator 16
basic device | 41078 | Float | RD | - | Constant Comparator 16
basic device | 41080 | Boolean | RD | - | Value comparator 17
basic device | 41081 | Float | RD | - | Constant Comparator 17
basic device | 41083 | Boolean | RD | - | Value Comparator 18
basic device | 41084 | Float | RD | - | Constant Comparator 18
basic device | 41086 | Boolean | RD | - | Value comparator 19
basic device | 41087 | Float | RD | - | Constant Comparator 19
basic device | 41089 | Boolean | RD | - | Value comparator 20
basic device | 41090 | Float | RD | - | Constant Comparator 20
basic device | 41092 | Boolean | RD | - | Value comparator 21
basic device | 41093 | Float | RD | - | Constant Comparator 21
basic device | 41095 | Boolean | RD | - | Value comparator 22
basic device | 41096 | Float | RD | - | Constant Comparator 22
basic device | 41098 | Boolean | RD | - | Value comparator 23
basic device | 41099 | Float | RD | - | Constant Comparator 23
basic device | 41101 | Boolean | RD | - | Value comparator 24
basic device | 41102 | Float | RD | - | Constant Comparator 24
basic device | 41104 | Boolean | RD | - | Value comparator 25
basic device | 41105 | Float | RD | - | Constant Comparator 25
basic device | 41107 | Boolean | RD | - | Value comparator 26
basic device | 41108 | Float | RD | - | Constant Comparator 26
basic device | 41110 | Boolean | RD | - | Value comparator 27
basic device | 41111 | Float | RD | - | Constant Comparator 27
basic device | 41113 | Boolean | RD | - | Value comparator 28
basic device | 41114 | Float | RD | - | Constant Comparator 28
basic device | 41116 | Boolean | RD | - | Value comparator 29
basic device | 41117 | Float | RD | - | Constant Comparator 29
basic device | 41119 | Boolean | RD | - | Value comparator 30
basic device | 41120 | Float | RD | - | Constant Comparator 30
basic device | 41122 | Boolean | RD | - | Value comparator 31
basic device | 41123 | Float | RD | - | Constant Comparator 31
basic device | 41125 | Boolean | RD | - | Value comparator 32
basic device | 41126 | Float | RD | - | Constant Comparator 32
basic device | 41128 | Boolean | RD | - | Value comparator 33
basic device | 41129 | Float | RD | - | Constant Comparator 33
basic device | 41131 | Boolean | RD | - | Value comparator 34
basic device | 41132 | Float | RD | - | Constant Comparator 34
basic device | 41134 | Boolean | RD | - | Value comparator 35
basic device | 41135 | Float | RD | - | Constant Comparator 35
basic device | 41137 | Boolean | RD | - | Value Comparator 36
basic device | 41138 | Float | RD | - | Constant Comparator 36
basic device | 41140 | Boolean | RD | - | Value comparator 37
basic device | 41141 | Float | RD | - | Constant Comparator 37
basic device | 41143 | Boolean | RD | - | Value comparator 38
basic device | 41144 | Float | RD | - | Constant Comparator 38
basic device | 41146 | Boolean | RD | - | Value comparator 39
basic device | 41147 | Float | RD | - | Constant Comparator 39
basic device | 41149 | Boolean | RD | - | Value comparator 40
basic device | 41150 | Float | RD | - | Constant Comparator 40
basic device | 41152 | Boolean | RD | - | Value comparator 41
basic device | 41153 | Float | RD | - | Constant Comparator 41
basic device | 41155 | Boolean | RD | - | Value comparator 42
basic device | 41156 | Float | RD | - | Constant Comparator 42
basic device | 41158 | Boolean | RD | - | Value comparator 43
basic device | 41159 | Float | RD | - | Constant Comparator 43
basic device | 41161 | Boolean | RD | - | Value Comparator 44
basic device | 41162 | Float | RD | - | Constant Comparator 44
basic device | 41164 | Boolean | RD | - | Value comparator 45
basic device | 41165 | Float | RD | - | Constant Comparator 45
basic device | 41167 | Boolean | RD | - | Value comparator 46
basic device | 41168 | Float | RD | - | Constant Comparator 46
basic device | 41170 | Boolean | RD | - | Value comparator 47
basic device | 41171 | Float | RD | - | Constant Comparator 47
basic device | 41173 | Boolean | RD | - | Value Comparator 48
basic device | 41174 | Float | RD | - | Constant Comparator 48
basic device | 41176 | Boolean | RD | - | Value Comparator 49
basic device | 41177 | Float | RD | - | Constant Comparator 49
basic device | 41179 | Boolean | RD | - | Value comparator 50
basic device | 41180 | Float | RD | - | Constant Comparator 50
basic device | 41182 | Boolean | RD | - | Value comparator 51
basic device | 41183 | Float | RD | - | Constant Comparator 51
basic device | 41185 | Boolean | RD | - | Value comparator 52
basic device | 41186 | Float | RD | - | Constant Comparator 52
basic device | 41188 | Boolean | RD | - | Value comparator 53
basic device | 41189 | Float | RD | - | Constant Comparator 53
basic device | 41191 | Boolean | RD | - | Value Comparator 54
basic device | 41192 | Float | RD | - | Constant Comparator 54
basic device | 41194 | Boolean | RD | - | Value comparator 55
basic device | 41195 | Float | RD | - | Constant Comparator 55
basic device | 41197 | Boolean | RD | - | Value comparator 56
basic device | 41198 | Float | RD | - | Constant Comparator 56
basic device | 41200 | Boolean | RD | - | Value comparator 57
basic device | 41201 | Float | RD | - | Constant Comparator 57
basic device | 41203 | Boolean | RD | - | Value comparator 58
basic device | 41204 | Float | RD | - | Constant Comparator 58
basic device | 41206 | Boolean | RD | - | Value comparator 59
basic device | 41207 | Float | RD | - | Constant Comparator 59
basic device | 41209 | Boolean | RD | - | Value comparator 60
basic device | 41210 | Float | RD | - | Constant Comparator 60
basic device | 41212 | Boolean | RD | - | Value comparator 61
basic device | 41213 | Float | RD | - | Constant Comparator 61
basic device | 41215 | Boolean | RD | - | Value comparator 62
basic device | 41216 | Float | RD | - | Constant Comparator 62
basic device | 41218 | Boolean | RD | - | Value comparator 63
basic device | 41219 | Float | RD | - | Constant Comparator 63
basic device | 41221 | Boolean | RD | - | Value comparator 64
basic device | 41222 | Float | RD | - | Constant Comparator 64
basic device | 41224 | Boolean | RD | - | Value comparator 65
basic device | 41225 | Float | RD | - | Constant Comparator 65
basic device | 41227 | Boolean | RD | - | Value comparator 66
basic device | 41228 | Float | RD | - | Constant Comparator 66
basic device | 41230 | Boolean | RD | - | Value comparator 67
basic device | 41231 | Float | RD | - | Constant Comparator 67
basic device | 41233 | Boolean | RD | - | Value comparator 68
basic device | 41234 | Float | RD | - | Constant Comparator 68
basic device | 41236 | Boolean | RD | - | Value comparator 69
basic device | 41237 | Float | RD | - | Constant Comparator 69
basic device | 41239 | Boolean | RD | - | Value comparator 70
basic device | 41240 | Float | RD | - | Constant Comparator 70
basic device | 41242 | Boolean | RD | - | Value comparator 71
basic device | 41243 | Float | RD | - | Constant Comparator 71
basic device | 41245 | Boolean | RD | - | Value comparator 72
basic device | 41246 | Float | RD | - | Constant Comparator 72
basic device | 41248 | Boolean | RD | - | Value comparator 73
basic device | 41249 | Float | RD | - | Constant Comparator 73
basic device | 41251 | Boolean | RD | - | Value comparator 74
basic device | 41252 | Float | RD | - | Constant Comparator 74
basic device | 41254 | Boolean | RD | - | Value comparator 75
basic device | 41255 | Float | RD | - | Constant Comparator 75
basic device | 41257 | Boolean | RD | - | Value comparator 76
basic device | 41258 | Float | RD | - | Constant comparator 76
basic device | 41260 | Boolean | RD | - | Value comparator 77
basic device | 41261 | Float | RD | - | Constant Comparator 77
basic device | 41263 | Boolean | RD | - | Value comparator 78
basic device | 41264 | Float | RD | - | Constant Comparator 78
basic device | 41266 | Boolean | RD | - | Value comparator 79
basic device | 41267 | Float | RD | - | Constant Comparator 79
basic device | 41269 | Boolean | RD | - | Value comparator 80
basic device | 41270 | Float | RD | - | Constant Comparator 80
basic device | 41272 | Boolean | RD | - | Value comparator 81
basic device | 41273 | Float | RD | - | Constant Comparator 81
basic device | 41275 | Boolean | RD | - | Value comparator 82
basic device | 41276 | Float | RD | - | Constant Comparator 82
basic device | 41278 | Boolean | RD | - | Value comparator 83
basic device | 41279 | Float | RD | - | Constant Comparator 83
basic device | 41281 | Boolean | RD | - | Value comparator 84
basic device | 41282 | Float | RD | - | Constant Comparator 84
basic device | 41284 | Boolean | RD | - | Value comparator 85
basic device | 41285 | Float | RD | - | Constant Comparator 85
basic device | 41287 | Boolean | RD | - | Value comparator 86
basic device | 41288 | Float | RD | - | Constant Comparator 86
basic device | 41290 | Boolean | RD | - | Value comparator 87
basic device | 41291 | Float | RD | - | Constant Comparator 87
basic device | 41293 | Boolean | RD | - | Value comparator 88
basic device | 41294 | Float | RD | - | Constant Comparator 88
basic device | 41296 | Boolean | RD | - | Value comparator 89
basic device | 41297 | Float | RD | - | Constant Comparator 89
basic device | 41299 | Boolean | RD | - | Value comparator 90
basic device | 41300 | Float | RD | - | Constant Comparator 90
basic device | 41302 | Boolean | RD | - | Value comparator 91
basic device | 41303 | Float | RD | - | Constant Comparator 91
basic device | 41305 | Boolean | RD | - | Value comparator 92
basic device | 41306 | Float | RD | - | Constant Comparator 92
basic device | 41308 | Boolean | RD | - | Value comparator 93
basic device | 41309 | Float | RD | - | Constant Comparator 93
basic device | 41311 | Boolean | RD | - | Value comparator 94
basic device | 41312 | Float | RD | - | Constant Comparator 94
basic device | 41314 | Boolean | RD | - | Value comparator 95
basic device | 41315 | Float | RD | - | Constant Comparator 95
basic device | 41317 | Boolean | RD | - | Value comparator 96
basic device | 41318 | Float | RD | - | Constant Comparator 96
basic device | 41320 | Boolean | RD | - | Value comparator 97
basic device | 41321 | Float | RD | - | Constant Comparator 97
basic device | 41323 | Boolean | RD | - | Value comparator 98
basic device | 41324 | Float | RD | - | Constant Comparator 98
basic device | 41326 | Boolean | RD | - | Value comparator 99
basic device | 41327 | Float | RD | - | Constant Comparator 99
basic device | 41329 | Boolean | RD | - | Value comparator 100
basic device | 41330 | Float | RD | - | Constant Comparator 100
basic device | 41332 | Boolean | RD | - | Value comparator 101
basic device | 41333 | Float | RD | - | Constant Comparator 101
basic device | 41335 | Boolean | RD | - | Value comparator 102
basic device | 41336 | Float | RD | - | Constant Comparator 102
basic device | 41338 | Boolean | RD | - | Value comparator 103
basic device | 41339 | Float | RD | - | Constant Comparator 103
basic device | 41341 | Boolean | RD | - | Value comparator 104
basic device | 41342 | Float | RD | - | Constant Comparator 104
basic device | 41344 | Boolean | RD | - | Value comparator 105
basic device | 41345 | Float | RD | - | Constant comparator 105
basic device | 41347 | Boolean | RD | - | Value comparator 106
basic device | 41348 | Float | RD | - | Constant comparator 106
basic device | 41350 | Boolean | RD | - | Value comparator 107
basic device | 41351 | Float | RD | - | Constant comparator 107
basic device | 41353 | Boolean | RD | - | Value comparator 108
basic device | 41354 | Float | RD | - | Constant Comparator 108
basic device | 41356 | Boolean | RD | - | Value comparator 109
basic device | 41357 | Float | RD | - | Constant Comparator 109
basic device | 41359 | Boolean | RD | - | Value comparator 110
basic device | 41360 | Float | RD | - | Constant Comparator 110
basic device | 41362 | Boolean | RD | - | Value comparator 111
basic device | 41363 | Float | RD | - | Constant Comparator 111
basic device | 41365 | Boolean | RD | - | Value comparator 112
basic device | 41366 | Float | RD | - | Constant Comparator 112
basic device | 41368 | Boolean | RD | - | Value comparator 113
basic device | 41369 | Float | RD | - | Constant Comparator 113
basic device | 41371 | Boolean | RD | - | Value comparator 114
basic device | 41372 | Float | RD | - | Constant Comparator 114
basic device | 41374 | Boolean | RD | - | Value comparator 115
basic device | 41375 | Float | RD | - | Constant Comparator 115
basic device | 41377 | Boolean | RD | - | Value comparator 116
basic device | 41378 | Float | RD | - | Constant Comparator 116
basic device | 41380 | Boolean | RD | - | Value comparator 117
basic device | 41381 | Float | RD | - | Constant Comparator 117
basic device | 41383 | Boolean | RD | - | Value comparator 118
basic device | 41384 | Float | RD | - | Constant Comparator 118
basic device | 41386 | Boolean | RD | - | Value comparator 119
basic device | 41387 | Float | RD | - | Constant comparator 119
basic device | 41389 | Boolean | RD | - | Value comparator 120
basic device | 41390 | Float | RD | - | Constant Comparator 120
basic device | 41392 | Boolean | RD | - | Value comparator 121
basic device | 41393 | Float | RD | - | Constant Comparator 121
basic device | 41395 | Boolean | RD | - | Value comparator 122
basic device | 41396 | Float | RD | - | Constant Comparator 122
basic device | 41398 | Boolean | RD | - | Value comparator 123
basic device | 41399 | Float | RD | - | Constant Comparator 123
basic device | 41401 | Boolean | RD | - | Value Comparator 124
basic device | 41402 | Float | RD | - | Constant Comparator 124
basic device | 41404 | Boolean | RD | - | Value comparator 125
basic device | 41405 | Float | RD | - | Constant Comparator 125
basic device | 41500 | UInt32 | RD | - | Event Counter
basic device | 41502 | UInt32 | RD | - | Transients Counter
basic device | 41504 | UInt64 | RD | - | Active Event Flags
basic device | 42000 | Float | RD | V | 1. harmonic UL1N
basic device | 42002 | Float | RD | V | 2. harmonic UL1N
basic device | 42004 | Float | RD | V | 3. harmonic UL1N
basic device | 42006 | Float | RD | V | 4. harmonic UL1N
basic device | 42008 | Float | RD | V | 5. harmonic UL1N
basic device | 42010 | Float | RD | V | 6. harmonic UL1N
basic device | 42012 | Float | RD | V | 7. harmonic UL1N
basic device | 42014 | Float | RD | V | 8. harmonic UL1N
basic device | 42016 | Float | RD | V | 9. harmonic UL1N
basic device | 42018 | Float | RD | V | 10. harmonic UL1N
basic device | 42020 | Float | RD | V | 11. harmonic UL1N
basic device | 42022 | Float | RD | V | 12. harmonic UL1N
basic device | 42024 | Float | RD | V | 13. harmonic UL1N
basic device | 42026 | Float | RD | V | 14. harmonic UL1N
basic device | 42028 | Float | RD | V | 15. harmonic UL1N
basic device | 42030 | Float | RD | V | 16. harmonic UL1N
basic device | 42032 | Float | RD | V | 17. harmonic UL1N
basic device | 42034 | Float | RD | V | 18. harmonic UL1N
basic device | 42036 | Float | RD | V | 19. harmonic UL1N
basic device | 42038 | Float | RD | V | 20. harmonic UL1N
basic device | 42040 | Float | RD | V | 21. harmonic UL1N
basic device | 42042 | Float | RD | V | 22. harmonic UL1N
basic device | 42044 | Float | RD | V | 23. harmonic UL1N
basic device | 42046 | Float | RD | V | 24. harmonic UL1N
basic device | 42048 | Float | RD | V | 25. harmonic UL1N
basic device | 42050 | Float | RD | V | 26. harmonic UL1N
basic device | 42052 | Float | RD | V | 27. harmonic UL1N
basic device | 42054 | Float | RD | V | 28. harmonic UL1N
basic device | 42056 | Float | RD | V | 29. harmonic UL1N
basic device | 42058 | Float | RD | V | 30. harmonic UL1N
basic device | 42060 | Float | RD | V | 31. harmonic UL1N
basic device | 42062 | Float | RD | V | 32. harmonic UL1N
basic device | 42064 | Float | RD | V | 33. harmonic UL1N
basic device | 42066 | Float | RD | V | 34. harmonic UL1N
basic device | 42068 | Float | RD | V | 35. harmonic UL1N
basic device | 42070 | Float | RD | V | 36. harmonic UL1N
basic device | 42072 | Float | RD | V | 37. harmonic UL1N
basic device | 42074 | Float | RD | V | 38. harmonic UL1N
basic device | 42076 | Float | RD | V | 39. harmonic UL1N
basic device | 42078 | Float | RD | V | 40. harmonic UL1N
basic device | 42080 | Float | RD | V | 1. harmonic UL2N
basic device | 42082 | Float | RD | V | 2. harmonic UL2N
basic device | 42084 | Float | RD | V | 3. harmonic UL2N
basic device | 42086 | Float | RD | V | 4. harmonic UL2N
basic device | 42088 | Float | RD | V | 5. harmonic UL2N
basic device | 42090 | Float | RD | V | 6. harmonic UL2N
basic device | 42092 | Float | RD | V | 7. harmonic UL2N
basic device | 42094 | Float | RD | V | 8. harmonic UL2N
basic device | 42096 | Float | RD | V | 9. harmonic UL2N
basic device | 42098 | Float | RD | V | 10. harmonic UL2N
basic device | 42100 | Float | RD | V | 11. harmonic UL2N
basic device | 42102 | Float | RD | V | 12. harmonic UL2N
basic device | 42104 | Float | RD | V | 13. harmonic UL2N
basic device | 42106 | Float | RD | V | 14. harmonic UL2N
basic device | 42108 | Float | RD | V | 15. harmonic UL2N
basic device | 42110 | Float | RD | V | 16. harmonic UL2N
basic device | 42112 | Float | RD | V | 17. harmonic UL2N
basic device | 42114 | Float | RD | V | 18. harmonic UL2N
basic device | 42116 | Float | RD | V | 19. harmonic UL2N
basic device | 42118 | Float | RD | V | 20. harmonic UL2N
basic device | 42120 | Float | RD | V | 21. harmonic UL2N
basic device | 42122 | Float | RD | V | 22. harmonic UL2N
basic device | 42124 | Float | RD | V | 23. harmonic UL2N
basic device | 42126 | Float | RD | V | 24. harmonic UL2N
basic device | 42128 | Float | RD | V | 25. harmonic UL2N
basic device | 42130 | Float | RD | V | 26. harmonic UL2N
basic device | 42132 | Float | RD | V | 27. harmonic UL2N
basic device | 42134 | Float | RD | V | 28. harmonic UL2N
basic device | 42136 | Float | RD | V | 29. harmonic UL2N
basic device | 42138 | Float | RD | V | 30. harmonic UL2N
basic device | 42140 | Float | RD | V | 31. harmonic UL2N
basic device | 42142 | Float | RD | V | 32. harmonic UL2N
basic device | 42144 | Float | RD | V | 33. harmonic UL2N
basic device | 42146 | Float | RD | V | 34. harmonic UL2N
basic device | 42148 | Float | RD | V | 35. harmonic UL2N
basic device | 42150 | Float | RD | V | 36. harmonic UL2N
basic device | 42152 | Float | RD | V | 37. harmonic UL2N
basic device | 42154 | Float | RD | V | 38. harmonic UL2N
basic device | 42156 | Float | RD | V | 39. harmonic UL2N
basic device | 42158 | Float | RD | V | 40. harmonic UL2N
basic device | 42160 | Float | RD | V | 1. harmonic UL3N
basic device | 42162 | Float | RD | V | 2. harmonic UL3N
basic device | 42164 | Float | RD | V | 3. harmonic UL3N
basic device | 42166 | Float | RD | V | 4. harmonic UL3N
basic device | 42168 | Float | RD | V | 5. harmonic UL3N
basic device | 42170 | Float | RD | V | 6. harmonic UL3N
basic device | 42172 | Float | RD | V | 7. harmonic UL3N
basic device | 42174 | Float | RD | V | 8. harmonic UL3N
basic device | 42176 | Float | RD | V | 9. harmonic UL3N
basic device | 42178 | Float | RD | V | 10. harmonic UL3N
basic device | 42180 | Float | RD | V | 11. harmonic UL3N
basic device | 42182 | Float | RD | V | 12. harmonic UL3N
basic device | 42184 | Float | RD | V | 13. harmonic UL3N
basic device | 42186 | Float | RD | V | 14. harmonic UL3N
basic device | 42188 | Float | RD | V | 15. harmonic UL3N
basic device | 42190 | Float | RD | V | 16. harmonic UL3N
basic device | 42192 | Float | RD | V | 17. harmonic UL3N
basic device | 42194 | Float | RD | V | 18. harmonic UL3N
basic device | 42196 | Float | RD | V | 19. harmonic UL3N
basic device | 42198 | Float | RD | V | 20. harmonic UL3N
basic device | 42200 | Float | RD | V | 21. harmonic UL3N
basic device | 42202 | Float | RD | V | 22. harmonic UL3N
basic device | 42204 | Float | RD | V | 23. harmonic UL3N
basic device | 42206 | Float | RD | V | 24. harmonic UL3N
basic device | 42208 | Float | RD | V | 25. harmonic UL3N
basic device | 42210 | Float | RD | V | 26. harmonic UL3N
basic device | 42212 | Float | RD | V | 27. harmonic UL3N
basic device | 42214 | Float | RD | V | 28. harmonic UL3N
basic device | 42216 | Float | RD | V | 29. harmonic UL3N
basic device | 42218 | Float | RD | V | 30. harmonic UL3N
basic device | 42220 | Float | RD | V | 31. harmonic UL3N
basic device | 42222 | Float | RD | V | 32. harmonic UL3N
basic device | 42224 | Float | RD | V | 33. harmonic UL3N
basic device | 42226 | Float | RD | V | 34. harmonic UL3N
basic device | 42228 | Float | RD | V | 35. harmonic UL3N
basic device | 42230 | Float | RD | V | 36. harmonic UL3N
basic device | 42232 | Float | RD | V | 37. harmonic UL3N
basic device | 42234 | Float | RD | V | 38. harmonic UL3N
basic device | 42236 | Float | RD | V | 39. harmonic UL3N
basic device | 42238 | Float | RD | V | 40. harmonic UL3N
Basic device measuring group 01 | 42240 | Float | RD | A | 1. harmonic I1
Basic device measuring group 01 | 42242 | Float | RD | A | 2. harmonic I1
Basic device measuring group 01 | 42244 | Float | RD | A | 3. harmonic I1
Basic device measuring group 01 | 42246 | Float | RD | A | 4. harmonic I1
Basic device measuring group 01 | 42248 | Float | RD | A | 5. harmonic I1
Basic device measuring group 01 | 42250 | Float | RD | A | 6. harmonic I1
Basic device measuring group 01 | 42252 | Float | RD | A | 7. harmonic I1
Basic device measuring group 01 | 42254 | Float | RD | A | 8. harmonic I1
Basic device measuring group 01 | 42256 | Float | RD | A | 9. harmonic I1
Basic device measuring group 01 | 42258 | Float | RD | A | 10. harmonic I1
Basic device measuring group 01 | 42260 | Float | RD | A | 11. harmonic I1
Basic device measuring group 01 | 42262 | Float | RD | A | 12. harmonic I1
Basic device measuring group 01 | 42264 | Float | RD | A | 13. harmonic I1
Basic device measuring group 01 | 42266 | Float | RD | A | 14. harmonic I1
Basic device measuring group 01 | 42268 | Float | RD | A | 15. harmonic I1
Basic device measuring group 01 | 42270 | Float | RD | A | 16. harmonic I1
Basic device measuring group 01 | 42272 | Float | RD | A | 17. harmonic I1
Basic device measuring group 01 | 42274 | Float | RD | A | 18. harmonic I1
Basic device measuring group 01 | 42276 | Float | RD | A | 19. harmonic I1
Basic device measuring group 01 | 42278 | Float | RD | A | 20. harmonic I1
Basic device measuring group 01 | 42280 | Float | RD | A | 21. harmonic I1
Basic device measuring group 01 | 42282 | Float | RD | A | 22. harmonic I1
Basic device measuring group 01 | 42284 | Float | RD | A | 23. harmonic I1
Basic device measuring group 01 | 42286 | Float | RD | A | 24. harmonic I1
Basic device measuring group 01 | 42288 | Float | RD | A | 25. harmonic I1
Basic device measuring group 01 | 42290 | Float | RD | A | 26. harmonic I1
Basic device measuring group 01 | 42292 | Float | RD | A | 27. harmonic I1
Basic device measuring group 01 | 42294 | Float | RD | A | 28. harmonic I1
Basic device measuring group 01 | 42296 | Float | RD | A | 29. harmonic I1
Basic device measuring group 01 | 42298 | Float | RD | A | 30. harmonic I1
Basic device measuring group 01 | 42300 | Float | RD | A | 31. harmonic I1
Basic device measuring group 01 | 42302 | Float | RD | A | 32. harmonic I1
Basic device measuring group 01 | 42304 | Float | RD | A | 33. harmonic I1
Basic device measuring group 01 | 42306 | Float | RD | A | 34. harmonic I1
Basic device measuring group 01 | 42308 | Float | RD | A | 35. harmonic I1
Basic device measuring group 01 | 42310 | Float | RD | A | 36. harmonic I1
Basic device measuring group 01 | 42312 | Float | RD | A | 37. harmonic I1
Basic device measuring group 01 | 42314 | Float | RD | A | 38. harmonic I1
Basic device measuring group 01 | 42316 | Float | RD | A | 39. harmonic I1
Basic device measuring group 01 | 42318 | Float | RD | A | 40. harmonic I1
Basic device measuring group 01 | 42320 | Float | RD | A | 1. harmonic I2
Basic device measuring group 01 | 42322 | Float | RD | A | 2. harmonic I2
Basic device measuring group 01 | 42324 | Float | RD | A | 3. harmonic I2
Basic device measuring group 01 | 42326 | Float | RD | A | 4. harmonic I2
Basic device measuring group 01 | 42328 | Float | RD | A | 5. harmonic I2
Basic device measuring group 01 | 42330 | Float | RD | A | 6. harmonic I2
Basic device measuring group 01 | 42332 | Float | RD | A | 7. harmonic I2
Basic device measuring group 01 | 42334 | Float | RD | A | 8. harmonic I2
Basic device measuring group 01 | 42336 | Float | RD | A | 9. harmonic I2
Basic device measuring group 01 | 42338 | Float | RD | A | 10. harmonic I2
Basic device measuring group 01 | 42340 | Float | RD | A | 11. harmonic I2
Basic device measuring group 01 | 42342 | Float | RD | A | 12. harmonic I2
Basic device measuring group 01 | 42344 | Float | RD | A | 13. harmonic I2
Basic device measuring group 01 | 42346 | Float | RD | A | 14. harmonic I2
Basic device measuring group 01 | 42348 | Float | RD | A | 15. harmonic I2
Basic device measuring group 01 | 42350 | Float | RD | A | 16. harmonic I2
Basic device measuring group 01 | 42352 | Float | RD | A | 17. harmonic I2
Basic device measuring group 01 | 42354 | Float | RD | A | 18. harmonic I2
Basic device measuring group 01 | 42356 | Float | RD | A | 19. harmonic I2
Basic device measuring group 01 | 42358 | Float | RD | A | 20. harmonic I2
Basic device measuring group 01 | 42360 | Float | RD | A | 21. harmonic I2
Basic device measuring group 01 | 42362 | Float | RD | A | 22. harmonic I2
Basic device measuring group 01 | 42364 | Float | RD | A | 23. harmonic I2
Basic device measuring group 01 | 42366 | Float | RD | A | 24. harmonic I2
Basic device measuring group 01 | 42368 | Float | RD | A | 25. harmonic I2
Basic device measuring group 01 | 42370 | Float | RD | A | 26. harmonic I2
Basic device measuring group 01 | 42372 | Float | RD | A | 27. harmonic I2
Basic device measuring group 01 | 42374 | Float | RD | A | 28. harmonic I2
Basic device measuring group 01 | 42376 | Float | RD | A | 29. harmonic I2
Basic device measuring group 01 | 42378 | Float | RD | A | 30. harmonic I2
Basic device measuring group 01 | 42380 | Float | RD | A | 31. harmonic I2
Basic device measuring group 01 | 42382 | Float | RD | A | 32. harmonic I2
Basic device measuring group 01 | 42384 | Float | RD | A | 33. harmonic I2
Basic device measuring group 01 | 42386 | Float | RD | A | 34. harmonic I2
Basic device measuring group 01 | 42388 | Float | RD | A | 35. harmonic I2
Basic device measuring group 01 | 42390 | Float | RD | A | 36. harmonic I2
Basic device measuring group 01 | 42392 | Float | RD | A | 37. harmonic I2
Basic device measuring group 01 | 42394 | Float | RD | A | 38. harmonic I2
Basic device measuring group 01 | 42396 | Float | RD | A | 39. harmonic I2
Basic device measuring group 01 | 42398 | Float | RD | A | 40. harmonic I2
Basic device measuring group 01 | 42400 | Float | RD | A | 1. harmonic I3
Basic device measuring group 01 | 42402 | Float | RD | A | 2. harmonic I3
Basic device measuring group 01 | 42404 | Float | RD | A | 3. harmonic I3
Basic device measuring group 01 | 42406 | Float | RD | A | 4. harmonic I3
Basic device measuring group 01 | 42408 | Float | RD | A | 5. harmonic I3
Basic device measuring group 01 | 42410 | Float | RD | A | 6. harmonic I3
Basic device measuring group 01 | 42412 | Float | RD | A | 7. harmonic I3
Basic device measuring group 01 | 42414 | Float | RD | A | 8. harmonic I3
Basic device measuring group 01 | 42416 | Float | RD | A | 9. harmonic I3
Basic device measuring group 01 | 42418 | Float | RD | A | 10. harmonic I3
Basic device measuring group 01 | 42420 | Float | RD | A | 11. harmonic I3
Basic device measuring group 01 | 42422 | Float | RD | A | 12. harmonic I3
Basic device measuring group 01 | 42424 | Float | RD | A | 13. harmonic I3
Basic device measuring group 01 | 42426 | Float | RD | A | 14. harmonic I3
Basic device measuring group 01 | 42428 | Float | RD | A | 15. harmonic I3
Basic device measuring group 01 | 42430 | Float | RD | A | 16. harmonic I3
Basic device measuring group 01 | 42432 | Float | RD | A | 17. harmonic I3
Basic device measuring group 01 | 42434 | Float | RD | A | 18. harmonic I3
Basic device measuring group 01 | 42436 | Float | RD | A | 19. harmonic I3
Basic device measuring group 01 | 42438 | Float | RD | A | 20. harmonic I3
Basic device measuring group 01 | 42440 | Float | RD | A | 21. harmonic I3
Basic device measuring group 01 | 42442 | Float | RD | A | 22. harmonic I3
Basic device measuring group 01 | 42444 | Float | RD | A | 23. harmonic I3
Basic device measuring group 01 | 42446 | Float | RD | A | 24. harmonic I3
Basic device measuring group 01 | 42448 | Float | RD | A | 25. harmonic I3
Basic device measuring group 01 | 42450 | Float | RD | A | 26. harmonic I3
Basic device measuring group 01 | 42452 | Float | RD | A | 27. harmonic I3
Basic device measuring group 01 | 42454 | Float | RD | A | 28. harmonic I3
Basic device measuring group 01 | 42456 | Float | RD | A | 29. harmonic I3
Basic device measuring group 01 | 42458 | Float | RD | A | 30. harmonic I3
Basic device measuring group 01 | 42460 | Float | RD | A | 31. harmonic I3
Basic device measuring group 01 | 42462 | Float | RD | A | 32. harmonic I3
Basic device measuring group 01 | 42464 | Float | RD | A | 33. harmonic I3
Basic device measuring group 01 | 42466 | Float | RD | A | 34. harmonic I3
Basic device measuring group 01 | 42468 | Float | RD | A | 35. harmonic I3
Basic device measuring group 01 | 42470 | Float | RD | A | 36. harmonic I3
Basic device measuring group 01 | 42472 | Float | RD | A | 37. harmonic I3
Basic device measuring group 01 | 42474 | Float | RD | A | 38. harmonic I3
Basic device measuring group 01 | 42476 | Float | RD | A | 39. harmonic I3
Basic device measuring group 01 | 42478 | Float | RD | A | 40. harmonic I3
Basic device measuring group 01 | 42480 | Float | RD | A | 1. harmonic I4
Basic device measuring group 01 | 42482 | Float | RD | A | 2. harmonic I4
Basic device measuring group 01 | 42484 | Float | RD | A | 3. harmonic I4
Basic device measuring group 01 | 42486 | Float | RD | A | 4. harmonic I4
Basic device measuring group 01 | 42488 | Float | RD | A | 5. harmonic I4
Basic device measuring group 01 | 42490 | Float | RD | A | 6. harmonic I4
Basic device measuring group 01 | 42492 | Float | RD | A | 7. harmonic I4
Basic device measuring group 01 | 42494 | Float | RD | A | 8. harmonic I4
Basic device measuring group 01 | 42496 | Float | RD | A | 9. harmonic I4
Basic device measuring group 01 | 42498 | Float | RD | A | 10. harmonic I4
Basic device measuring group 01 | 42500 | Float | RD | A | 11. harmonic I4
Basic device measuring group 01 | 42502 | Float | RD | A | 12. harmonic I4
Basic device measuring group 01 | 42504 | Float | RD | A | 13. harmonic I4
Basic device measuring group 01 | 42506 | Float | RD | A | 14. harmonic I4
Basic device measuring group 01 | 42508 | Float | RD | A | 15. harmonic I4
Basic device measuring group 01 | 42510 | Float | RD | A | 16. harmonic I4
Basic device measuring group 01 | 42512 | Float | RD | A | 17. harmonic I4
Basic device measuring group 01 | 42514 | Float | RD | A | 18. harmonic I4
Basic device measuring group 01 | 42516 | Float | RD | A | 19. harmonic I4
Basic device measuring group 01 | 42518 | Float | RD | A | 20. harmonic I4
Basic device measuring group 01 | 42520 | Float | RD | A | 21. harmonic I4
Basic device measuring group 01 | 42522 | Float | RD | A | 22. harmonic I4
Basic device measuring group 01 | 42524 | Float | RD | A | 23. harmonic I4
Basic device measuring group 01 | 42526 | Float | RD | A | 24. harmonic I4
Basic device measuring group 01 | 42528 | Float | RD | A | 25. harmonic I4
Basic device measuring group 01 | 42530 | Float | RD | A | 26. harmonic I4
Basic device measuring group 01 | 42532 | Float | RD | A | 27. harmonic I4
Basic device measuring group 01 | 42534 | Float | RD | A | 28. harmonic I4
Basic device measuring group 01 | 42536 | Float | RD | A | 29. harmonic I4
Basic device measuring group 01 | 42538 | Float | RD | A | 30. harmonic I4
Basic device measuring group 01 | 42540 | Float | RD | A | 31. harmonic I4
Basic device measuring group 01 | 42542 | Float | RD | A | 32. harmonic I4
Basic device measuring group 01 | 42544 | Float | RD | A | 33. harmonic I4
Basic device measuring group 01 | 42546 | Float | RD | A | 34. harmonic I4
Basic device measuring group 01 | 42548 | Float | RD | A | 35. harmonic I4
Basic device measuring group 01 | 42550 | Float | RD | A | 36. harmonic I4
Basic device measuring group 01 | 42552 | Float | RD | A | 37. harmonic I4
Basic device measuring group 01 | 42554 | Float | RD | A | 38. harmonic I4
Basic device measuring group 01 | 42556 | Float | RD | A | 39. harmonic I4
Basic device measuring group 01 | 42558 | Float | RD | A | 40. harmonic I4
Basic device measuring group 02 | 42560 | Float | RD | A | 1. harmonic I1
Basic device measuring group 02 | 42562 | Float | RD | A | 2. harmonic I1
Basic device measuring group 02 | 42564 | Float | RD | A | 3. harmonic I1
Basic device measuring group 02 | 42566 | Float | RD | A | 4. harmonic I1
Basic device measuring group 02 | 42568 | Float | RD | A | 5. harmonic I1
Basic device measuring group 02 | 42570 | Float | RD | A | 6. harmonic I1
Basic device measuring group 02 | 42572 | Float | RD | A | 7. harmonic I1
Basic device measuring group 02 | 42574 | Float | RD | A | 8. harmonic I1
Basic device measuring group 02 | 42576 | Float | RD | A | 9. harmonic I1
Basic device measuring group 02 | 42578 | Float | RD | A | 10. harmonic I1
Basic device measuring group 02 | 42580 | Float | RD | A | 11. harmonic I1
Basic device measuring group 02 | 42582 | Float | RD | A | 12. harmonic I1
Basic device measuring group 02 | 42584 | Float | RD | A | 13. harmonic I1
Basic device measuring group 02 | 42586 | Float | RD | A | 14. harmonic I1
Basic device measuring group 02 | 42588 | Float | RD | A | 15. harmonic I1
Basic device measuring group 02 | 42590 | Float | RD | A | 16. harmonic I1
Basic device measuring group 02 | 42592 | Float | RD | A | 17. harmonic I1
Basic device measuring group 02 | 42594 | Float | RD | A | 18. harmonic I1
Basic device measuring group 02 | 42596 | Float | RD | A | 19. harmonic I1
Basic device measuring group 02 | 42598 | Float | RD | A | 20. harmonic I1
Basic device measuring group 02 | 42600 | Float | RD | A | 21. harmonic I1
Basic device measuring group 02 | 42602 | Float | RD | A | 22. harmonic I1
Basic device measuring group 02 | 42604 | Float | RD | A | 23. harmonic I1
Basic device measuring group 02 | 42606 | Float | RD | A | 24. harmonic I1
Basic device measuring group 02 | 42608 | Float | RD | A | 25. harmonic I1
Basic device measuring group 02 | 42610 | Float | RD | A | 26. harmonic I1
Basic device measuring group 02 | 42612 | Float | RD | A | 27. harmonic I1
Basic device measuring group 02 | 42614 | Float | RD | A | 28. harmonic I1
Basic device measuring group 02 | 42616 | Float | RD | A | 29. harmonic I1
Basic device measuring group 02 | 42618 | Float | RD | A | 30. harmonic I1
Basic device measuring group 02 | 42620 | Float | RD | A | 31. harmonic I1
Basic device measuring group 02 | 42622 | Float | RD | A | 32. harmonic I1
Basic device measuring group 02 | 42624 | Float | RD | A | 33. harmonic I1
Basic device measuring group 02 | 42626 | Float | RD | A | 34. harmonic I1
Basic device measuring group 02 | 42628 | Float | RD | A | 35. harmonic I1
Basic device measuring group 02 | 42630 | Float | RD | A | 36. harmonic I1
Basic device measuring group 02 | 42632 | Float | RD | A | 37. harmonic I1
Basic device measuring group 02 | 42634 | Float | RD | A | 38. harmonic I1
Basic device measuring group 02 | 42636 | Float | RD | A | 39. harmonic I1
Basic device measuring group 02 | 42638 | Float | RD | A | 40. harmonic I1
Basic device measuring group 02 | 42640 | Float | RD | A | 1. harmonic I2
Basic device measuring group 02 | 42642 | Float | RD | A | 2. harmonic I2
Basic device measuring group 02 | 42644 | Float | RD | A | 3. harmonic I2
Basic device measuring group 02 | 42646 | Float | RD | A | 4. harmonic I2
Basic device measuring group 02 | 42648 | Float | RD | A | 5. harmonic I2
Basic device measuring group 02 | 42650 | Float | RD | A | 6. harmonic I2
Basic device measuring group 02 | 42652 | Float | RD | A | 7. harmonic I2
Basic device measuring group 02 | 42654 | Float | RD | A | 8. harmonic I2
Basic device measuring group 02 | 42656 | Float | RD | A | 9. harmonic I2
Basic device measuring group 02 | 42658 | Float | RD | A | 10. harmonic I2
Basic device measuring group 02 | 42660 | Float | RD | A | 11. harmonic I2
Basic device measuring group 02 | 42662 | Float | RD | A | 12. harmonic I2
Basic device measuring group 02 | 42664 | Float | RD | A | 13. harmonic I2
Basic device measuring group 02 | 42666 | Float | RD | A | 14. harmonic I2
Basic device measuring group 02 | 42668 | Float | RD | A | 15. harmonic I2
Basic device measuring group 02 | 42670 | Float | RD | A | 16. harmonic I2
Basic device measuring group 02 | 42672 | Float | RD | A | 17. harmonic I2
Basic device measuring group 02 | 42674 | Float | RD | A | 18. harmonic I2
Basic device measuring group 02 | 42676 | Float | RD | A | 19. harmonic I2
Basic device measuring group 02 | 42678 | Float | RD | A | 20. harmonic I2
Basic device measuring group 02 | 42680 | Float | RD | A | 21. harmonic I2
Basic device measuring group 02 | 42682 | Float | RD | A | 22. harmonic I2
Basic device measuring group 02 | 42684 | Float | RD | A | 23. harmonic I2
Basic device measuring group 02 | 42686 | Float | RD | A | 24. harmonic I2
Basic device measuring group 02 | 42688 | Float | RD | A | 25. harmonic I2
Basic device measuring group 02 | 42690 | Float | RD | A | 26. harmonic I2
Basic device measuring group 02 | 42692 | Float | RD | A | 27. harmonic I2
Basic device measuring group 02 | 42694 | Float | RD | A | 28. harmonic I2
Basic device measuring group 02 | 42696 | Float | RD | A | 29. harmonic I2
Basic device measuring group 02 | 42698 | Float | RD | A | 30. harmonic I2
Basic device measuring group 02 | 42700 | Float | RD | A | 31. harmonic I2
Basic device measuring group 02 | 42702 | Float | RD | A | 32. harmonic I2
Basic device measuring group 02 | 42704 | Float | RD | A | 33. harmonic I2
Basic device measuring group 02 | 42706 | Float | RD | A | 34. harmonic I2
Basic device measuring group 02 | 42708 | Float | RD | A | 35. harmonic I2
Basic device measuring group 02 | 42710 | Float | RD | A | 36. harmonic I2
Basic device measuring group 02 | 42712 | Float | RD | A | 37. harmonic I2
Basic device measuring group 02 | 42714 | Float | RD | A | 38. harmonic I2
Basic device measuring group 02 | 42716 | Float | RD | A | 39. harmonic I2
Basic device measuring group 02 | 42718 | Float | RD | A | 40. harmonic I2
Basic device measuring group 02 | 42720 | Float | RD | A | 1. harmonic I3
Basic device measuring group 02 | 42722 | Float | RD | A | 2. harmonic I3
Basic device measuring group 02 | 42724 | Float | RD | A | 3. harmonic I3
Basic device measuring group 02 | 42726 | Float | RD | A | 4. harmonic I3
Basic device measuring group 02 | 42728 | Float | RD | A | 5. harmonic I3
Basic device measuring group 02 | 42730 | Float | RD | A | 6. harmonic I3
Basic device measuring group 02 | 42732 | Float | RD | A | 7. harmonic I3
Basic device measuring group 02 | 42734 | Float | RD | A | 8. harmonic I3
Basic device measuring group 02 | 42736 | Float | RD | A | 9. harmonic I3
Basic device measuring group 02 | 42738 | Float | RD | A | 10. harmonic I3
Basic device measuring group 02 | 42740 | Float | RD | A | 11. harmonic I3
Basic device measuring group 02 | 42742 | Float | RD | A | 12. harmonic I3
Basic device measuring group 02 | 42744 | Float | RD | A | 13. harmonic I3
Basic device measuring group 02 | 42746 | Float | RD | A | 14. harmonic I3
Basic device measuring group 02 | 42748 | Float | RD | A | 15. harmonic I3
Basic device measuring group 02 | 42750 | Float | RD | A | 16. harmonic I3
Basic device measuring group 02 | 42752 | Float | RD | A | 17. harmonic I3
Basic device measuring group 02 | 42754 | Float | RD | A | 18. harmonic I3
Basic device measuring group 02 | 42756 | Float | RD | A | 19. harmonic I3
Basic device measuring group 02 | 42758 | Float | RD | A | 20. harmonic I3
Basic device measuring group 02 | 42760 | Float | RD | A | 21. harmonic I3
Basic device measuring group 02 | 42762 | Float | RD | A | 22. harmonic I3
Basic device measuring group 02 | 42764 | Float | RD | A | 23. harmonic I3
Basic device measuring group 02 | 42766 | Float | RD | A | 24. harmonic I3
Basic device measuring group 02 | 42768 | Float | RD | A | 25. harmonic I3
Basic device measuring group 02 | 42770 | Float | RD | A | 26. harmonic I3
Basic device measuring group 02 | 42772 | Float | RD | A | 27. harmonic I3
Basic device measuring group 02 | 42774 | Float | RD | A | 28. harmonic I3
Basic device measuring group 02 | 42776 | Float | RD | A | 29. harmonic I3
Basic device measuring group 02 | 42778 | Float | RD | A | 30. harmonic I3
Basic device measuring group 02 | 42780 | Float | RD | A | 31. harmonic I3
Basic device measuring group 02 | 42782 | Float | RD | A | 32. harmonic I3
Basic device measuring group 02 | 42784 | Float | RD | A | 33. harmonic I3
Basic device measuring group 02 | 42786 | Float | RD | A | 34. harmonic I3
Basic device measuring group 02 | 42788 | Float | RD | A | 35. harmonic I3
Basic device measuring group 02 | 42790 | Float | RD | A | 36. harmonic I3
Basic device measuring group 02 | 42792 | Float | RD | A | 37. harmonic I3
Basic device measuring group 02 | 42794 | Float | RD | A | 38. harmonic I3
Basic device measuring group 02 | 42796 | Float | RD | A | 39. harmonic I3
Basic device measuring group 02 | 42798 | Float | RD | A | 40. harmonic I3
Basic device measuring group 02 | 42800 | Float | RD | A | 1. harmonic I4
Basic device measuring group 02 | 42802 | Float | RD | A | 2. harmonic I4
Basic device measuring group 02 | 42804 | Float | RD | A | 3. harmonic I4
Basic device measuring group 02 | 42806 | Float | RD | A | 4. harmonic I4
Basic device measuring group 02 | 42808 | Float | RD | A | 5. harmonic I4
Basic device measuring group 02 | 42810 | Float | RD | A | 6. harmonic I4
Basic device measuring group 02 | 42812 | Float | RD | A | 7. harmonic I4
Basic device measuring group 02 | 42814 | Float | RD | A | 8. harmonic I4
Basic device measuring group 02 | 42816 | Float | RD | A | 9. harmonic I4
Basic device measuring group 02 | 42818 | Float | RD | A | 10. harmonic I4
Basic device measuring group 02 | 42820 | Float | RD | A | 11. harmonic I4
Basic device measuring group 02 | 42822 | Float | RD | A | 12. harmonic I4
Basic device measuring group 02 | 42824 | Float | RD | A | 13. harmonic I4
Basic device measuring group 02 | 42826 | Float | RD | A | 14. harmonic I4
Basic device measuring group 02 | 42828 | Float | RD | A | 15. harmonic I4
Basic device measuring group 02 | 42830 | Float | RD | A | 16. harmonic I4
Basic device measuring group 02 | 42832 | Float | RD | A | 17. harmonic I4
Basic device measuring group 02 | 42834 | Float | RD | A | 18. harmonic I4
Basic device measuring group 02 | 42836 | Float | RD | A | 19. harmonic I4
Basic device measuring group 02 | 42838 | Float | RD | A | 20. harmonic I4
Basic device measuring group 02 | 42840 | Float | RD | A | 21. harmonic I4
Basic device measuring group 02 | 42842 | Float | RD | A | 22. harmonic I4
Basic device measuring group 02 | 42844 | Float | RD | A | 23. harmonic I4
Basic device measuring group 02 | 42846 | Float | RD | A | 24. harmonic I4
Basic device measuring group 02 | 42848 | Float | RD | A | 25. harmonic I4
Basic device measuring group 02 | 42850 | Float | RD | A | 26. harmonic I4
Basic device measuring group 02 | 42852 | Float | RD | A | 27. harmonic I4
Basic device measuring group 02 | 42854 | Float | RD | A | 28. harmonic I4
Basic device measuring group 02 | 42856 | Float | RD | A | 29. harmonic I4
Basic device measuring group 02 | 42858 | Float | RD | A | 30. harmonic I4
Basic device measuring group 02 | 42860 | Float | RD | A | 31. harmonic I4
Basic device measuring group 02 | 42862 | Float | RD | A | 32. harmonic I4
Basic device measuring group 02 | 42864 | Float | RD | A | 33. harmonic I4
Basic device measuring group 02 | 42866 | Float | RD | A | 34. harmonic I4
Basic device measuring group 02 | 42868 | Float | RD | A | 35. harmonic I4
Basic device measuring group 02 | 42870 | Float | RD | A | 36. harmonic I4
Basic device measuring group 02 | 42872 | Float | RD | A | 37. harmonic I4
Basic device measuring group 02 | 42874 | Float | RD | A | 38. harmonic I4
Basic device measuring group 02 | 42876 | Float | RD | A | 39. harmonic I4
Basic device measuring group 02 | 42878 | Float | RD | A | 40. harmonic I4
Basic device measuring group 03 | 42880 | Float | RD | A | 1. harmonic I1
Basic device measuring group 03 | 42882 | Float | RD | A | 2. harmonic I1
Basic device measuring group 03 | 42884 | Float | RD | A | 3. harmonic I1
Basic device measuring group 03 | 42886 | Float | RD | A | 4. harmonic I1
Basic device measuring group 03 | 42888 | Float | RD | A | 5. harmonic I1
Basic device measuring group 03 | 42890 | Float | RD | A | 6. harmonic I1
Basic device measuring group 03 | 42892 | Float | RD | A | 7. harmonic I1
Basic device measuring group 03 | 42894 | Float | RD | A | 8. harmonic I1
Basic device measuring group 03 | 42896 | Float | RD | A | 9. harmonic I1
Basic device measuring group 03 | 42898 | Float | RD | A | 10. harmonic I1
Basic device measuring group 03 | 42900 | Float | RD | A | 11. harmonic I1
Basic device measuring group 03 | 42902 | Float | RD | A | 12. harmonic I1
Basic device measuring group 03 | 42904 | Float | RD | A | 13. harmonic I1
Basic device measuring group 03 | 42906 | Float | RD | A | 14. harmonic I1
Basic device measuring group 03 | 42908 | Float | RD | A | 15. harmonic I1
Basic device measuring group 03 | 42910 | Float | RD | A | 16. harmonic I1
Basic device measuring group 03 | 42912 | Float | RD | A | 17. harmonic I1
Basic device measuring group 03 | 42914 | Float | RD | A | 18. harmonic I1
Basic device measuring group 03 | 42916 | Float | RD | A | 19. harmonic I1
Basic device measuring group 03 | 42918 | Float | RD | A | 20. harmonic I1
Basic device measuring group 03 | 42920 | Float | RD | A | 21. harmonic I1
Basic device measuring group 03 | 42922 | Float | RD | A | 22. harmonic I1
Basic device measuring group 03 | 42924 | Float | RD | A | 23. harmonic I1
Basic device measuring group 03 | 42926 | Float | RD | A | 24. harmonic I1
Basic device measuring group 03 | 42928 | Float | RD | A | 25. harmonic I1
Basic device measuring group 03 | 42930 | Float | RD | A | 26. harmonic I1
Basic device measuring group 03 | 42932 | Float | RD | A | 27. harmonic I1
Basic device measuring group 03 | 42934 | Float | RD | A | 28. harmonic I1
Basic device measuring group 03 | 42936 | Float | RD | A | 29. harmonic I1
Basic device measuring group 03 | 42938 | Float | RD | A | 30. harmonic I1
Basic device measuring group 03 | 42940 | Float | RD | A | 31. harmonic I1
Basic device measuring group 03 | 42942 | Float | RD | A | 32. harmonic I1
Basic device measuring group 03 | 42944 | Float | RD | A | 33. harmonic I1
Basic device measuring group 03 | 42946 | Float | RD | A | 34. harmonic I1
Basic device measuring group 03 | 42948 | Float | RD | A | 35. harmonic I1
Basic device measuring group 03 | 42950 | Float | RD | A | 36. harmonic I1
Basic device measuring group 03 | 42952 | Float | RD | A | 37. harmonic I1
Basic device measuring group 03 | 42954 | Float | RD | A | 38. harmonic I1
Basic device measuring group 03 | 42956 | Float | RD | A | 39. harmonic I1
Basic device measuring group 03 | 42958 | Float | RD | A | 40. harmonic I1
Basic device measuring group 03 | 42960 | Float | RD | A | 1. harmonic I2
Basic device measuring group 03 | 42962 | Float | RD | A | 2. harmonic I2
Basic device measuring group 03 | 42964 | Float | RD | A | 3. harmonic I2
Basic device measuring group 03 | 42966 | Float | RD | A | 4. harmonic I2
Basic device measuring group 03 | 42968 | Float | RD | A | 5. harmonic I2
Basic device measuring group 03 | 42970 | Float | RD | A | 6. harmonic I2
Basic device measuring group 03 | 42972 | Float | RD | A | 7. harmonic I2
Basic device measuring group 03 | 42974 | Float | RD | A | 8. harmonic I2
Basic device measuring group 03 | 42976 | Float | RD | A | 9. harmonic I2
Basic device measuring group 03 | 42978 | Float | RD | A | 10. harmonic I2
Basic device measuring group 03 | 42980 | Float | RD | A | 11. harmonic I2
Basic device measuring group 03 | 42982 | Float | RD | A | 12. harmonic I2
Basic device measuring group 03 | 42984 | Float | RD | A | 13. harmonic I2
Basic device measuring group 03 | 42986 | Float | RD | A | 14. harmonic I2
Basic device measuring group 03 | 42988 | Float | RD | A | 15. harmonic I2
Basic device measuring group 03 | 42990 | Float | RD | A | 16. harmonic I2
Basic device measuring group 03 | 42992 | Float | RD | A | 17. harmonic I2
Basic device measuring group 03 | 42994 | Float | RD | A | 18. harmonic I2
Basic device measuring group 03 | 42996 | Float | RD | A | 19. harmonic I2
Basic device measuring group 03 | 42998 | Float | RD | A | 20. harmonic I2
Basic device measuring group 03 | 43000 | Float | RD | A | 21. harmonic I2
Basic device measuring group 03 | 43002 | Float | RD | A | 22. harmonic I2
Basic device measuring group 03 | 43004 | Float | RD | A | 23. harmonic I2
Basic device measuring group 03 | 43006 | Float | RD | A | 24. harmonic I2
Basic device measuring group 03 | 43008 | Float | RD | A | 25. harmonic I2
Basic device measuring group 03 | 43010 | Float | RD | A | 26. harmonic I2
Basic device measuring group 03 | 43012 | Float | RD | A | 27. harmonic I2
Basic device measuring group 03 | 43014 | Float | RD | A | 28. harmonic I2
Basic device measuring group 03 | 43016 | Float | RD | A | 29. harmonic I2
Basic device measuring group 03 | 43018 | Float | RD | A | 30. harmonic I2
Basic device measuring group 03 | 43020 | Float | RD | A | 31. harmonic I2
Basic device measuring group 03 | 43022 | Float | RD | A | 32. harmonic I2
Basic device measuring group 03 | 43024 | Float | RD | A | 33. harmonic I2
Basic device measuring group 03 | 43026 | Float | RD | A | 34. harmonic I2
Basic device measuring group 03 | 43028 | Float | RD | A | 35. harmonic I2
Basic device measuring group 03 | 43030 | Float | RD | A | 36. harmonic I2
Basic device measuring group 03 | 43032 | Float | RD | A | 37. harmonic I2
Basic device measuring group 03 | 43034 | Float | RD | A | 38. harmonic I2
Basic device measuring group 03 | 43036 | Float | RD | A | 39. harmonic I2
Basic device measuring group 03 | 43038 | Float | RD | A | 40. harmonic I2
Basic device measuring group 03 | 43040 | Float | RD | A | 1. harmonic I3
Basic device measuring group 03 | 43042 | Float | RD | A | 2. harmonic I3
Basic device measuring group 03 | 43044 | Float | RD | A | 3. harmonic I3
Basic device measuring group 03 | 43046 | Float | RD | A | 4. harmonic I3
Basic device measuring group 03 | 43048 | Float | RD | A | 5. harmonic I3
Basic device measuring group 03 | 43050 | Float | RD | A | 6. harmonic I3
Basic device measuring group 03 | 43052 | Float | RD | A | 7. harmonic I3
Basic device measuring group 03 | 43054 | Float | RD | A | 8. harmonic I3
Basic device measuring group 03 | 43056 | Float | RD | A | 9. harmonic I3
Basic device measuring group 03 | 43058 | Float | RD | A | 10. harmonic I3
Basic device measuring group 03 | 43060 | Float | RD | A | 11. harmonic I3
Basic device measuring group 03 | 43062 | Float | RD | A | 12. harmonic I3
Basic device measuring group 03 | 43064 | Float | RD | A | 13. harmonic I3
Basic device measuring group 03 | 43066 | Float | RD | A | 14. harmonic I3
Basic device measuring group 03 | 43068 | Float | RD | A | 15. harmonic I3
Basic device measuring group 03 | 43070 | Float | RD | A | 16. harmonic I3
Basic device measuring group 03 | 43072 | Float | RD | A | 17. harmonic I3
Basic device measuring group 03 | 43074 | Float | RD | A | 18. harmonic I3
Basic device measuring group 03 | 43076 | Float | RD | A | 19. harmonic I3
Basic device measuring group 03 | 43078 | Float | RD | A | 20. harmonic I3
Basic device measuring group 03 | 43080 | Float | RD | A | 21. harmonic I3
Basic device measuring group 03 | 43082 | Float | RD | A | 22. harmonic I3
Basic device measuring group 03 | 43084 | Float | RD | A | 23. harmonic I3
Basic device measuring group 03 | 43086 | Float | RD | A | 24. harmonic I3
Basic device measuring group 03 | 43088 | Float | RD | A | 25. harmonic I3
Basic device measuring group 03 | 43090 | Float | RD | A | 26. harmonic I3
Basic device measuring group 03 | 43092 | Float | RD | A | 27. harmonic I3
Basic device measuring group 03 | 43094 | Float | RD | A | 28. harmonic I3
Basic device measuring group 03 | 43096 | Float | RD | A | 29. harmonic I3
Basic device measuring group 03 | 43098 | Float | RD | A | 30. harmonic I3
Basic device measuring group 03 | 43100 | Float | RD | A | 31. harmonic I3
Basic device measuring group 03 | 43102 | Float | RD | A | 32. harmonic I3
Basic device measuring group 03 | 43104 | Float | RD | A | 33. harmonic I3
Basic device measuring group 03 | 43106 | Float | RD | A | 34. harmonic I3
Basic device measuring group 03 | 43108 | Float | RD | A | 35. harmonic I3
Basic device measuring group 03 | 43110 | Float | RD | A | 36. harmonic I3
Basic device measuring group 03 | 43112 | Float | RD | A | 37. harmonic I3
Basic device measuring group 03 | 43114 | Float | RD | A | 38. harmonic I3
Basic device measuring group 03 | 43116 | Float | RD | A | 39. harmonic I3
Basic device measuring group 03 | 43118 | Float | RD | A | 40. harmonic I3
Basic device measuring group 03 | 43120 | Float | RD | A | 1. harmonic I4
Basic device measuring group 03 | 43122 | Float | RD | A | 2. harmonic I4
Basic device measuring group 03 | 43124 | Float | RD | A | 3. harmonic I4
Basic device measuring group 03 | 43126 | Float | RD | A | 4. harmonic I4
Basic device measuring group 03 | 43128 | Float | RD | A | 5. harmonic I4
Basic device measuring group 03 | 43130 | Float | RD | A | 6. harmonic I4
Basic device measuring group 03 | 43132 | Float | RD | A | 7. harmonic I4
Basic device measuring group 03 | 43134 | Float | RD | A | 8. harmonic I4
Basic device measuring group 03 | 43136 | Float | RD | A | 9. harmonic I4
Basic device measuring group 03 | 43138 | Float | RD | A | 10. harmonic I4
Basic device measuring group 03 | 43140 | Float | RD | A | 11. harmonic I4
Basic device measuring group 03 | 43142 | Float | RD | A | 12. harmonic I4
Basic device measuring group 03 | 43144 | Float | RD | A | 13. harmonic I4
Basic device measuring group 03 | 43146 | Float | RD | A | 14. harmonic I4
Basic device measuring group 03 | 43148 | Float | RD | A | 15. harmonic I4
Basic device measuring group 03 | 43150 | Float | RD | A | 16. harmonic I4
Basic device measuring group 03 | 43152 | Float | RD | A | 17. harmonic I4
Basic device measuring group 03 | 43154 | Float | RD | A | 18. harmonic I4
Basic device measuring group 03 | 43156 | Float | RD | A | 19. harmonic I4
Basic device measuring group 03 | 43158 | Float | RD | A | 20. harmonic I4
Basic device measuring group 03 | 43160 | Float | RD | A | 21. harmonic I4
Basic device measuring group 03 | 43162 | Float | RD | A | 22. harmonic I4
Basic device measuring group 03 | 43164 | Float | RD | A | 23. harmonic I4
Basic device measuring group 03 | 43166 | Float | RD | A | 24. harmonic I4
Basic device measuring group 03 | 43168 | Float | RD | A | 25. harmonic I4
Basic device measuring group 03 | 43170 | Float | RD | A | 26. harmonic I4
Basic device measuring group 03 | 43172 | Float | RD | A | 27. harmonic I4
Basic device measuring group 03 | 43174 | Float | RD | A | 28. harmonic I4
Basic device measuring group 03 | 43176 | Float | RD | A | 29. harmonic I4
Basic device measuring group 03 | 43178 | Float | RD | A | 30. harmonic I4
Basic device measuring group 03 | 43180 | Float | RD | A | 31. harmonic I4
Basic device measuring group 03 | 43182 | Float | RD | A | 32. harmonic I4
Basic device measuring group 03 | 43184 | Float | RD | A | 33. harmonic I4
Basic device measuring group 03 | 43186 | Float | RD | A | 34. harmonic I4
Basic device measuring group 03 | 43188 | Float | RD | A | 35. harmonic I4
Basic device measuring group 03 | 43190 | Float | RD | A | 36. harmonic I4
Basic device measuring group 03 | 43192 | Float | RD | A | 37. harmonic I4
Basic device measuring group 03 | 43194 | Float | RD | A | 38. harmonic I4
Basic device measuring group 03 | 43196 | Float | RD | A | 39. harmonic I4
Basic device measuring group 03 | 43198 | Float | RD | A | 40. harmonic I4
Current measurement module group 01 | 43200 | Float | RD | A | 1. harmonic I1
Current measurement module group 01 | 43202 | Float | RD | A | 3. harmonic I1
Current measurement module group 01 | 43204 | Float | RD | A | 5. harmonic I1
Current measurement module group 01 | 43206 | Float | RD | A | 7. harmonic I1
Current measurement module group 01 | 43208 | Float | RD | A | 9. harmonic I1
Current measurement module group 01 | 43210 | Float | RD | A | 11. harmonic I1
Current measurement module group 01 | 43212 | Float | RD | A | 13. harmonic I1
Current measurement module group 01 | 43214 | Float | RD | A | 15. harmonic I1
Current measurement module group 01 | 43216 | Float | RD | A | 17. harmonic I1
Current measurement module group 01 | 43218 | Float | RD | A | 19. harmonic I1
Current measurement module group 01 | 43220 | Float | RD | A | 21. harmonic I1
Current measurement module group 01 | 43222 | Float | RD | A | 23. harmonic I1
Current measurement module group 01 | 43224 | Float | RD | A | 25. harmonic I1
Current measurement module group 01 | 43226 | Float | RD | A | 1. harmonic I2
Current measurement module group 01 | 43228 | Float | RD | A | 3. harmonic I2
Current measurement module group 01 | 43230 | Float | RD | A | 5. harmonic I2
Current measurement module group 01 | 43232 | Float | RD | A | 7. harmonic I2
Current measurement module group 01 | 43234 | Float | RD | A | 9. harmonic I2
Current measurement module group 01 | 43236 | Float | RD | A | 11. harmonic I2
Current measurement module group 01 | 43238 | Float | RD | A | 13. harmonic I2
Current measurement module group 01 | 43240 | Float | RD | A | 15. harmonic I2
Current measurement module group 01 | 43242 | Float | RD | A | 17. harmonic I2
Current measurement module group 01 | 43244 | Float | RD | A | 19. harmonic I2
Current measurement module group 01 | 43246 | Float | RD | A | 21. harmonic I2
Current measurement module group 01 | 43248 | Float | RD | A | 23. harmonic I2
Current measurement module group 01 | 43250 | Float | RD | A | 25. harmonic I2
Current measurement module group 01 | 43252 | Float | RD | A | 1. harmonic I3
Current measurement module group 01 | 43254 | Float | RD | A | 3. harmonic I3
Current measurement module group 01 | 43256 | Float | RD | A | 5. harmonic I3
Current measurement module group 01 | 43258 | Float | RD | A | 7. harmonic I3
Current measurement module group 01 | 43260 | Float | RD | A | 9. harmonic I3
Current measurement module group 01 | 43262 | Float | RD | A | 11. harmonic I3
Current measurement module group 01 | 43264 | Float | RD | A | 13. harmonic I3
Current measurement module group 01 | 43266 | Float | RD | A | 15. harmonic I3
Current measurement module group 01 | 43268 | Float | RD | A | 17. harmonic I3
Current measurement module group 01 | 43270 | Float | RD | A | 19. harmonic I3
Current measurement module group 01 | 43272 | Float | RD | A | 21. harmonic I3
Current measurement module group 01 | 43274 | Float | RD | A | 23. harmonic I3
Current measurement module group 01 | 43276 | Float | RD | A | 25. harmonic I3
Current measurement module group 01 | 43278 | Float | RD | A | 1. harmonic I4
Current measurement module group 01 | 43280 | Float | RD | A | 3. harmonic I4
Current measurement module group 01 | 43282 | Float | RD | A | 5. harmonic I4
Current measurement module group 01 | 43284 | Float | RD | A | 7. harmonic I4
Current measurement module group 01 | 43286 | Float | RD | A | 9. harmonic I4
Current measurement module group 01 | 43288 | Float | RD | A | 11. harmonic I4
Current measurement module group 01 | 43290 | Float | RD | A | 13. harmonic I4
Current measurement module group 01 | 43292 | Float | RD | A | 15. harmonic I4
Current measurement module group 01 | 43294 | Float | RD | A | 17. harmonic I4
Current measurement module group 01 | 43296 | Float | RD | A | 19. harmonic I4
Current measurement module group 01 | 43298 | Float | RD | A | 21. harmonic I4
Current measurement module group 01 | 43300 | Float | RD | A | 23. harmonic I4
Current measurement module group 01 | 43302 | Float | RD | A | 25. harmonic I4
Current measurement module group 02 | 43304 | Float | RD | A | 1. harmonic I1
Current measurement module group 02 | 43306 | Float | RD | A | 3. harmonic I1
Current measurement module group 02 | 43308 | Float | RD | A | 5. harmonic I1
Current measurement module group 02 | 43310 | Float | RD | A | 7. harmonic I1
Current measurement module group 02 | 43312 | Float | RD | A | 9. harmonic I1
Current measurement module group 02 | 43314 | Float | RD | A | 11. harmonic I1
Current measurement module group 02 | 43316 | Float | RD | A | 13. harmonic I1
Current measurement module group 02 | 43318 | Float | RD | A | 15. harmonic I1
Current measurement module group 02 | 43320 | Float | RD | A | 17. harmonic I1
Current measurement module group 02 | 43322 | Float | RD | A | 19. harmonic I1
Current measurement module group 02 | 43324 | Float | RD | A | 21. harmonic I1
Current measurement module group 02 | 43326 | Float | RD | A | 23. harmonic I1
Current measurement module group 02 | 43328 | Float | RD | A | 25. harmonic I1
Current measurement module group 02 | 43330 | Float | RD | A | 1. harmonic I2
Current measurement module group 02 | 43332 | Float | RD | A | 3. harmonic I2
Current measurement module group 02 | 43334 | Float | RD | A | 5. harmonic I2
Current measurement module group 02 | 43336 | Float | RD | A | 7. harmonic I2
Current measurement module group 02 | 43338 | Float | RD | A | 9. harmonic I2
Current measurement module group 02 | 43340 | Float | RD | A | 11. harmonic I2
Current measurement module group 02 | 43342 | Float | RD | A | 13. harmonic I2
Current measurement module group 02 | 43344 | Float | RD | A | 15. harmonic I2
Current measurement module group 02 | 43346 | Float | RD | A | 17. harmonic I2
Current measurement module group 02 | 43348 | Float | RD | A | 19. harmonic I2
Current measurement module group 02 | 43350 | Float | RD | A | 21. harmonic I2
Current measurement module group 02 | 43352 | Float | RD | A | 23. harmonic I2
Current measurement module group 02 | 43354 | Float | RD | A | 25. harmonic I2
Current measurement module group 02 | 43356 | Float | RD | A | 1. harmonic I3
Current measurement module group 02 | 43358 | Float | RD | A | 3. harmonic I3
Current measurement module group 02 | 43360 | Float | RD | A | 5. harmonic I3
Current measurement module group 02 | 43362 | Float | RD | A | 7. harmonic I3
Current measurement module group 02 | 43364 | Float | RD | A | 9. harmonic I3
Current measurement module group 02 | 43366 | Float | RD | A | 11. harmonic I3
Current measurement module group 02 | 43368 | Float | RD | A | 13. harmonic I3
Current measurement module group 02 | 43370 | Float | RD | A | 15. harmonic I3
Current measurement module group 02 | 43372 | Float | RD | A | 17. harmonic I3
Current measurement module group 02 | 43374 | Float | RD | A | 19. harmonic I3
Current measurement module group 02 | 43376 | Float | RD | A | 21. harmonic I3
Current measurement module group 02 | 43378 | Float | RD | A | 23. harmonic I3
Current measurement module group 02 | 43380 | Float | RD | A | 25. harmonic I3
Current measurement module group 02 | 43382 | Float | RD | A | 1. harmonic I4
Current measurement module group 02 | 43384 | Float | RD | A | 3. harmonic I4
Current measurement module group 02 | 43386 | Float | RD | A | 5. harmonic I4
Current measurement module group 02 | 43388 | Float | RD | A | 7. harmonic I4
Current measurement module group 02 | 43390 | Float | RD | A | 9. harmonic I4
Current measurement module group 02 | 43392 | Float | RD | A | 11. harmonic I4
Current measurement module group 02 | 43394 | Float | RD | A | 13. harmonic I4
Current measurement module group 02 | 43396 | Float | RD | A | 15. harmonic I4
Current measurement module group 02 | 43398 | Float | RD | A | 17. harmonic I4
Current measurement module group 02 | 43400 | Float | RD | A | 19. harmonic I4
Current measurement module group 02 | 43402 | Float | RD | A | 21. harmonic I4
Current measurement module group 02 | 43404 | Float | RD | A | 23. harmonic I4
Current measurement module group 02 | 43406 | Float | RD | A | 25. harmonic I4
Current measurement module group 03 | 43408 | Float | RD | A | 1. harmonic I1
Current measurement module group 03 | 43410 | Float | RD | A | 3. harmonic I1
Current measurement module group 03 | 43412 | Float | RD | A | 5. harmonic I1
Current measurement module group 03 | 43414 | Float | RD | A | 7. harmonic I1
Current measurement module group 03 | 43416 | Float | RD | A | 9. harmonic I1
Current measurement module group 03 | 43418 | Float | RD | A | 11. harmonic I1
Current measurement module group 03 | 43420 | Float | RD | A | 13. harmonic I1
Current measurement module group 03 | 43422 | Float | RD | A | 15. harmonic I1
Current measurement module group 03 | 43424 | Float | RD | A | 17. harmonic I1
Current measurement module group 03 | 43426 | Float | RD | A | 19. harmonic I1
Current measurement module group 03 | 43428 | Float | RD | A | 21. harmonic I1
Current measurement module group 03 | 43430 | Float | RD | A | 23. harmonic I1
Current measurement module group 03 | 43432 | Float | RD | A | 25. harmonic I1
Current measurement module group 03 | 43434 | Float | RD | A | 1. harmonic I2
Current measurement module group 03 | 43436 | Float | RD | A | 3. harmonic I2
Current measurement module group 03 | 43438 | Float | RD | A | 5. harmonic I2
Current measurement module group 03 | 43440 | Float | RD | A | 7. harmonic I2
Current measurement module group 03 | 43442 | Float | RD | A | 9. harmonic I2
Current measurement module group 03 | 43444 | Float | RD | A | 11. harmonic I2
Current measurement module group 03 | 43446 | Float | RD | A | 13. harmonic I2
Current measurement module group 03 | 43448 | Float | RD | A | 15. harmonic I2
Current measurement module group 03 | 43450 | Float | RD | A | 17. harmonic I2
Current measurement module group 03 | 43452 | Float | RD | A | 19. harmonic I2
Current measurement module group 03 | 43454 | Float | RD | A | 21. harmonic I2
Current measurement module group 03 | 43456 | Float | RD | A | 23. harmonic I2
Current measurement module group 03 | 43458 | Float | RD | A | 25. harmonic I2
Current measurement module group 03 | 43460 | Float | RD | A | 1. harmonic I3
Current measurement module group 03 | 43462 | Float | RD | A | 3. harmonic I3
Current measurement module group 03 | 43464 | Float | RD | A | 5. harmonic I3
Current measurement module group 03 | 43466 | Float | RD | A | 7. harmonic I3
Current measurement module group 03 | 43468 | Float | RD | A | 9. harmonic I3
Current measurement module group 03 | 43470 | Float | RD | A | 11. harmonic I3
Current measurement module group 03 | 43472 | Float | RD | A | 13. harmonic I3
Current measurement module group 03 | 43474 | Float | RD | A | 15. harmonic I3
Current measurement module group 03 | 43476 | Float | RD | A | 17. harmonic I3
Current measurement module group 03 | 43478 | Float | RD | A | 19. harmonic I3
Current measurement module group 03 | 43480 | Float | RD | A | 21. harmonic I3
Current measurement module group 03 | 43482 | Float | RD | A | 23. harmonic I3
Current measurement module group 03 | 43484 | Float | RD | A | 25. harmonic I3
Current measurement module group 03 | 43486 | Float | RD | A | 1. harmonic I4
Current measurement module group 03 | 43488 | Float | RD | A | 3. harmonic I4
Current measurement module group 03 | 43490 | Float | RD | A | 5. harmonic I4
Current measurement module group 03 | 43492 | Float | RD | A | 7. harmonic I4
Current measurement module group 03 | 43494 | Float | RD | A | 9. harmonic I4
Current measurement module group 03 | 43496 | Float | RD | A | 11. harmonic I4
Current measurement module group 03 | 43498 | Float | RD | A | 13. harmonic I4
Current measurement module group 03 | 43500 | Float | RD | A | 15. harmonic I4
Current measurement module group 03 | 43502 | Float | RD | A | 17. harmonic I4
Current measurement module group 03 | 43504 | Float | RD | A | 19. harmonic I4
Current measurement module group 03 | 43506 | Float | RD | A | 21. harmonic I4
Current measurement module group 03 | 43508 | Float | RD | A | 23. harmonic I4
Current measurement module group 03 | 43510 | Float | RD | A | 25. harmonic I4
Current measurement module group 04 | 43512 | Float | RD | A | 1. harmonic I1
Current measurement module group 04 | 43514 | Float | RD | A | 3. harmonic I1
Current measurement module group 04 | 43516 | Float | RD | A | 5. harmonic I1
Current measurement module group 04 | 43518 | Float | RD | A | 7. harmonic I1
Current measurement module group 04 | 43520 | Float | RD | A | 9. harmonic I1
Current measurement module group 04 | 43522 | Float | RD | A | 11. harmonic I1
Current measurement module group 04 | 43524 | Float | RD | A | 13. harmonic I1
Current measurement module group 04 | 43526 | Float | RD | A | 15. harmonic I1
Current measurement module group 04 | 43528 | Float | RD | A | 17. harmonic I1
Current measurement module group 04 | 43530 | Float | RD | A | 19. harmonic I1
Current measurement module group 04 | 43532 | Float | RD | A | 21. harmonic I1
Current measurement module group 04 | 43534 | Float | RD | A | 23. harmonic I1
Current measurement module group 04 | 43536 | Float | RD | A | 25. harmonic I1
Current measurement module group 04 | 43538 | Float | RD | A | 1. harmonic I2
Current measurement module group 04 | 43540 | Float | RD | A | 3. harmonic I2
Current measurement module group 04 | 43542 | Float | RD | A | 5. harmonic I2
Current measurement module group 04 | 43544 | Float | RD | A | 7. harmonic I2
Current measurement module group 04 | 43546 | Float | RD | A | 9. harmonic I2
Current measurement module group 04 | 43548 | Float | RD | A | 11. harmonic I2
Current measurement module group 04 | 43550 | Float | RD | A | 13. harmonic I2
Current measurement module group 04 | 43552 | Float | RD | A | 15. harmonic I2
Current measurement module group 04 | 43554 | Float | RD | A | 17. harmonic I2
Current measurement module group 04 | 43556 | Float | RD | A | 19. harmonic I2
Current measurement module group 04 | 43558 | Float | RD | A | 21. harmonic I2
Current measurement module group 04 | 43560 | Float | RD | A | 23. harmonic I2
Current measurement module group 04 | 43562 | Float | RD | A | 25. harmonic I2
Current measurement module group 04 | 43564 | Float | RD | A | 1. harmonic I3
Current measurement module group 04 | 43566 | Float | RD | A | 3. harmonic I3
Current measurement module group 04 | 43568 | Float | RD | A | 5. harmonic I3
Current measurement module group 04 | 43570 | Float | RD | A | 7. harmonic I3
Current measurement module group 04 | 43572 | Float | RD | A | 9. harmonic I3
Current measurement module group 04 | 43574 | Float | RD | A | 11. harmonic I3
Current measurement module group 04 | 43576 | Float | RD | A | 13. harmonic I3
Current measurement module group 04 | 43578 | Float | RD | A | 15. harmonic I3
Current measurement module group 04 | 43580 | Float | RD | A | 17. harmonic I3
Current measurement module group 04 | 43582 | Float | RD | A | 19. harmonic I3
Current measurement module group 04 | 43584 | Float | RD | A | 21. harmonic I3
Current measurement module group 04 | 43586 | Float | RD | A | 23. harmonic I3
Current measurement module group 04 | 43588 | Float | RD | A | 25. harmonic I3
Current measurement module group 04 | 43590 | Float | RD | A | 1. harmonic I4
Current measurement module group 04 | 43592 | Float | RD | A | 3. harmonic I4
Current measurement module group 04 | 43594 | Float | RD | A | 5. harmonic I4
Current measurement module group 04 | 43596 | Float | RD | A | 7. harmonic I4
Current measurement module group 04 | 43598 | Float | RD | A | 9. harmonic I4
Current measurement module group 04 | 43600 | Float | RD | A | 11. harmonic I4
Current measurement module group 04 | 43602 | Float | RD | A | 13. harmonic I4
Current measurement module group 04 | 43604 | Float | RD | A | 15. harmonic I4
Current measurement module group 04 | 43606 | Float | RD | A | 17. harmonic I4
Current measurement module group 04 | 43608 | Float | RD | A | 19. harmonic I4
Current measurement module group 04 | 43610 | Float | RD | A | 21. harmonic I4
Current measurement module group 04 | 43612 | Float | RD | A | 23. harmonic I4
Current measurement module group 04 | 43614 | Float | RD | A | 25. harmonic I4
Current measurement module group 05 | 43616 | Float | RD | A | 1. harmonic I1
Current measurement module group 05 | 43618 | Float | RD | A | 3. harmonic I1
Current measurement module group 05 | 43620 | Float | RD | A | 5. harmonic I1
Current measurement module group 05 | 43622 | Float | RD | A | 7. harmonic I1
Current measurement module group 05 | 43624 | Float | RD | A | 9. harmonic I1
Current measurement module group 05 | 43626 | Float | RD | A | 11. harmonic I1
Current measurement module group 05 | 43628 | Float | RD | A | 13. harmonic I1
Current measurement module group 05 | 43630 | Float | RD | A | 15. harmonic I1
Current measurement module group 05 | 43632 | Float | RD | A | 17. harmonic I1
Current measurement module group 05 | 43634 | Float | RD | A | 19. harmonic I1
Current measurement module group 05 | 43636 | Float | RD | A | 21. harmonic I1
Current measurement module group 05 | 43638 | Float | RD | A | 23. harmonic I1
Current measurement module group 05 | 43640 | Float | RD | A | 25. harmonic I1
Current measurement module group 05 | 43642 | Float | RD | A | 1. harmonic I2
Current measurement module group 05 | 43644 | Float | RD | A | 3. harmonic I2
Current measurement module group 05 | 43646 | Float | RD | A | 5. harmonic I2
Current measurement module group 05 | 43648 | Float | RD | A | 7. harmonic I2
Current measurement module group 05 | 43650 | Float | RD | A | 9. harmonic I2
Current measurement module group 05 | 43652 | Float | RD | A | 11. harmonic I2
Current measurement module group 05 | 43654 | Float | RD | A | 13. harmonic I2
Current measurement module group 05 | 43656 | Float | RD | A | 15. harmonic I2
Current measurement module group 05 | 43658 | Float | RD | A | 17. harmonic I2
Current measurement module group 05 | 43660 | Float | RD | A | 19. harmonic I2
Current measurement module group 05 | 43662 | Float | RD | A | 21. harmonic I2
Current measurement module group 05 | 43664 | Float | RD | A | 23. harmonic I2
Current measurement module group 05 | 43666 | Float | RD | A | 25. harmonic I2
Current measurement module group 05 | 43668 | Float | RD | A | 1. harmonic I3
Current measurement module group 05 | 43670 | Float | RD | A | 3. harmonic I3
Current measurement module group 05 | 43672 | Float | RD | A | 5. harmonic I3
Current measurement module group 05 | 43674 | Float | RD | A | 7. harmonic I3
Current measurement module group 05 | 43676 | Float | RD | A | 9. harmonic I3
Current measurement module group 05 | 43678 | Float | RD | A | 11. harmonic I3
Current measurement module group 05 | 43680 | Float | RD | A | 13. harmonic I3
Current measurement module group 05 | 43682 | Float | RD | A | 15. harmonic I3
Current measurement module group 05 | 43684 | Float | RD | A | 17. harmonic I3
Current measurement module group 05 | 43686 | Float | RD | A | 19. harmonic I3
Current measurement module group 05 | 43688 | Float | RD | A | 21. harmonic I3
Current measurement module group 05 | 43690 | Float | RD | A | 23. harmonic I3
Current measurement module group 05 | 43692 | Float | RD | A | 25. harmonic I3
Current measurement module group 05 | 43694 | Float | RD | A | 1. harmonic I4
Current measurement module group 05 | 43696 | Float | RD | A | 3. harmonic I4
Current measurement module group 05 | 43698 | Float | RD | A | 5. harmonic I4
Current measurement module group 05 | 43700 | Float | RD | A | 7. harmonic I4
Current measurement module group 05 | 43702 | Float | RD | A | 9. harmonic I4
Current measurement module group 05 | 43704 | Float | RD | A | 11. harmonic I4
Current measurement module group 05 | 43706 | Float | RD | A | 13. harmonic I4
Current measurement module group 05 | 43708 | Float | RD | A | 15. harmonic I4
Current measurement module group 05 | 43710 | Float | RD | A | 17. harmonic I4
Current measurement module group 05 | 43712 | Float | RD | A | 19. harmonic I4
Current measurement module group 05 | 43714 | Float | RD | A | 21. harmonic I4
Current measurement module group 05 | 43716 | Float | RD | A | 23. harmonic I4
Current measurement module group 05 | 43718 | Float | RD | A | 25. harmonic I4
Current measurement module group 06 | 43720 | Float | RD | A | 1. harmonic I1
Current measurement module group 06 | 43722 | Float | RD | A | 3. harmonic I1
Current measurement module group 06 | 43724 | Float | RD | A | 5. harmonic I1
Current measurement module group 06 | 43726 | Float | RD | A | 7. harmonic I1
Current measurement module group 06 | 43728 | Float | RD | A | 9. harmonic I1
Current measurement module group 06 | 43730 | Float | RD | A | 11. harmonic I1
Current measurement module group 06 | 43732 | Float | RD | A | 13. harmonic I1
Current measurement module group 06 | 43734 | Float | RD | A | 15. harmonic I1
Current measurement module group 06 | 43736 | Float | RD | A | 17. harmonic I1
Current measurement module group 06 | 43738 | Float | RD | A | 19. harmonic I1
Current measurement module group 06 | 43740 | Float | RD | A | 21. harmonic I1
Current measurement module group 06 | 43742 | Float | RD | A | 23. harmonic I1
Current measurement module group 06 | 43744 | Float | RD | A | 25. harmonic I1
Current measurement module group 06 | 43746 | Float | RD | A | 1. harmonic I2
Current measurement module group 06 | 43748 | Float | RD | A | 3. harmonic I2
Current measurement module group 06 | 43750 | Float | RD | A | 5. harmonic I2
Current measurement module group 06 | 43752 | Float | RD | A | 7. harmonic I2
Current measurement module group 06 | 43754 | Float | RD | A | 9. harmonic I2
Current measurement module group 06 | 43756 | Float | RD | A | 11. harmonic I2
Current measurement module group 06 | 43758 | Float | RD | A | 13. harmonic I2
Current measurement module group 06 | 43760 | Float | RD | A | 15. harmonic I2
Current measurement module group 06 | 43762 | Float | RD | A | 17. harmonic I2
Current measurement module group 06 | 43764 | Float | RD | A | 19. harmonic I2
Current measurement module group 06 | 43766 | Float | RD | A | 21. harmonic I2
Current measurement module group 06 | 43768 | Float | RD | A | 23. harmonic I2
Current measurement module group 06 | 43770 | Float | RD | A | 25. harmonic I2
Current measurement module group 06 | 43772 | Float | RD | A | 1. harmonic I3
Current measurement module group 06 | 43774 | Float | RD | A | 3. harmonic I3
Current measurement module group 06 | 43776 | Float | RD | A | 5. harmonic I3
Current measurement module group 06 | 43778 | Float | RD | A | 7. harmonic I3
Current measurement module group 06 | 43780 | Float | RD | A | 9. harmonic I3
Current measurement module group 06 | 43782 | Float | RD | A | 11. harmonic I3
Current measurement module group 06 | 43784 | Float | RD | A | 13. harmonic I3
Current measurement module group 06 | 43786 | Float | RD | A | 15. harmonic I3
Current measurement module group 06 | 43788 | Float | RD | A | 17. harmonic I3
Current measurement module group 06 | 43790 | Float | RD | A | 19. harmonic I3
Current measurement module group 06 | 43792 | Float | RD | A | 21. harmonic I3
Current measurement module group 06 | 43794 | Float | RD | A | 23. harmonic I3
Current measurement module group 06 | 43796 | Float | RD | A | 25. harmonic I3
Current measurement module group 06 | 43798 | Float | RD | A | 1. harmonic I4
Current measurement module group 06 | 43800 | Float | RD | A | 3. harmonic I4
Current measurement module group 06 | 43802 | Float | RD | A | 5. harmonic I4
Current measurement module group 06 | 43804 | Float | RD | A | 7. harmonic I4
Current measurement module group 06 | 43806 | Float | RD | A | 9. harmonic I4
Current measurement module group 06 | 43808 | Float | RD | A | 11. harmonic I4
Current measurement module group 06 | 43810 | Float | RD | A | 13. harmonic I4
Current measurement module group 06 | 43812 | Float | RD | A | 15. harmonic I4
Current measurement module group 06 | 43814 | Float | RD | A | 17. harmonic I4
Current measurement module group 06 | 43816 | Float | RD | A | 19. harmonic I4
Current measurement module group 06 | 43818 | Float | RD | A | 21. harmonic I4
Current measurement module group 06 | 43820 | Float | RD | A | 23. harmonic I4
Current measurement module group 06 | 43822 | Float | RD | A | 25. harmonic I4
Current measurement module group 07 | 43824 | Float | RD | A | 1. harmonic I1
Current measurement module group 07 | 43826 | Float | RD | A | 3. harmonic I1
Current measurement module group 07 | 43828 | Float | RD | A | 5. harmonic I1
Current measurement module group 07 | 43830 | Float | RD | A | 7. harmonic I1
Current measurement module group 07 | 43832 | Float | RD | A | 9. harmonic I1
Current measurement module group 07 | 43834 | Float | RD | A | 11. harmonic I1
Current measurement module group 07 | 43836 | Float | RD | A | 13. harmonic I1
Current measurement module group 07 | 43838 | Float | RD | A | 15. harmonic I1
Current measurement module group 07 | 43840 | Float | RD | A | 17. harmonic I1
Current measurement module group 07 | 43842 | Float | RD | A | 19. harmonic I1
Current measurement module group 07 | 43844 | Float | RD | A | 21. harmonic I1
Current measurement module group 07 | 43846 | Float | RD | A | 23. harmonic I1
Current measurement module group 07 | 43848 | Float | RD | A | 25. harmonic I1
Current measurement module group 07 | 43850 | Float | RD | A | 1. harmonic I2
Current measurement module group 07 | 43852 | Float | RD | A | 3. harmonic I2
Current measurement module group 07 | 43854 | Float | RD | A | 5. harmonic I2
Current measurement module group 07 | 43856 | Float | RD | A | 7. harmonic I2
Current measurement module group 07 | 43858 | Float | RD | A | 9. harmonic I2
Current measurement module group 07 | 43860 | Float | RD | A | 11. harmonic I2
Current measurement module group 07 | 43862 | Float | RD | A | 13. harmonic I2
Current measurement module group 07 | 43864 | Float | RD | A | 15. harmonic I2
Current measurement module group 07 | 43866 | Float | RD | A | 17. harmonic I2
Current measurement module group 07 | 43868 | Float | RD | A | 19. harmonic I2
Current measurement module group 07 | 43870 | Float | RD | A | 21. harmonic I2
Current measurement module group 07 | 43872 | Float | RD | A | 23. harmonic I2
Current measurement module group 07 | 43874 | Float | RD | A | 25. harmonic I2
Current measurement module group 07 | 43876 | Float | RD | A | 1. harmonic I3
Current measurement module group 07 | 43878 | Float | RD | A | 3. harmonic I3
Current measurement module group 07 | 43880 | Float | RD | A | 5. harmonic I3
Current measurement module group 07 | 43882 | Float | RD | A | 7. harmonic I3
Current measurement module group 07 | 43884 | Float | RD | A | 9. harmonic I3
Current measurement module group 07 | 43886 | Float | RD | A | 11. harmonic I3
Current measurement module group 07 | 43888 | Float | RD | A | 13. harmonic I3
Current measurement module group 07 | 43890 | Float | RD | A | 15. harmonic I3
Current measurement module group 07 | 43892 | Float | RD | A | 17. harmonic I3
Current measurement module group 07 | 43894 | Float | RD | A | 19. harmonic I3
Current measurement module group 07 | 43896 | Float | RD | A | 21. harmonic I3
Current measurement module group 07 | 43898 | Float | RD | A | 23. harmonic I3
Current measurement module group 07 | 43900 | Float | RD | A | 25. harmonic I3
Current measurement module group 07 | 43902 | Float | RD | A | 1. harmonic I4
Current measurement module group 07 | 43904 | Float | RD | A | 3. harmonic I4
Current measurement module group 07 | 43906 | Float | RD | A | 5. harmonic I4
Current measurement module group 07 | 43908 | Float | RD | A | 7. harmonic I4
Current measurement module group 07 | 43910 | Float | RD | A | 9. harmonic I4
Current measurement module group 07 | 43912 | Float | RD | A | 11. harmonic I4
Current measurement module group 07 | 43914 | Float | RD | A | 13. harmonic I4
Current measurement module group 07 | 43916 | Float | RD | A | 15. harmonic I4
Current measurement module group 07 | 43918 | Float | RD | A | 17. harmonic I4
Current measurement module group 07 | 43920 | Float | RD | A | 19. harmonic I4
Current measurement module group 07 | 43922 | Float | RD | A | 21. harmonic I4
Current measurement module group 07 | 43924 | Float | RD | A | 23. harmonic I4
Current measurement module group 07 | 43926 | Float | RD | A | 25. harmonic I4
Current measurement module group 08 | 43928 | Float | RD | A | 1. harmonic I1
Current measurement module group 08 | 43930 | Float | RD | A | 3. harmonic I1
Current measurement module group 08 | 43932 | Float | RD | A | 5. harmonic I1
Current measurement module group 08 | 43934 | Float | RD | A | 7. harmonic I1
Current measurement module group 08 | 43936 | Float | RD | A | 9. harmonic I1
Current measurement module group 08 | 43938 | Float | RD | A | 11. harmonic I1
Current measurement module group 08 | 43940 | Float | RD | A | 13. harmonic I1
Current measurement module group 08 | 43942 | Float | RD | A | 15. harmonic I1
Current measurement module group 08 | 43944 | Float | RD | A | 17. harmonic I1
Current measurement module group 08 | 43946 | Float | RD | A | 19. harmonic I1
Current measurement module group 08 | 43948 | Float | RD | A | 21. harmonic I1
Current measurement module group 08 | 43950 | Float | RD | A | 23. harmonic I1
Current measurement module group 08 | 43952 | Float | RD | A | 25. harmonic I1
Current measurement module group 08 | 43954 | Float | RD | A | 1. harmonic I2
Current measurement module group 08 | 43956 | Float | RD | A | 3. harmonic I2
Current measurement module group 08 | 43958 | Float | RD | A | 5. harmonic I2
Current measurement module group 08 | 43960 | Float | RD | A | 7. harmonic I2
Current measurement module group 08 | 43962 | Float | RD | A | 9. harmonic I2
Current measurement module group 08 | 43964 | Float | RD | A | 11. harmonic I2
Current measurement module group 08 | 43966 | Float | RD | A | 13. harmonic I2
Current measurement module group 08 | 43968 | Float | RD | A | 15. harmonic I2
Current measurement module group 08 | 43970 | Float | RD | A | 17. harmonic I2
Current measurement module group 08 | 43972 | Float | RD | A | 19. harmonic I2
Current measurement module group 08 | 43974 | Float | RD | A | 21. harmonic I2
Current measurement module group 08 | 43976 | Float | RD | A | 23. harmonic I2
Current measurement module group 08 | 43978 | Float | RD | A | 25. harmonic I2
Current measurement module group 08 | 43980 | Float | RD | A | 1. harmonic I3
Current measurement module group 08 | 43982 | Float | RD | A | 3. harmonic I3
Current measurement module group 08 | 43984 | Float | RD | A | 5. harmonic I3
Current measurement module group 08 | 43986 | Float | RD | A | 7. harmonic I3
Current measurement module group 08 | 43988 | Float | RD | A | 9. harmonic I3
Current measurement module group 08 | 43990 | Float | RD | A | 11. harmonic I3
Current measurement module group 08 | 43992 | Float | RD | A | 13. harmonic I3
Current measurement module group 08 | 43994 | Float | RD | A | 15. harmonic I3
Current measurement module group 08 | 43996 | Float | RD | A | 17. harmonic I3
Current measurement module group 08 | 43998 | Float | RD | A | 19. harmonic I3
Current measurement module group 08 | 44000 | Float | RD | A | 21. harmonic I3
Current measurement module group 08 | 44002 | Float | RD | A | 23. harmonic I3
Current measurement module group 08 | 44004 | Float | RD | A | 25. harmonic I3
Current measurement module group 08 | 44006 | Float | RD | A | 1. harmonic I4
Current measurement module group 08 | 44008 | Float | RD | A | 3. harmonic I4
Current measurement module group 08 | 44010 | Float | RD | A | 5. harmonic I4
Current measurement module group 08 | 44012 | Float | RD | A | 7. harmonic I4
Current measurement module group 08 | 44014 | Float | RD | A | 9. harmonic I4
Current measurement module group 08 | 44016 | Float | RD | A | 11. harmonic I4
Current measurement module group 08 | 44018 | Float | RD | A | 13. harmonic I4
Current measurement module group 08 | 44020 | Float | RD | A | 15. harmonic I4
Current measurement module group 08 | 44022 | Float | RD | A | 17. harmonic I4
Current measurement module group 08 | 44024 | Float | RD | A | 19. harmonic I4
Current measurement module group 08 | 44026 | Float | RD | A | 21. harmonic I4
Current measurement module group 08 | 44028 | Float | RD | A | 23. harmonic I4
Current measurement module group 08 | 44030 | Float | RD | A | 25. harmonic I4
Current measurement module group 09 | 44032 | Float | RD | A | 1. harmonic I1
Current measurement module group 09 | 44034 | Float | RD | A | 3. harmonic I1
Current measurement module group 09 | 44036 | Float | RD | A | 5. harmonic I1
Current measurement module group 09 | 44038 | Float | RD | A | 7. harmonic I1
Current measurement module group 09 | 44040 | Float | RD | A | 9. harmonic I1
Current measurement module group 09 | 44042 | Float | RD | A | 11. harmonic I1
Current measurement module group 09 | 44044 | Float | RD | A | 13. harmonic I1
Current measurement module group 09 | 44046 | Float | RD | A | 15. harmonic I1
Current measurement module group 09 | 44048 | Float | RD | A | 17. harmonic I1
Current measurement module group 09 | 44050 | Float | RD | A | 19. harmonic I1
Current measurement module group 09 | 44052 | Float | RD | A | 21. harmonic I1
Current measurement module group 09 | 44054 | Float | RD | A | 23. harmonic I1
Current measurement module group 09 | 44056 | Float | RD | A | 25. harmonic I1
Current measurement module group 09 | 44058 | Float | RD | A | 1. harmonic I2
Current measurement module group 09 | 44060 | Float | RD | A | 3. harmonic I2
Current measurement module group 09 | 44062 | Float | RD | A | 5. harmonic I2
Current measurement module group 09 | 44064 | Float | RD | A | 7. harmonic I2
Current measurement module group 09 | 44066 | Float | RD | A | 9. harmonic I2
Current measurement module group 09 | 44068 | Float | RD | A | 11. harmonic I2
Current measurement module group 09 | 44070 | Float | RD | A | 13. harmonic I2
Current measurement module group 09 | 44072 | Float | RD | A | 15. harmonic I2
Current measurement module group 09 | 44074 | Float | RD | A | 17. harmonic I2
Current measurement module group 09 | 44076 | Float | RD | A | 19. harmonic I2
Current measurement module group 09 | 44078 | Float | RD | A | 21. harmonic I2
Current measurement module group 09 | 44080 | Float | RD | A | 23. harmonic I2
Current measurement module group 09 | 44082 | Float | RD | A | 25. harmonic I2
Current measurement module group 09 | 44084 | Float | RD | A | 1. harmonic I3
Current measurement module group 09 | 44086 | Float | RD | A | 3. harmonic I3
Current measurement module group 09 | 44088 | Float | RD | A | 5. harmonic I3
Current measurement module group 09 | 44090 | Float | RD | A | 7. harmonic I3
Current measurement module group 09 | 44092 | Float | RD | A | 9. harmonic I3
Current measurement module group 09 | 44094 | Float | RD | A | 11. harmonic I3
Current measurement module group 09 | 44096 | Float | RD | A | 13. harmonic I3
Current measurement module group 09 | 44098 | Float | RD | A | 15. harmonic I3
Current measurement module group 09 | 44100 | Float | RD | A | 17. harmonic I3
Current measurement module group 09 | 44102 | Float | RD | A | 19. harmonic I3
Current measurement module group 09 | 44104 | Float | RD | A | 21. harmonic I3
Current measurement module group 09 | 44106 | Float | RD | A | 23. harmonic I3
Current measurement module group 09 | 44108 | Float | RD | A | 25. harmonic I3
Current measurement module group 09 | 44110 | Float | RD | A | 1. harmonic I4
Current measurement module group 09 | 44112 | Float | RD | A | 3. harmonic I4
Current measurement module group 09 | 44114 | Float | RD | A | 5. harmonic I4
Current measurement module group 09 | 44116 | Float | RD | A | 7. harmonic I4
Current measurement module group 09 | 44118 | Float | RD | A | 9. harmonic I4
Current measurement module group 09 | 44120 | Float | RD | A | 11. harmonic I4
Current measurement module group 09 | 44122 | Float | RD | A | 13. harmonic I4
Current measurement module group 09 | 44124 | Float | RD | A | 15. harmonic I4
Current measurement module group 09 | 44126 | Float | RD | A | 17. harmonic I4
Current measurement module group 09 | 44128 | Float | RD | A | 19. harmonic I4
Current measurement module group 09 | 44130 | Float | RD | A | 21. harmonic I4
Current measurement module group 09 | 44132 | Float | RD | A | 23. harmonic I4
Current measurement module group 09 | 44134 | Float | RD | A | 25. harmonic I4
Current measurement module group 10 | 44136 | Float | RD | A | 1. harmonic I1
Current measurement module group 10 | 44138 | Float | RD | A | 3. harmonic I1
Current measurement module group 10 | 44140 | Float | RD | A | 5. harmonic I1
Current measurement module group 10 | 44142 | Float | RD | A | 7. harmonic I1
Current measurement module group 10 | 44144 | Float | RD | A | 9. harmonic I1
Current measurement module group 10 | 44146 | Float | RD | A | 11. harmonic I1
Current measurement module group 10 | 44148 | Float | RD | A | 13. harmonic I1
Current measurement module group 10 | 44150 | Float | RD | A | 15. harmonic I1
Current measurement module group 10 | 44152 | Float | RD | A | 17. harmonic I1
Current measurement module group 10 | 44154 | Float | RD | A | 19. harmonic I1
Current measurement module group 10 | 44156 | Float | RD | A | 21. harmonic I1
Current measurement module group 10 | 44158 | Float | RD | A | 23. harmonic I1
Current measurement module group 10 | 44160 | Float | RD | A | 25. harmonic I1
Current measurement module group 10 | 44162 | Float | RD | A | 1. harmonic I2
Current measurement module group 10 | 44164 | Float | RD | A | 3. harmonic I2
Current measurement module group 10 | 44166 | Float | RD | A | 5. harmonic I2
Current measurement module group 10 | 44168 | Float | RD | A | 7. harmonic I2
Current measurement module group 10 | 44170 | Float | RD | A | 9. harmonic I2
Current measurement module group 10 | 44172 | Float | RD | A | 11. harmonic I2
Current measurement module group 10 | 44174 | Float | RD | A | 13. harmonic I2
Current measurement module group 10 | 44176 | Float | RD | A | 15. harmonic I2
Current measurement module group 10 | 44178 | Float | RD | A | 17. harmonic I2
Current measurement module group 10 | 44180 | Float | RD | A | 19. harmonic I2
Current measurement module group 10 | 44182 | Float | RD | A | 21. harmonic I2
Current measurement module group 10 | 44184 | Float | RD | A | 23. harmonic I2
Current measurement module group 10 | 44186 | Float | RD | A | 25. harmonic I2
Current measurement module group 10 | 44188 | Float | RD | A | 1. harmonic I3
Current measurement module group 10 | 44190 | Float | RD | A | 3. harmonic I3
Current measurement module group 10 | 44192 | Float | RD | A | 5. harmonic I3
Current measurement module group 10 | 44194 | Float | RD | A | 7. harmonic I3
Current measurement module group 10 | 44196 | Float | RD | A | 9. harmonic I3
Current measurement module group 10 | 44198 | Float | RD | A | 11. harmonic I3
Current measurement module group 10 | 44200 | Float | RD | A | 13. harmonic I3
Current measurement module group 10 | 44202 | Float | RD | A | 15. harmonic I3
Current measurement module group 10 | 44204 | Float | RD | A | 17. harmonic I3
Current measurement module group 10 | 44206 | Float | RD | A | 19. harmonic I3
Current measurement module group 10 | 44208 | Float | RD | A | 21. harmonic I3
Current measurement module group 10 | 44210 | Float | RD | A | 23. harmonic I3
Current measurement module group 10 | 44212 | Float | RD | A | 25. harmonic I3
Current measurement module group 10 | 44214 | Float | RD | A | 1. harmonic I4
Current measurement module group 10 | 44216 | Float | RD | A | 3. harmonic I4
Current measurement module group 10 | 44218 | Float | RD | A | 5. harmonic I4
Current measurement module group 10 | 44220 | Float | RD | A | 7. harmonic I4
Current measurement module group 10 | 44222 | Float | RD | A | 9. harmonic I4
Current measurement module group 10 | 44224 | Float | RD | A | 11. harmonic I4
Current measurement module group 10 | 44226 | Float | RD | A | 13. harmonic I4
Current measurement module group 10 | 44228 | Float | RD | A | 15. harmonic I4
Current measurement module group 10 | 44230 | Float | RD | A | 17. harmonic I4
Current measurement module group 10 | 44232 | Float | RD | A | 19. harmonic I4
Current measurement module group 10 | 44234 | Float | RD | A | 21. harmonic I4
Current measurement module group 10 | 44236 | Float | RD | A | 23. harmonic I4
Current measurement module group 10 | 44238 | Float | RD | A | 25. harmonic I4
Current measurement module group 11 | 44240 | Float | RD | A | 1. harmonic I1
Current measurement module group 11 | 44242 | Float | RD | A | 3. harmonic I1
Current measurement module group 11 | 44244 | Float | RD | A | 5. harmonic I1
Current measurement module group 11 | 44246 | Float | RD | A | 7. harmonic I1
Current measurement module group 11 | 44248 | Float | RD | A | 9. harmonic I1
Current measurement module group 11 | 44250 | Float | RD | A | 11. harmonic I1
Current measurement module group 11 | 44252 | Float | RD | A | 13. harmonic I1
Current measurement module group 11 | 44254 | Float | RD | A | 15. harmonic I1
Current measurement module group 11 | 44256 | Float | RD | A | 17. harmonic I1
Current measurement module group 11 | 44258 | Float | RD | A | 19. harmonic I1
Current measurement module group 11 | 44260 | Float | RD | A | 21. harmonic I1
Current measurement module group 11 | 44262 | Float | RD | A | 23. harmonic I1
Current measurement module group 11 | 44264 | Float | RD | A | 25. harmonic I1
Current measurement module group 11 | 44266 | Float | RD | A | 1. harmonic I2
Current measurement module group 11 | 44268 | Float | RD | A | 3. harmonic I2
Current measurement module group 11 | 44270 | Float | RD | A | 5. harmonic I2
Current measurement module group 11 | 44272 | Float | RD | A | 7. harmonic I2
Current measurement module group 11 | 44274 | Float | RD | A | 9. harmonic I2
Current measurement module group 11 | 44276 | Float | RD | A | 11. harmonic I2
Current measurement module group 11 | 44278 | Float | RD | A | 13. harmonic I2
Current measurement module group 11 | 44280 | Float | RD | A | 15. harmonic I2
Current measurement module group 11 | 44282 | Float | RD | A | 17. harmonic I2
Current measurement module group 11 | 44284 | Float | RD | A | 19. harmonic I2
Current measurement module group 11 | 44286 | Float | RD | A | 21. harmonic I2
Current measurement module group 11 | 44288 | Float | RD | A | 23. harmonic I2
Current measurement module group 11 | 44290 | Float | RD | A | 25. harmonic I2
Current measurement module group 11 | 44292 | Float | RD | A | 1. harmonic I3
Current measurement module group 11 | 44294 | Float | RD | A | 3. harmonic I3
Current measurement module group 11 | 44296 | Float | RD | A | 5. harmonic I3
Current measurement module group 11 | 44298 | Float | RD | A | 7. harmonic I3
Current measurement module group 11 | 44300 | Float | RD | A | 9. harmonic I3
Current measurement module group 11 | 44302 | Float | RD | A | 11. harmonic I3
Current measurement module group 11 | 44304 | Float | RD | A | 13. harmonic I3
Current measurement module group 11 | 44306 | Float | RD | A | 15. harmonic I3
Current measurement module group 11 | 44308 | Float | RD | A | 17. harmonic I3
Current measurement module group 11 | 44310 | Float | RD | A | 19. harmonic I3
Current measurement module group 11 | 44312 | Float | RD | A | 21. harmonic I3
Current measurement module group 11 | 44314 | Float | RD | A | 23. harmonic I3
Current measurement module group 11 | 44316 | Float | RD | A | 25. harmonic I3
Current measurement module group 11 | 44318 | Float | RD | A | 1. harmonic I4
Current measurement module group 11 | 44320 | Float | RD | A | 3. harmonic I4
Current measurement module group 11 | 44322 | Float | RD | A | 5. harmonic I4
Current measurement module group 11 | 44324 | Float | RD | A | 7. harmonic I4
Current measurement module group 11 | 44326 | Float | RD | A | 9. harmonic I4
Current measurement module group 11 | 44328 | Float | RD | A | 11. harmonic I4
Current measurement module group 11 | 44330 | Float | RD | A | 13. harmonic I4
Current measurement module group 11 | 44332 | Float | RD | A | 15. harmonic I4
Current measurement module group 11 | 44334 | Float | RD | A | 17. harmonic I4
Current measurement module group 11 | 44336 | Float | RD | A | 19. harmonic I4
Current measurement module group 11 | 44338 | Float | RD | A | 21. harmonic I4
Current measurement module group 11 | 44340 | Float | RD | A | 23. harmonic I4
Current measurement module group 11 | 44342 | Float | RD | A | 25. harmonic I4
Current measurement module group 12 | 44344 | Float | RD | A | 1. harmonic I1
Current measurement module group 12 | 44346 | Float | RD | A | 3. harmonic I1
Current measurement module group 12 | 44348 | Float | RD | A | 5. harmonic I1
Current measurement module group 12 | 44350 | Float | RD | A | 7. harmonic I1
Current measurement module group 12 | 44352 | Float | RD | A | 9. harmonic I1
Current measurement module group 12 | 44354 | Float | RD | A | 11. harmonic I1
Current measurement module group 12 | 44356 | Float | RD | A | 13. harmonic I1
Current measurement module group 12 | 44358 | Float | RD | A | 15. harmonic I1
Current measurement module group 12 | 44360 | Float | RD | A | 17. harmonic I1
Current measurement module group 12 | 44362 | Float | RD | A | 19. harmonic I1
Current measurement module group 12 | 44364 | Float | RD | A | 21. harmonic I1
Current measurement module group 12 | 44366 | Float | RD | A | 23. harmonic I1
Current measurement module group 12 | 44368 | Float | RD | A | 25. harmonic I1
Current measurement module group 12 | 44370 | Float | RD | A | 1. harmonic I2
Current measurement module group 12 | 44372 | Float | RD | A | 3. harmonic I2
Current measurement module group 12 | 44374 | Float | RD | A | 5. harmonic I2
Current measurement module group 12 | 44376 | Float | RD | A | 7. harmonic I2
Current measurement module group 12 | 44378 | Float | RD | A | 9. harmonic I2
Current measurement module group 12 | 44380 | Float | RD | A | 11. harmonic I2
Current measurement module group 12 | 44382 | Float | RD | A | 13. harmonic I2
Current measurement module group 12 | 44384 | Float | RD | A | 15. harmonic I2
Current measurement module group 12 | 44386 | Float | RD | A | 17. harmonic I2
Current measurement module group 12 | 44388 | Float | RD | A | 19. harmonic I2
Current measurement module group 12 | 44390 | Float | RD | A | 21. harmonic I2
Current measurement module group 12 | 44392 | Float | RD | A | 23. harmonic I2
Current measurement module group 12 | 44394 | Float | RD | A | 25. harmonic I2
Current measurement module group 12 | 44396 | Float | RD | A | 1. harmonic I3
Current measurement module group 12 | 44398 | Float | RD | A | 3. harmonic I3
Current measurement module group 12 | 44400 | Float | RD | A | 5. harmonic I3
Current measurement module group 12 | 44402 | Float | RD | A | 7. harmonic I3
Current measurement module group 12 | 44404 | Float | RD | A | 9. harmonic I3
Current measurement module group 12 | 44406 | Float | RD | A | 11. harmonic I3
Current measurement module group 12 | 44408 | Float | RD | A | 13. harmonic I3
Current measurement module group 12 | 44410 | Float | RD | A | 15. harmonic I3
Current measurement module group 12 | 44412 | Float | RD | A | 17. harmonic I3
Current measurement module group 12 | 44414 | Float | RD | A | 19. harmonic I3
Current measurement module group 12 | 44416 | Float | RD | A | 21. harmonic I3
Current measurement module group 12 | 44418 | Float | RD | A | 23. harmonic I3
Current measurement module group 12 | 44420 | Float | RD | A | 25. harmonic I3
Current measurement module group 12 | 44422 | Float | RD | A | 1. harmonic I4
Current measurement module group 12 | 44424 | Float | RD | A | 3. harmonic I4
Current measurement module group 12 | 44426 | Float | RD | A | 5. harmonic I4
Current measurement module group 12 | 44428 | Float | RD | A | 7. harmonic I4
Current measurement module group 12 | 44430 | Float | RD | A | 9. harmonic I4
Current measurement module group 12 | 44432 | Float | RD | A | 11. harmonic I4
Current measurement module group 12 | 44434 | Float | RD | A | 13. harmonic I4
Current measurement module group 12 | 44436 | Float | RD | A | 15. harmonic I4
Current measurement module group 12 | 44438 | Float | RD | A | 17. harmonic I4
Current measurement module group 12 | 44440 | Float | RD | A | 19. harmonic I4
Current measurement module group 12 | 44442 | Float | RD | A | 21. harmonic I4
Current measurement module group 12 | 44444 | Float | RD | A | 23. harmonic I4
Current measurement module group 12 | 44446 | Float | RD | A | 25. harmonic I4
Current measurement module group 13 | 44448 | Float | RD | A | 1. harmonic I1
Current measurement module group 13 | 44450 | Float | RD | A | 3. harmonic I1
Current measurement module group 13 | 44452 | Float | RD | A | 5. harmonic I1
Current measurement module group 13 | 44454 | Float | RD | A | 7. harmonic I1
Current measurement module group 13 | 44456 | Float | RD | A | 9. harmonic I1
Current measurement module group 13 | 44458 | Float | RD | A | 11. harmonic I1
Current measurement module group 13 | 44460 | Float | RD | A | 13. harmonic I1
Current measurement module group 13 | 44462 | Float | RD | A | 15. harmonic I1
Current measurement module group 13 | 44464 | Float | RD | A | 17. harmonic I1
Current measurement module group 13 | 44466 | Float | RD | A | 19. harmonic I1
Current measurement module group 13 | 44468 | Float | RD | A | 21. harmonic I1
Current measurement module group 13 | 44470 | Float | RD | A | 23. harmonic I1
Current measurement module group 13 | 44472 | Float | RD | A | 25. harmonic I1
Current measurement module group 13 | 44474 | Float | RD | A | 1. harmonic I2
Current measurement module group 13 | 44476 | Float | RD | A | 3. harmonic I2
Current measurement module group 13 | 44478 | Float | RD | A | 5. harmonic I2
Current measurement module group 13 | 44480 | Float | RD | A | 7. harmonic I2
Current measurement module group 13 | 44482 | Float | RD | A | 9. harmonic I2
Current measurement module group 13 | 44484 | Float | RD | A | 11. harmonic I2
Current measurement module group 13 | 44486 | Float | RD | A | 13. harmonic I2
Current measurement module group 13 | 44488 | Float | RD | A | 15. harmonic I2
Current measurement module group 13 | 44490 | Float | RD | A | 17. harmonic I2
Current measurement module group 13 | 44492 | Float | RD | A | 19. harmonic I2
Current measurement module group 13 | 44494 | Float | RD | A | 21. harmonic I2
Current measurement module group 13 | 44496 | Float | RD | A | 23. harmonic I2
Current measurement module group 13 | 44498 | Float | RD | A | 25. harmonic I2
Current measurement module group 13 | 44500 | Float | RD | A | 1. harmonic I3
Current measurement module group 13 | 44502 | Float | RD | A | 3. harmonic I3
Current measurement module group 13 | 44504 | Float | RD | A | 5. harmonic I3
Current measurement module group 13 | 44506 | Float | RD | A | 7. harmonic I3
Current measurement module group 13 | 44508 | Float | RD | A | 9. harmonic I3
Current measurement module group 13 | 44510 | Float | RD | A | 11. harmonic I3
Current measurement module group 13 | 44512 | Float | RD | A | 13. harmonic I3
Current measurement module group 13 | 44514 | Float | RD | A | 15. harmonic I3
Current measurement module group 13 | 44516 | Float | RD | A | 17. harmonic I3
Current measurement module group 13 | 44518 | Float | RD | A | 19. harmonic I3
Current measurement module group 13 | 44520 | Float | RD | A | 21. harmonic I3
Current measurement module group 13 | 44522 | Float | RD | A | 23. harmonic I3
Current measurement module group 13 | 44524 | Float | RD | A | 25. harmonic I3
Current measurement module group 13 | 44526 | Float | RD | A | 1. harmonic I4
Current measurement module group 13 | 44528 | Float | RD | A | 3. harmonic I4
Current measurement module group 13 | 44530 | Float | RD | A | 5. harmonic I4
Current measurement module group 13 | 44532 | Float | RD | A | 7. harmonic I4
Current measurement module group 13 | 44534 | Float | RD | A | 9. harmonic I4
Current measurement module group 13 | 44536 | Float | RD | A | 11. harmonic I4
Current measurement module group 13 | 44538 | Float | RD | A | 13. harmonic I4
Current measurement module group 13 | 44540 | Float | RD | A | 15. harmonic I4
Current measurement module group 13 | 44542 | Float | RD | A | 17. harmonic I4
Current measurement module group 13 | 44544 | Float | RD | A | 19. harmonic I4
Current measurement module group 13 | 44546 | Float | RD | A | 21. harmonic I4
Current measurement module group 13 | 44548 | Float | RD | A | 23. harmonic I4
Current measurement module group 13 | 44550 | Float | RD | A | 25. harmonic I4
Current measurement module group 14 | 44552 | Float | RD | A | 1. harmonic I1
Current measurement module group 14 | 44554 | Float | RD | A | 3. harmonic I1
Current measurement module group 14 | 44556 | Float | RD | A | 5. harmonic I1
Current measurement module group 14 | 44558 | Float | RD | A | 7. harmonic I1
Current measurement module group 14 | 44560 | Float | RD | A | 9. harmonic I1
Current measurement module group 14 | 44562 | Float | RD | A | 11. harmonic I1
Current measurement module group 14 | 44564 | Float | RD | A | 13. harmonic I1
Current measurement module group 14 | 44566 | Float | RD | A | 15. harmonic I1
Current measurement module group 14 | 44568 | Float | RD | A | 17. harmonic I1
Current measurement module group 14 | 44570 | Float | RD | A | 19. harmonic I1
Current measurement module group 14 | 44572 | Float | RD | A | 21. harmonic I1
Current measurement module group 14 | 44574 | Float | RD | A | 23. harmonic I1
Current measurement module group 14 | 44576 | Float | RD | A | 25. harmonic I1
Current measurement module group 14 | 44578 | Float | RD | A | 1. harmonic I2
Current measurement module group 14 | 44580 | Float | RD | A | 3. harmonic I2
Current measurement module group 14 | 44582 | Float | RD | A | 5. harmonic I2
Current measurement module group 14 | 44584 | Float | RD | A | 7. harmonic I2
Current measurement module group 14 | 44586 | Float | RD | A | 9. harmonic I2
Current measurement module group 14 | 44588 | Float | RD | A | 11. harmonic I2
Current measurement module group 14 | 44590 | Float | RD | A | 13. harmonic I2
Current measurement module group 14 | 44592 | Float | RD | A | 15. harmonic I2
Current measurement module group 14 | 44594 | Float | RD | A | 17. harmonic I2
Current measurement module group 14 | 44596 | Float | RD | A | 19. harmonic I2
Current measurement module group 14 | 44598 | Float | RD | A | 21. harmonic I2
Current measurement module group 14 | 44600 | Float | RD | A | 23. harmonic I2
Current measurement module group 14 | 44602 | Float | RD | A | 25. harmonic I2
Current measurement module group 14 | 44604 | Float | RD | A | 1. harmonic I3
Current measurement module group 14 | 44606 | Float | RD | A | 3. harmonic I3
Current measurement module group 14 | 44608 | Float | RD | A | 5. harmonic I3
Current measurement module group 14 | 44610 | Float | RD | A | 7. harmonic I3
Current measurement module group 14 | 44612 | Float | RD | A | 9. harmonic I3
Current measurement module group 14 | 44614 | Float | RD | A | 11. harmonic I3
Current measurement module group 14 | 44616 | Float | RD | A | 13. harmonic I3
Current measurement module group 14 | 44618 | Float | RD | A | 15. harmonic I3
Current measurement module group 14 | 44620 | Float | RD | A | 17. harmonic I3
Current measurement module group 14 | 44622 | Float | RD | A | 19. harmonic I3
Current measurement module group 14 | 44624 | Float | RD | A | 21. harmonic I3
Current measurement module group 14 | 44626 | Float | RD | A | 23. harmonic I3
Current measurement module group 14 | 44628 | Float | RD | A | 25. harmonic I3
Current measurement module group 14 | 44630 | Float | RD | A | 1. harmonic I4
Current measurement module group 14 | 44632 | Float | RD | A | 3. harmonic I4
Current measurement module group 14 | 44634 | Float | RD | A | 5. harmonic I4
Current measurement module group 14 | 44636 | Float | RD | A | 7. harmonic I4
Current measurement module group 14 | 44638 | Float | RD | A | 9. harmonic I4
Current measurement module group 14 | 44640 | Float | RD | A | 11. harmonic I4
Current measurement module group 14 | 44642 | Float | RD | A | 13. harmonic I4
Current measurement module group 14 | 44644 | Float | RD | A | 15. harmonic I4
Current measurement module group 14 | 44646 | Float | RD | A | 17. harmonic I4
Current measurement module group 14 | 44648 | Float | RD | A | 19. harmonic I4
Current measurement module group 14 | 44650 | Float | RD | A | 21. harmonic I4
Current measurement module group 14 | 44652 | Float | RD | A | 23. harmonic I4
Current measurement module group 14 | 44654 | Float | RD | A | 25. harmonic I4
Current measurement module group 15 | 44656 | Float | RD | A | 1. harmonic I1
Current measurement module group 15 | 44658 | Float | RD | A | 3. harmonic I1
Current measurement module group 15 | 44660 | Float | RD | A | 5. harmonic I1
Current measurement module group 15 | 44662 | Float | RD | A | 7. harmonic I1
Current measurement module group 15 | 44664 | Float | RD | A | 9. harmonic I1
Current measurement module group 15 | 44666 | Float | RD | A | 11. harmonic I1
Current measurement module group 15 | 44668 | Float | RD | A | 13. harmonic I1
Current measurement module group 15 | 44670 | Float | RD | A | 15. harmonic I1
Current measurement module group 15 | 44672 | Float | RD | A | 17. harmonic I1
Current measurement module group 15 | 44674 | Float | RD | A | 19. harmonic I1
Current measurement module group 15 | 44676 | Float | RD | A | 21. harmonic I1
Current measurement module group 15 | 44678 | Float | RD | A | 23. harmonic I1
Current measurement module group 15 | 44680 | Float | RD | A | 25. harmonic I1
Current measurement module group 15 | 44682 | Float | RD | A | 1. harmonic I2
Current measurement module group 15 | 44684 | Float | RD | A | 3. harmonic I2
Current measurement module group 15 | 44686 | Float | RD | A | 5. harmonic I2
Current measurement module group 15 | 44688 | Float | RD | A | 7. harmonic I2
Current measurement module group 15 | 44690 | Float | RD | A | 9. harmonic I2
Current measurement module group 15 | 44692 | Float | RD | A | 11. harmonic I2
Current measurement module group 15 | 44694 | Float | RD | A | 13. harmonic I2
Current measurement module group 15 | 44696 | Float | RD | A | 15. harmonic I2
Current measurement module group 15 | 44698 | Float | RD | A | 17. harmonic I2
Current measurement module group 15 | 44700 | Float | RD | A | 19. harmonic I2
Current measurement module group 15 | 44702 | Float | RD | A | 21. harmonic I2
Current measurement module group 15 | 44704 | Float | RD | A | 23. harmonic I2
Current measurement module group 15 | 44706 | Float | RD | A | 25. harmonic I2
Current measurement module group 15 | 44708 | Float | RD | A | 1. harmonic I3
Current measurement module group 15 | 44710 | Float | RD | A | 3. harmonic I3
Current measurement module group 15 | 44712 | Float | RD | A | 5. harmonic I3
Current measurement module group 15 | 44714 | Float | RD | A | 7. harmonic I3
Current measurement module group 15 | 44716 | Float | RD | A | 9. harmonic I3
Current measurement module group 15 | 44718 | Float | RD | A | 11. harmonic I3
Current measurement module group 15 | 44720 | Float | RD | A | 13. harmonic I3
Current measurement module group 15 | 44722 | Float | RD | A | 15. harmonic I3
Current measurement module group 15 | 44724 | Float | RD | A | 17. harmonic I3
Current measurement module group 15 | 44726 | Float | RD | A | 19. harmonic I3
Current measurement module group 15 | 44728 | Float | RD | A | 21. harmonic I3
Current measurement module group 15 | 44730 | Float | RD | A | 23. harmonic I3
Current measurement module group 15 | 44732 | Float | RD | A | 25. harmonic I3
Current measurement module group 15 | 44734 | Float | RD | A | 1. harmonic I4
Current measurement module group 15 | 44736 | Float | RD | A | 3. harmonic I4
Current measurement module group 15 | 44738 | Float | RD | A | 5. harmonic I4
Current measurement module group 15 | 44740 | Float | RD | A | 7. harmonic I4
Current measurement module group 15 | 44742 | Float | RD | A | 9. harmonic I4
Current measurement module group 15 | 44744 | Float | RD | A | 11. harmonic I4
Current measurement module group 15 | 44746 | Float | RD | A | 13. harmonic I4
Current measurement module group 15 | 44748 | Float | RD | A | 15. harmonic I4
Current measurement module group 15 | 44750 | Float | RD | A | 17. harmonic I4
Current measurement module group 15 | 44752 | Float | RD | A | 19. harmonic I4
Current measurement module group 15 | 44754 | Float | RD | A | 21. harmonic I4
Current measurement module group 15 | 44756 | Float | RD | A | 23. harmonic I4
Current measurement module group 15 | 44758 | Float | RD | A | 25. harmonic I4
Current measurement module group 16 | 44760 | Float | RD | A | 1. harmonic I1
Current measurement module group 16 | 44762 | Float | RD | A | 3. harmonic I1
Current measurement module group 16 | 44764 | Float | RD | A | 5. harmonic I1
Current measurement module group 16 | 44766 | Float | RD | A | 7. harmonic I1
Current measurement module group 16 | 44768 | Float | RD | A | 9. harmonic I1
Current measurement module group 16 | 44770 | Float | RD | A | 11. harmonic I1
Current measurement module group 16 | 44772 | Float | RD | A | 13. harmonic I1
Current measurement module group 16 | 44774 | Float | RD | A | 15. harmonic I1
Current measurement module group 16 | 44776 | Float | RD | A | 17. harmonic I1
Current measurement module group 16 | 44778 | Float | RD | A | 19. harmonic I1
Current measurement module group 16 | 44780 | Float | RD | A | 21. harmonic I1
Current measurement module group 16 | 44782 | Float | RD | A | 23. harmonic I1
Current measurement module group 16 | 44784 | Float | RD | A | 25. harmonic I1
Current measurement module group 16 | 44786 | Float | RD | A | 1. harmonic I2
Current measurement module group 16 | 44788 | Float | RD | A | 3. harmonic I2
Current measurement module group 16 | 44790 | Float | RD | A | 5. harmonic I2
Current measurement module group 16 | 44792 | Float | RD | A | 7. harmonic I2
Current measurement module group 16 | 44794 | Float | RD | A | 9. harmonic I2
Current measurement module group 16 | 44796 | Float | RD | A | 11. harmonic I2
Current measurement module group 16 | 44798 | Float | RD | A | 13. harmonic I2
Current measurement module group 16 | 44800 | Float | RD | A | 15. harmonic I2
Current measurement module group 16 | 44802 | Float | RD | A | 17. harmonic I2
Current measurement module group 16 | 44804 | Float | RD | A | 19. harmonic I2
Current measurement module group 16 | 44806 | Float | RD | A | 21. harmonic I2
Current measurement module group 16 | 44808 | Float | RD | A | 23. harmonic I2
Current measurement module group 16 | 44810 | Float | RD | A | 25. harmonic I2
Current measurement module group 16 | 44812 | Float | RD | A | 1. harmonic I3
Current measurement module group 16 | 44814 | Float | RD | A | 3. harmonic I3
Current measurement module group 16 | 44816 | Float | RD | A | 5. harmonic I3
Current measurement module group 16 | 44818 | Float | RD | A | 7. harmonic I3
Current measurement module group 16 | 44820 | Float | RD | A | 9. harmonic I3
Current measurement module group 16 | 44822 | Float | RD | A | 11. harmonic I3
Current measurement module group 16 | 44824 | Float | RD | A | 13. harmonic I3
Current measurement module group 16 | 44826 | Float | RD | A | 15. harmonic I3
Current measurement module group 16 | 44828 | Float | RD | A | 17. harmonic I3
Current measurement module group 16 | 44830 | Float | RD | A | 19. harmonic I3
Current measurement module group 16 | 44832 | Float | RD | A | 21. harmonic I3
Current measurement module group 16 | 44834 | Float | RD | A | 23. harmonic I3
Current measurement module group 16 | 44836 | Float | RD | A | 25. harmonic I3
Current measurement module group 16 | 44838 | Float | RD | A | 1. harmonic I4
Current measurement module group 16 | 44840 | Float | RD | A | 3. harmonic I4
Current measurement module group 16 | 44842 | Float | RD | A | 5. harmonic I4
Current measurement module group 16 | 44844 | Float | RD | A | 7. harmonic I4
Current measurement module group 16 | 44846 | Float | RD | A | 9. harmonic I4
Current measurement module group 16 | 44848 | Float | RD | A | 11. harmonic I4
Current measurement module group 16 | 44850 | Float | RD | A | 13. harmonic I4
Current measurement module group 16 | 44852 | Float | RD | A | 15. harmonic I4
Current measurement module group 16 | 44854 | Float | RD | A | 17. harmonic I4
Current measurement module group 16 | 44856 | Float | RD | A | 19. harmonic I4
Current measurement module group 16 | 44858 | Float | RD | A | 21. harmonic I4
Current measurement module group 16 | 44860 | Float | RD | A | 23. harmonic I4
Current measurement module group 16 | 44862 | Float | RD | A | 25. harmonic I4
Current measurement module group 17 | 44864 | Float | RD | A | 1. harmonic I1
Current measurement module group 17 | 44866 | Float | RD | A | 3. harmonic I1
Current measurement module group 17 | 44868 | Float | RD | A | 5. harmonic I1
Current measurement module group 17 | 44870 | Float | RD | A | 7. harmonic I1
Current measurement module group 17 | 44872 | Float | RD | A | 9. harmonic I1
Current measurement module group 17 | 44874 | Float | RD | A | 11. harmonic I1
Current measurement module group 17 | 44876 | Float | RD | A | 13. harmonic I1
Current measurement module group 17 | 44878 | Float | RD | A | 15. harmonic I1
Current measurement module group 17 | 44880 | Float | RD | A | 17. harmonic I1
Current measurement module group 17 | 44882 | Float | RD | A | 19. harmonic I1
Current measurement module group 17 | 44884 | Float | RD | A | 21. harmonic I1
Current measurement module group 17 | 44886 | Float | RD | A | 23. harmonic I1
Current measurement module group 17 | 44888 | Float | RD | A | 25. harmonic I1
Current measurement module group 17 | 44890 | Float | RD | A | 1. harmonic I2
Current measurement module group 17 | 44892 | Float | RD | A | 3. harmonic I2
Current measurement module group 17 | 44894 | Float | RD | A | 5. harmonic I2
Current measurement module group 17 | 44896 | Float | RD | A | 7. harmonic I2
Current measurement module group 17 | 44898 | Float | RD | A | 9. harmonic I2
Current measurement module group 17 | 44900 | Float | RD | A | 11. harmonic I2
Current measurement module group 17 | 44902 | Float | RD | A | 13. harmonic I2
Current measurement module group 17 | 44904 | Float | RD | A | 15. harmonic I2
Current measurement module group 17 | 44906 | Float | RD | A | 17. harmonic I2
Current measurement module group 17 | 44908 | Float | RD | A | 19. harmonic I2
Current measurement module group 17 | 44910 | Float | RD | A | 21. harmonic I2
Current measurement module group 17 | 44912 | Float | RD | A | 23. harmonic I2
Current measurement module group 17 | 44914 | Float | RD | A | 25. harmonic I2
Current measurement module group 17 | 44916 | Float | RD | A | 1. harmonic I3
Current measurement module group 17 | 44918 | Float | RD | A | 3. harmonic I3
Current measurement module group 17 | 44920 | Float | RD | A | 5. harmonic I3
Current measurement module group 17 | 44922 | Float | RD | A | 7. harmonic I3
Current measurement module group 17 | 44924 | Float | RD | A | 9. harmonic I3
Current measurement module group 17 | 44926 | Float | RD | A | 11. harmonic I3
Current measurement module group 17 | 44928 | Float | RD | A | 13. harmonic I3
Current measurement module group 17 | 44930 | Float | RD | A | 15. harmonic I3
Current measurement module group 17 | 44932 | Float | RD | A | 17. harmonic I3
Current measurement module group 17 | 44934 | Float | RD | A | 19. harmonic I3
Current measurement module group 17 | 44936 | Float | RD | A | 21. harmonic I3
Current measurement module group 17 | 44938 | Float | RD | A | 23. harmonic I3
Current measurement module group 17 | 44940 | Float | RD | A | 25. harmonic I3
Current measurement module group 17 | 44942 | Float | RD | A | 1. harmonic I4
Current measurement module group 17 | 44944 | Float | RD | A | 3. harmonic I4
Current measurement module group 17 | 44946 | Float | RD | A | 5. harmonic I4
Current measurement module group 17 | 44948 | Float | RD | A | 7. harmonic I4
Current measurement module group 17 | 44950 | Float | RD | A | 9. harmonic I4
Current measurement module group 17 | 44952 | Float | RD | A | 11. harmonic I4
Current measurement module group 17 | 44954 | Float | RD | A | 13. harmonic I4
Current measurement module group 17 | 44956 | Float | RD | A | 15. harmonic I4
Current measurement module group 17 | 44958 | Float | RD | A | 17. harmonic I4
Current measurement module group 17 | 44960 | Float | RD | A | 19. harmonic I4
Current measurement module group 17 | 44962 | Float | RD | A | 21. harmonic I4
Current measurement module group 17 | 44964 | Float | RD | A | 23. harmonic I4
Current measurement module group 17 | 44966 | Float | RD | A | 25. harmonic I4
Current measurement module group 18 | 44968 | Float | RD | A | 1. harmonic I1
Current measurement module group 18 | 44970 | Float | RD | A | 3. harmonic I1
Current measurement module group 18 | 44972 | Float | RD | A | 5. harmonic I1
Current measurement module group 18 | 44974 | Float | RD | A | 7. harmonic I1
Current measurement module group 18 | 44976 | Float | RD | A | 9. harmonic I1
Current measurement module group 18 | 44978 | Float | RD | A | 11. harmonic I1
Current measurement module group 18 | 44980 | Float | RD | A | 13. harmonic I1
Current measurement module group 18 | 44982 | Float | RD | A | 15. harmonic I1
Current measurement module group 18 | 44984 | Float | RD | A | 17. harmonic I1
Current measurement module group 18 | 44986 | Float | RD | A | 19. harmonic I1
Current measurement module group 18 | 44988 | Float | RD | A | 21. harmonic I1
Current measurement module group 18 | 44990 | Float | RD | A | 23. harmonic I1
Current measurement module group 18 | 44992 | Float | RD | A | 25. harmonic I1
Current measurement module group 18 | 44994 | Float | RD | A | 1. harmonic I2
Current measurement module group 18 | 44996 | Float | RD | A | 3. harmonic I2
Current measurement module group 18 | 44998 | Float | RD | A | 5. harmonic I2
Current measurement module group 18 | 45000 | Float | RD | A | 7. harmonic I2
Current measurement module group 18 | 45002 | Float | RD | A | 9. harmonic I2
Current measurement module group 18 | 45004 | Float | RD | A | 11. harmonic I2
Current measurement module group 18 | 45006 | Float | RD | A | 13. harmonic I2
Current measurement module group 18 | 45008 | Float | RD | A | 15. harmonic I2
Current measurement module group 18 | 45010 | Float | RD | A | 17. harmonic I2
Current measurement module group 18 | 45012 | Float | RD | A | 19. harmonic I2
Current measurement module group 18 | 45014 | Float | RD | A | 21. harmonic I2
Current measurement module group 18 | 45016 | Float | RD | A | 23. harmonic I2
Current measurement module group 18 | 45018 | Float | RD | A | 25. harmonic I2
Current measurement module group 18 | 45020 | Float | RD | A | 1. harmonic I3
Current measurement module group 18 | 45022 | Float | RD | A | 3. harmonic I3
Current measurement module group 18 | 45024 | Float | RD | A | 5. harmonic I3
Current measurement module group 18 | 45026 | Float | RD | A | 7. harmonic I3
Current measurement module group 18 | 45028 | Float | RD | A | 9. harmonic I3
Current measurement module group 18 | 45030 | Float | RD | A | 11. harmonic I3
Current measurement module group 18 | 45032 | Float | RD | A | 13. harmonic I3
Current measurement module group 18 | 45034 | Float | RD | A | 15. harmonic I3
Current measurement module group 18 | 45036 | Float | RD | A | 17. harmonic I3
Current measurement module group 18 | 45038 | Float | RD | A | 19. harmonic I3
Current measurement module group 18 | 45040 | Float | RD | A | 21. harmonic I3
Current measurement module group 18 | 45042 | Float | RD | A | 23. harmonic I3
Current measurement module group 18 | 45044 | Float | RD | A | 25. harmonic I3
Current measurement module group 18 | 45046 | Float | RD | A | 1. harmonic I4
Current measurement module group 18 | 45048 | Float | RD | A | 3. harmonic I4
Current measurement module group 18 | 45050 | Float | RD | A | 5. harmonic I4
Current measurement module group 18 | 45052 | Float | RD | A | 7. harmonic I4
Current measurement module group 18 | 45054 | Float | RD | A | 9. harmonic I4
Current measurement module group 18 | 45056 | Float | RD | A | 11. harmonic I4
Current measurement module group 18 | 45058 | Float | RD | A | 13. harmonic I4
Current measurement module group 18 | 45060 | Float | RD | A | 15. harmonic I4
Current measurement module group 18 | 45062 | Float | RD | A | 17. harmonic I4
Current measurement module group 18 | 45064 | Float | RD | A | 19. harmonic I4
Current measurement module group 18 | 45066 | Float | RD | A | 21. harmonic I4
Current measurement module group 18 | 45068 | Float | RD | A | 23. harmonic I4
Current measurement module group 18 | 45070 | Float | RD | A | 25. harmonic I4
Current measurement module group 19 | 45072 | Float | RD | A | 1. harmonic I1
Current measurement module group 19 | 45074 | Float | RD | A | 3. harmonic I1
Current measurement module group 19 | 45076 | Float | RD | A | 5. harmonic I1
Current measurement module group 19 | 45078 | Float | RD | A | 7. harmonic I1
Current measurement module group 19 | 45080 | Float | RD | A | 9. harmonic I1
Current measurement module group 19 | 45082 | Float | RD | A | 11. harmonic I1
Current measurement module group 19 | 45084 | Float | RD | A | 13. harmonic I1
Current measurement module group 19 | 45086 | Float | RD | A | 15. harmonic I1
Current measurement module group 19 | 45088 | Float | RD | A | 17. harmonic I1
Current measurement module group 19 | 45090 | Float | RD | A | 19. harmonic I1
Current measurement module group 19 | 45092 | Float | RD | A | 21. harmonic I1
Current measurement module group 19 | 45094 | Float | RD | A | 23. harmonic I1
Current measurement module group 19 | 45096 | Float | RD | A | 25. harmonic I1
Current measurement module group 19 | 45098 | Float | RD | A | 1. harmonic I2
Current measurement module group 19 | 45100 | Float | RD | A | 3. harmonic I2
Current measurement module group 19 | 45102 | Float | RD | A | 5. harmonic I2
Current measurement module group 19 | 45104 | Float | RD | A | 7. harmonic I2
Current measurement module group 19 | 45106 | Float | RD | A | 9. harmonic I2
Current measurement module group 19 | 45108 | Float | RD | A | 11. harmonic I2
Current measurement module group 19 | 45110 | Float | RD | A | 13. harmonic I2
Current measurement module group 19 | 45112 | Float | RD | A | 15. harmonic I2
Current measurement module group 19 | 45114 | Float | RD | A | 17. harmonic I2
Current measurement module group 19 | 45116 | Float | RD | A | 19. harmonic I2
Current measurement module group 19 | 45118 | Float | RD | A | 21. harmonic I2
Current measurement module group 19 | 45120 | Float | RD | A | 23. harmonic I2
Current measurement module group 19 | 45122 | Float | RD | A | 25. harmonic I2
Current measurement module group 19 | 45124 | Float | RD | A | 1. harmonic I3
Current measurement module group 19 | 45126 | Float | RD | A | 3. harmonic I3
Current measurement module group 19 | 45128 | Float | RD | A | 5. harmonic I3
Current measurement module group 19 | 45130 | Float | RD | A | 7. harmonic I3
Current measurement module group 19 | 45132 | Float | RD | A | 9. harmonic I3
Current measurement module group 19 | 45134 | Float | RD | A | 11. harmonic I3
Current measurement module group 19 | 45136 | Float | RD | A | 13. harmonic I3
Current measurement module group 19 | 45138 | Float | RD | A | 15. harmonic I3
Current measurement module group 19 | 45140 | Float | RD | A | 17. harmonic I3
Current measurement module group 19 | 45142 | Float | RD | A | 19. harmonic I3
Current measurement module group 19 | 45144 | Float | RD | A | 21. harmonic I3
Current measurement module group 19 | 45146 | Float | RD | A | 23. harmonic I3
Current measurement module group 19 | 45148 | Float | RD | A | 25. harmonic I3
Current measurement module group 19 | 45150 | Float | RD | A | 1. harmonic I4
Current measurement module group 19 | 45152 | Float | RD | A | 3. harmonic I4
Current measurement module group 19 | 45154 | Float | RD | A | 5. harmonic I4
Current measurement module group 19 | 45156 | Float | RD | A | 7. harmonic I4
Current measurement module group 19 | 45158 | Float | RD | A | 9. harmonic I4
Current measurement module group 19 | 45160 | Float | RD | A | 11. harmonic I4
Current measurement module group 19 | 45162 | Float | RD | A | 13. harmonic I4
Current measurement module group 19 | 45164 | Float | RD | A | 15. harmonic I4
Current measurement module group 19 | 45166 | Float | RD | A | 17. harmonic I4
Current measurement module group 19 | 45168 | Float | RD | A | 19. harmonic I4
Current measurement module group 19 | 45170 | Float | RD | A | 21. harmonic I4
Current measurement module group 19 | 45172 | Float | RD | A | 23. harmonic I4
Current measurement module group 19 | 45174 | Float | RD | A | 25. harmonic I4
Current measurement module group 20 | 45176 | Float | RD | A | 1. harmonic I1
Current measurement module group 20 | 45178 | Float | RD | A | 3. harmonic I1
Current measurement module group 20 | 45180 | Float | RD | A | 5. harmonic I1
Current measurement module group 20 | 45182 | Float | RD | A | 7. harmonic I1
Current measurement module group 20 | 45184 | Float | RD | A | 9. harmonic I1
Current measurement module group 20 | 45186 | Float | RD | A | 11. harmonic I1
Current measurement module group 20 | 45188 | Float | RD | A | 13. harmonic I1
Current measurement module group 20 | 45190 | Float | RD | A | 15. harmonic I1
Current measurement module group 20 | 45192 | Float | RD | A | 17. harmonic I1
Current measurement module group 20 | 45194 | Float | RD | A | 19. harmonic I1
Current measurement module group 20 | 45196 | Float | RD | A | 21. harmonic I1
Current measurement module group 20 | 45198 | Float | RD | A | 23. harmonic I1
Current measurement module group 20 | 45200 | Float | RD | A | 25. harmonic I1
Current measurement module group 20 | 45202 | Float | RD | A | 1. harmonic I2
Current measurement module group 20 | 45204 | Float | RD | A | 3. harmonic I2
Current measurement module group 20 | 45206 | Float | RD | A | 5. harmonic I2
Current measurement module group 20 | 45208 | Float | RD | A | 7. harmonic I2
Current measurement module group 20 | 45210 | Float | RD | A | 9. harmonic I2
Current measurement module group 20 | 45212 | Float | RD | A | 11. harmonic I2
Current measurement module group 20 | 45214 | Float | RD | A | 13. harmonic I2
Current measurement module group 20 | 45216 | Float | RD | A | 15. harmonic I2
Current measurement module group 20 | 45218 | Float | RD | A | 17. harmonic I2
Current measurement module group 20 | 45220 | Float | RD | A | 19. harmonic I2
Current measurement module group 20 | 45222 | Float | RD | A | 21. harmonic I2
Current measurement module group 20 | 45224 | Float | RD | A | 23. harmonic I2
Current measurement module group 20 | 45226 | Float | RD | A | 25. harmonic I2
Current measurement module group 20 | 45228 | Float | RD | A | 1. harmonic I3
Current measurement module group 20 | 45230 | Float | RD | A | 3. harmonic I3
Current measurement module group 20 | 45232 | Float | RD | A | 5. harmonic I3
Current measurement module group 20 | 45234 | Float | RD | A | 7. harmonic I3
Current measurement module group 20 | 45236 | Float | RD | A | 9. harmonic I3
Current measurement module group 20 | 45238 | Float | RD | A | 11. harmonic I3
Current measurement module group 20 | 45240 | Float | RD | A | 13. harmonic I3
Current measurement module group 20 | 45242 | Float | RD | A | 15. harmonic I3
Current measurement module group 20 | 45244 | Float | RD | A | 17. harmonic I3
Current measurement module group 20 | 45246 | Float | RD | A | 19. harmonic I3
Current measurement module group 20 | 45248 | Float | RD | A | 21. harmonic I3
Current measurement module group 20 | 45250 | Float | RD | A | 23. harmonic I3
Current measurement module group 20 | 45252 | Float | RD | A | 25. harmonic I3
Current measurement module group 20 | 45254 | Float | RD | A | 1. harmonic I4
Current measurement module group 20 | 45256 | Float | RD | A | 3. harmonic I4
Current measurement module group 20 | 45258 | Float | RD | A | 5. harmonic I4
Current measurement module group 20 | 45260 | Float | RD | A | 7. harmonic I4
Current measurement module group 20 | 45262 | Float | RD | A | 9. harmonic I4
Current measurement module group 20 | 45264 | Float | RD | A | 11. harmonic I4
Current measurement module group 20 | 45266 | Float | RD | A | 13. harmonic I4
Current measurement module group 20 | 45268 | Float | RD | A | 15. harmonic I4
Current measurement module group 20 | 45270 | Float | RD | A | 17. harmonic I4
Current measurement module group 20 | 45272 | Float | RD | A | 19. harmonic I4
Current measurement module group 20 | 45274 | Float | RD | A | 21. harmonic I4
Current measurement module group 20 | 45276 | Float | RD | A | 23. harmonic I4
Current measurement module group 20 | 45278 | Float | RD | A | 25. harmonic I4